## Slide 1: Start Counter Update
W. U. Boeglin, L. Guo, E. Pooser
Prototype production
Timing studies SiPM
Summary and Plans

## Slide 2: Prototype Production
Material bought from Eljen: EJ-2123x 15 x 600 mm2, 5 pieces
Sent to Plastic Craft (PC)
Sample of 5 machined elements ordered
Production delayed due to flooding this spring
Price for initial setup ~5K, order confirmed
Ordered sample of 1mm clear fibers from Kuraray (they do not make 3x3mm2 and do not have samples of 2x2mm2

## Slide 3: Timing Studies
Further studies:
Moved source to the end of the scintillator bar to study the worst case scenario (~55mm from detector)
2 x 15 mm2 scintillator, no wrapping, VM2000 at end
‘Standard’ SiPM array (4x4 element each 3x3mm2)
Improvised cooling using dry ice in dark box

## Slide 4: SIPM Signals (Single Trig.)
Amplified signal:
rise time ~20ns
Amplified and differentiated
t =10ns, rise time ~11ns

## Slide 5: Previous Result
SiPM
2x15
3x15
3x12
Last measurements
Best current result

## Slide 6: Results
Temperature | s | Amplifier | Comment
Room temp. ~20oC | 1075 | LeCroy | No Shaping
Room temp. ~20oC | 875 | LeCroy & 1711 | 10ns diff.
~ -2oC | 810 | LeCroy & 1711 | 10ns diff.
~ -5oC | 780 | 1711 | 10ns diff., removed the LeCroy amp.
~ -5oC | 580 | 1711 | Optimized CFD
~ -5oC | 753 | 1711 | Leading edge, software corr.
~ -5oC | 930 | 1711 | No diff., leading edge

## Slide 7
s = 753
s = 930
s does not improve if ADC> 500

## Slide 8: Summary
Cooling of SiPM in a reproducible manner seems to be important, more testing necessary
Time resolution very sensitive on signal size
Rise time is crucial for good timing
Best resolution of 580ns for 2mm material implies that 400ns at the end should be feasible with 3mm material and wrapping

## Slide 9: Plans
Get faster detector
Test PC manufactured paddles with PM and SiPM
Wrap detector with reflective material: need to be able to get VM2000 like material!
Build small light guide to connect to photo detector
Test fiber light guide with PMT for fall back solution

## Slide 10
achievable
13%
7%
3%
<1%