TOF Performance
By:
Bradford Cannon
Florida State
University
GlueX
Collaboration
Meeting
October 9, 2015

---
Outline of Talk
- TOF Calibration Procedure
- TOF Performance
- TOF Issues
- To Do List

---
TOF timing calibration procedure
Developed by Benni Zihlmann see GlueX-Doc-2767
There are 6 steps in TOF timing calibration:
1) Event selection: Prepare root tree with TOF hits.
2) Time-walk correction: Correction of time shift due to varying pulse amplitudes with
    Respect to the discriminator threshold.
3) Mean Time Calibration: Determine relative Mean Time for all double ended modules.
4) Time Difference Calibration: Determine relative Time Difference for all double ended
    modules.
5) Double Ended Paddle Timing Offsets: Obtain timing offsets for each individual PMT
    channel from the module's Mean Time and Time Difference.
6) Single Ended Paddle Timing Offsets: Determine relative timing offsets for single-ended
    paddles using calibrated double ended modules.
Beni wrote a fully automated calibration toolkit which we have used successfully.
+<?>
MT= { t_{L} + t_{R}} over {2}
TD= { t_{L} - t_{R}} over {2}

---
1) Event Selection
- Only events with ADC and TDC hits
- Create separate root trees for both ADC and TDC data
- Create separate root trees for double and single ended paddles
- Some of the information we collect is: Number of Hits, Paddle and Plane location, MT, TD, and ADC integral.

---
2) Time-walk calibration
ADC hit time is used as t0 for TDC time
Characteristic dependence of Δ(tTDC-tADC) on the pulse integral is clearly seen;
It is fitted with a functional form:
Important parameter C is typically in the range from -0.75 to -0.85, i.e.,
A much steeper dependence than 1/sqrt(ADC).

---
Investigation of alternative t0
ADC time is determined as a crossing of a fixed threshold. It should have its own time-walk
Possible alternatives:
1. Use 6x6cm2 intersection region with the other plane (as per GlueX-doc-1719).
A single fit with ~350 parameters did not converge well
6cm uncertainty in position translates into >0.4 ns uncertainty in t0
2. Use signal reflections.
Assume that reflection time is constant
  (twice the cable length at ~170 ns)
Assume that a huge first pulse (which
  is needed to observe the reflection)
  has no time walk
Study the timing of the reflection
  a function of its magnitude
  (more on next slide)

---
Using reflections to correct time-walks
Events with 2 hits in the same channel (both TDC and ADC) have been selected
fADC Mode 8
Reflections which trigger TDC-registered
  are sitting on top of the signal's tail and
  therefore have different time-walk shape
fADC Mode 7 with Npeaks=3
We set NSA to 45 samples (or 180ns)
This window was too long and caused
   the ADC to miss the reflection
true reflections

---
3) Mean Time Calibration
Calculate the difference in MT for a reference paddle
in plane A with all double ended paddles in plane B
Plot the MT difference versus paddle number
Select a paddle number and project onto the x-axis
Fit the resulting peak to get the timing offset between
the two paddles
Repeat steps 1-4 for all paddles in plane A
Select paddle 18 as a reference paddle for plane A
Calculate the difference in MT for paddle 18 with all
other double ended paddles in plane A
Plot the MT difference versus paddle number
Select a paddle number and project onto the x-axis
Fit the resulting peak to get the timing offset
between the two paddles

---
4) Time Difference Calibration
Same procedure as MT, except we use TD instead
After we collect the TD peak positions, we can convert a
paddle number into its geometric position on the TOF
This result of this plot gives us a slope which is inversely
proportional to the effective velocity of light inside the
paddle
Average ceff = 15.7 cm/ns

---
5) Double Ended Paddle Timing Offsets
- Knowing the MT and TD offsets for each module allows us to get the timing offsets for individual PMT's by adding/subtracting the MT and TD equations
- On average, the timing offset is about 1 ns, with max offsets of 3.5 ns
MT= { t_{L} + t_{R}} over {2}
TD= { t_{L} - t_{R}} over {2}

---
6) Single Ended Paddle Timing Offsets
1. Apply walk corrections and timing offset corrections to all full length paddles
2. Select events with a TD < 0.5 ns so that the hit took place near the middle of the double ended paddles, and thus overlapping with the single paddles
3. Calculate MT for all double ended paddles that geometrically overlap with a single paddle
4. Calculate timing offset of the single paddle by subtracting the MT of a double ended paddle with the recorded time of single paddle

---
Timing Resolution Estimate
1. Select tracks with
   TOF hits in both planes
2. Compare measured
    tmean (aka tTOF) times
   in both planes.
3. The difference of 2
    independent tTOF
   measurements has
   a width of 136.2 ps.
  This translates into per-plane timing resolution of 136.3/√2 = 96 ps
   We are almost at the design goal!

---
TOF performance: the bottom line
C.Meyer Hadron 2015 presentation
Proton band is clearly separated all the way to 3 GeV/c
Pion and Kaon bands are distinguishable upto 1.5 – 2 GeV/c

---
Issue #1: TDC vs ADC timing
Hit location along module's length can be determined both from TDC and ADC Δt.
TDC times are calibrated; ADC times are as is.
At first glance, there is a good linear correspondence of TDC and ADC time differences.
Closer inspection reveals a slight curvature, with a Cosine-like behavior of the slope.
Near the center, hits locations are in agreement. Near the edges, they differ by ~7cm.
The origin of this effect is still mysterious.

---
Issue #2:ADC data: Attenuation length
ADC pulse integral vs.
TDC-determined hit position
along the module length.
Minimum-ionizing particles
are clearly seen
Taking slices at different locations and fitting
them with a Landau convoluted with a
Gaussian provides the dependence of the
observed MIP peak on the location of the hit.

---
Issue #2: ADC data: Attenuation length
        MIP peak vs location fit
- Blue and pink points are from left and
  right PMTs of the same module.
- Two exponents with different att.lengths
   (73 cm and 426 cm here) are required.
- Central area is also fitted with a single
  exponent (cyan) giving 218 cm length.
- For all modules, the fitted attenuation
  lengths are 75±8 cm and 506±90 cm.
        - Mean value for single
        attenuation fit is 225 cm
        - Past lab measurements
        (unwrapped modules in the
        black box with UV LED) had
        307±51 cm average length.

---
Issue #3: FADC250 integration range
- Most TOF data were obtained with
  FADC250 integration range values
  of NSB/NSA at 10/45 samples.
- In this study, actual waveforms of
  MIP hits at a fixed location have
  been integrated within varying range.
- The distribution of such integrals
   is then fitted with a Landau.
- First plot shows fit's MVP value
  for different NSA/NSB combinations.
- The second plot is Landau's width
   to max ratio. The best integration
   resolution (blue) is for any NSB
   below 5 and any NSA above 20.
- Therefore, using too high NSB is
   not necessary.

---
Issue #4: PMT Gain Balancing
Observed MIP energy deposition as a function of module's location (both PMTs)
Module (PMT) location in the TOF wall
Landau MPV for MIP peak at moulde's center

---
Issue #4: PMT Gain Balancing
- Observed energy deposition from minimum
  ionizing particles shows an apparent variation
  of PMT gains up to 30% from module to
  module.
- At FSU labs, the gains were balanced with
  10% accuracy.
- The variation is not random:
   a) Left/Right PMTs seems to be balanced;
   b) The farther a PMT from TOF center,
       the higher its gain appears to be.
Why 10% balanced became 30% unbalanced? Why unbalanced in an orderly fashion?
Hardware-related “theories”: a) “jitter” of different HV PS at FSU and JLAB; b) fringe magnetic
        field effect on PMT; c) something else.
Beam-related causes: a) average path-length through scintillator variations due to angles;
        b) average momentum variations with MIP change per Bethe formula.
Short-term plan: use cosmic data to confirm the effect. If it's real then re-adjust HV.

---
To-do list for the near future
- Collect TOF cosmic data to: a) verify PMT gains; b) independently check
   timing calibration (i.e., how straight are cosmic tracks after calibration).
- Study TOF efficiencies with collected beam data.
- Work on CAEN TDC calibration by B.Moffit is in progress.
- Improve TOF code in hdview2 (use of calibrations, TOFPoint in addition
  to TOFHit, etc.).
- ADC-to-energy calibration of TOF hits (requires Monte Carlo studies).