## Slide 1: Upgrade of the Forward Calorimeter (FCAL 2)
Replace inner part of the lead glass
calorimeter by PbWO4 crystals
(the goal is to install an array of 50 x 50 crystals)
Property | Improvement factor
Energy σ | 2
Position σ | 2
Granularity | 4
Radiation-resistance | 10
Significantly improves reconstruction of GlueX channels of interest and allows to
study rare  decays such as 0

## Slide 2
PWO Calorimeter Prototype: Compton Calorimeter
Array of 12 x 12 PbWO4 crystals
Beam hole: 2 x 2 crystals
Water cooling (18 C), nitrogen purge
LED-based gain monitoring system
Positioned on X-Y movable platform
Used during the PrimEx -  experiment in Hall D for reconstruction of Compton event (Spring 2019)
built in collaboration with Hall C

## Slide 3
20.5 mm x 20.5 mm x 200 mm
PbWO4 crystal wrapped with a
ESR reflective foil and Tedlar
Hamamatsu R4125 PMT
PMT divider and amplifier
designed at JLab
Compton Calorimeter: Module Assembly
PbWO4 crystal
PMT housing
Brass strip
PMT divider & Amplifier
Popov, et. al, IEEE proceedings, Nov 2012
FCAL insert will be position in the fringe field of solenoid magnet ( Bz < 50 G, Bxy < 10 G)
Magnetic shielding simulated with TOSCA and measured using Helmholz coils
Use soft iron PMT housing and acrylic light guides GlueX-Doc-4166
light collection studied using the Hall D pair spectrometer test setup
https://halldweb.jlab.org/wiki/index.php/Ligh_Guide

## Slide 4
Compton Calorimeter: Performance



Compton
Calorimeter
Reconstruction of Compton events
in PrimEx
compared with Hall B hybrid calorimeter
(built from SICCAS crystals produced more
than15 years ago)
Two cluster events
Clean reconstruction of Compton candidates
Energy Resolution

## Slide 5: Status of Calorimeter Construction
https://halldweb.jlab.org/wiki/index.php/FCAL_Insert
PbWO crystals from two vendors: SICCAS and CRYTUR
- Ordered 1000 crystals. About 200 crystals delivered to the lab
- perform various QA checks
Received from Hamamatsu 500 R4125 PMTs.
Finalizing design of the PMT active base
Ordered readout / trigger electronics, HVs for 1600 modules
Preliminary engineering design of the calorimeter frame and installation procedures
Have already fabricated several modules. Ready to start mass production
NIM A956 (2020) 163375

## Slide 6
Testing of PbWO4 Crystals
Dimension measurements
Light transmission measurement
Pair Spectrometer test setup in Hall D
Light yield measurement (Na22 source)

## Slide 7: Installation Schedule
We expect to start the detector installation after the Spring run in 2023
- the size of the insert may vary depending on availability of funds. The project is
funded by the Jefferson Lab. Applied for the NSF grans (PI, Prof. L. Gan, UNCW)
The FCAL 2 can be integrated to the Hall D detector in about 6 months
- the installation plan has been developed
( T. Whiltatch presentation at GlueX collaboration meeting, May 2020 )
- main installation steps: Disassemble FCAL modules, modify FCAL frame,
Install lead glass and PbWO4 modules, connect cables

## Slide 8: Backup Slides

## Slide 9
PbWO4 crystal wrapped with a ESR reflective foil and Tedlar
Brass tension strips (25 μm) are brazed to flanges and hold the module together
PMT is placed inside G-10 housing, wrapped with mu-metal foil
Compton Calorimeter: Module Assembly
PbWO4 crystal
PMT housing
Brass strip
PMT divider & Amplifier
New module features:
- Soft iron PMT housing
- Acrylic light guide

## Slide 10: Upgrade of the Forward Calorimeter
Replace inner part of the lead glass calorimeter by PWO crystals
Property | Improvement factor
Energy σ | 2
Position σ | 2
Granularity | 4
Radiation-resistance | 10
MC simulation:
- Reconstructed 0 events during
1 day of taking data

## Slide 11: Project Overview
PbWO Crystals
Module Design
PMT and Active Base
Electronics
Frame Design / Engineering
https://halldweb.jlab.org/wiki/index.php/FCAL_Insert

## Slide 12: PWO Crystals: Procurement Status
64 crystals received from CRYTUR, all accepted
500 crystals ordered from SICCAS
- 183 crystals received (60 Hall C replacement + 123 Hall D)
- All produced crystals were checked at SICCAS by Wuhan group
Preparing a new contract for ordering 500 more crystals
- Paperwork is in place, coordinated with SICCAS (DOE approval will be needed)
- Expected delivery of all crystal - end of January 2021 (150, 150, 200 crystals)
Additional financial support - UNCW NSF grant is under review

## Slide 13: PWO Crystals
192 SICCAS crystal received in the middle of April (first batch)
Performed quick inventory and visual inspection check
Visually all crystals look Ok

## Slide 14
Crystal QA Tests
Visual inspection
Measurement of crystal dimensions
Measurement of light transmission
Measurement of light yield
Crystal properties and testing procedures
Nucl.Instrum.Meth. A956 (2020) 163375

## Slide 15
Visual Inspection
NPS Lab
Check crystal color
Use laser pointer to observe cloudy regions
Check edge quality for scratches and chips
Clear crystal
Cloudy regions
yellow discoloratiuon

## Slide 16
Crystal Dimensions
NPS test Lab
- Granite table
- Vertical height gage
Crystal measurement for NPS
(CCAL)
The lab and setup will be used to measure the Hall D crystals as well.

## Slide 17: Light Transmission Measurement
Setup at Jefferson Lab (Carl Zorn’s lab)
- automatic operation for measurements
- will be used by Hall D
Reproduce transmittance measurements
performed with the setup at CUA
Measurement procedure:
Correlation between crystal
transmittance and crystal color
crystal transmittance as a function of
wavelength
https://halldweb.jlab.org/wiki/index.ph/Measurement_Light_Transmission
crystal with yellow
discoloration

## Slide 18
Crystal Transmittance

## Slide 19: Light Yield Setup
New setup at TEDF (similar to NPS lab)
- crate with readout electronics
- CODA running on local PC
Trigger
LaBr(Ce) scint
PWO
Na-22 β+ and gamma source
- two 0.511 MeV photons. One used
for the trigger

## Slide 20: Light Yield Setup

## Slide 21
Crystal Testing Using Pair Spectrometer
Setup allows to test 12 modules, installed downstream the Pair Spectrometer
Run setup in parallel with GlueX data taking

## Slide 22
Crystal Testing Using Pair Spectrometer
‘Projection’ of the crystal to the PS
hodoscope
GlueX-doc-3590,
V. Berdnikov, A.Somov, J. Crafts
The setup has been aligned
Crystals can be easily installed during opportunistic access to the hall

## Slide 23
PbWO4 crystal wrapped with a ESR reflective foil (60 μm) and Tedlar (40 μm)
Brass tension strips (25 μm) are brazed to flanges and hold the module together
PMT is placed inside G-10 housing, wrapped with mu-metal foil
Module Assembly
PbWO4 crystal
PMT housing
Brass strip
PMT divider & Amplifier
CCAL module (design based on HyCal, Hall B)
New module features:
- Soft iron PMT housing
- Acrylic light guide

## Slide 24
Finalizing Module Design
Magnetic shielding studies
https://halldweb.jlab.org/wiki-private/index.php/Private:Eta_Rare_Decay,
TOSCA simulation
GlueX-Doc-4166, Shielding studies using Helmholz coils
Light guides
LG length: 3.5 cm (magnetic shielding)
Diameter: 18.5 mm
Glue LG to the PMT (UV glue)
Couple LG to the crystal using Silicon
cookie (RTV 615)

## Slide 25
Light losses of 25 % (compared with the direct coupling of PMT to the crystal)
Can improve light collection by a few present if use LG with the diameter larger
than 18 mm
Light Guide Studies
https://halldweb.jlab.org/wiki/index.php/Ligh_Guide
Installed light guides on 2 modules in the PS setup
- Studied relative light collection for several LG configurations
(unwrapped, wrapped with ESR, coupled to the crystal using optical grease
and Silicon cookie, use LG with different diameters)
- Beam tests using light guides with 2 diameters: 18 mm and 18.8 mm

## Slide 26
Clean Box to Produce Silicon Cookies
Use plexiglas plate with holes, produce 36 cookies at the same time
Cure silicon at room temperature (no heating)
Need to redo the fixture according to the final LG size of 18.5 mm

## Slide 27
Fabrication Procedures
Improving strap brazing setup
- extend base plate to accommodate 20 cm long crystals
- new heating nozzle caps, implement temperature control
- clean flanges before brazing (JLab chemist group) with acetic acid
Checked strength of 25 m strips, force analysis

## Slide 28
Item | Availability | Comments
PMT housing | 20 | Production rate 100/month
 |  | Produced at CMU, Annealed at Jlab
350 um mu-metal schileding cylinder | 20 | Ordered for 1600 modules
50 um mu-metal foil | 1600 modules | 
Light Guides | 5 | Wait for evaluation
 |  | Production rate: fast
ESR | 1600 modules | 
Flanges | 20 | Ordered for 500 modules
Brass strips | 1600 modules | 
PMT tension plate (G10) | 1600 modules | 
Kapton, Tedlar, Screws, Tape | Available | 
Preparation for Module Fabrication
Fabrication tools are ready (ESR pre-shaping, brazing, LG gluing, etc)
Most fabrication procedures are finalized (used during CCAL fabrication)
https://halldweb.jlab.org/wiki-private/index.php/CompCal_Construction
OSP for module fabrication (GlueX-doc-4379)
Ready to start module fabrication !

## Slide 29: Clean Room for Module Fabrication (TEDF)

## Slide 30: PMT and Active Base
Received 500 R4125 PMTs
Tested diameter of 250 PMTs (about 15 % have somewhat smaller diameter, trade in
with Hall C)
Check PMTs using an LED setup at JLAB
- have already tested a few hundred PMTs for CCAL and NPS, no defective PMTs found
https://halldweb.jlab.org/wiki-private/index.php/PMT_testing
PMT will be instrumented with an Active Base
- amplifier gain – 24 (reduce HV, small anode current)
- stable gain at high rate (up 2 – 3 MHz)
The Active Base was installed on CCAL and used during
PeimEx run in spring 2019
New Active PMT Base for PbWO4
(Popov, H. Mkrtchyan, presentation at IEEE, Nov 2012)

## Slide 31
PMT Active Base Linearity
Found somewhat non-linear performance of the active base during CCAL testing
The linearity can be improved by reducing the amplifier gain and increasing HV
Some tuning of the PMT base are required
V. Popov modified the active divider (increase the anode current, changed resistors
in the amplifier chain )
- three modified bases were installed on CCAL and tested in January 2020
https://logbooks.jlab.org/entry/3780745
700 V
2 V FADC range
825 V
2 V FADC range
Original PMT base, gain 25
Modified base, gain 6
Linearity of the FADC peak amplitude

## Slide 32: Sum of FADC amplitudes in 400 ns time window
30 muA / 24
15 muA / 24
5 muA / 24
Example:
Current seen by FADC in the 3rd row: 30 muA
Current through the divider: 30 muA / 24 =
1.25 muA
Current estimates for FCAL:
Beam hole: 2x2
1 layer covered by an absorber
The 3rd layer in FCAL – 5th layer
In CCAL (factor of two difference in Z)
Solid angle correction is 2x2 = 4
If no amplifier used: 5 x 4 = 20 muA
Discussion:
- Use active bases with the gain 3 (or 6) in the inner layers of the FCAL 2
- Don’t need amplifiers on other layers
Install 24 active bases with modified dividers (gain 3 and 6) on the CCAL inner layers
Notes: *

## Slide 33
Install Modified Bases on 24 CCAL Modules
Sssds https:/ /logbooks.jlab.org/entry/3807923, 3807924, 3808592
Voltages calibrated using a snake scan (CCAL as TAC)
Use CCAL during GlueX production for about 1.5 days
- stable performance
- looking on data now
Implemented CCAL to the trigger
x24 gain 670 V 260 A
x 3 gain 870 V 980 A (modified base)

## Slide 34
Performance of CCAL Modified Bases in GlueX
FADC channel
FADC channel
8 mV ADC Threshold (45 MeV)
Stable performance during data taking
Checking LED response (gain stability) now
Use CCAL with GlueX runs (instrument with dividers with gain x3)

## Slide 35
Next Steps
Perform detailed check of 192 crystals received from SICCAS
(visual inspection, dimension measurements, optical transmission, light yield)
- check transmission and light yield only for preselected crystals from future batches
- check some crystals using PS setup (?)
Check/order more light guides when the lab will be opened again
Start fabricating modules. Currently have parts for 5 modules (5 light guides)
Check PMTs (check dimensions of 250 PMTs, check gains for 100 – 150 PMTs)
Finalize design of the active base
- Test active bases with the gain of 3 during PrimEx run (to be coordinated)
- Fernando is currently working on the PCB layout (use dip switch or soldering point
to change between gain 3 and bypassed amplifier). Some coordination is needed
Test fabricated modules using the PS setup

## Slide 36
Backup Slides

## Slide 37
Energy Resolution Measured with CCAL



NPS prototype
HyCal (Hall B)
NPS Prototype
3x3 array
bypassed amplifier
Relatively good energy resolution, which is consistent with the resolution of the
Hall B HyCal calorimeter
Rate and PMT anode current have to be considered (some measurements were performed)
Preliminary plot, will be updated
More details can be found in
GlueX-doc-3590, GlueX-doc-3998, V. Berdnikov, A.Somov, J. Crafts

## Slide 38: Measurement of the PMT Anode Current
Threshold
FADC Pedestals
What amplifier gain do we need ?
Use CCAL to estimate the PMT anode current
Collect FADC waveforms using random and PS triggers
Two beam currents: 350 nA and 200 nA
- no data sparsification (no threshold applied)
- estimate current using waveforms acquired
with the random trigger
- Use PS trigger to monitor beam trips
Sum up FADC amplitudes above the threshold
(1 count above the baseline) in a 400 ns window
(100 fadc samples)
Estimate anode current as follows:
A = 5 x 10-4 V / 50 Om <A (FADC) > / 100 samples / Gain (Amp)
Notes: *

## Slide 39: Sum of FADC amplitudes in 400 ns time window
Notes: *

## Slide 40
Magnetic Shielding
20.6 mm
Housing D = 19.9mm
Mu-metal foil: 350 m
50 m
Steel 1020 (not annealed)
(350 m foil welded and
annealed)
https://halldweb.jlab.org/wiki-private/index.php/Private:Eta_Rare_Decay
1
2
3
4
5
Simulated magnetic shielding using TOSCA (S. Glamazdin)
Build a 3x3 prototype and measured magnetic field
and PMT performance using Helmholz coils
GlueX-Doc-4166
No significant degradation of performance of the shielded PMT observed

## Slide 41
Light Guide Studies
LG length: 3.5 cm (magnetic shielding)
Diameter: 18 mm
Glue LG to the PMT (UV glue)
Couple LG to the crystal using Silicon
cookie (RTV 615)

## Slide 42: PWO Crystals
https://halldweb.jlab.org/wiki-private/index.php/Crystal_Specifications
PWO crystals: 20.5 mm x 20.5 mm x 200 mm ( FCAL insert: an array of 40 x 40 crystals)
Two vendors:
SICCAS: Shanghai Institute of Ceramics
CRYTUR: Integrated crystal based solution, Czech Republic
Procurement status:
64 crystals received from CRYTUR, all accepted
500 crystals ordered from SICCAS
- 160 crystals are ready to be shipped to us
- All produced crystals were checked at SICCAS by Wuhan group
Preparing a PR for 500 more crystals
Collaborate with the Hall C NPS group
- building PbWO calorimeter, already have about 1100 crystals
Many crystal QA tests were performed during CCAL fabrication

## Slide 43
Test up to 12 PWO crystals (modules) at the same time
Install test bench on the platform behind the PS
Take data in parallel with GlueX (fadc in the ST crate)
Crystal Testing Using Pair Spectrometer
Module assembly for testing
PMT
Crystral
Divider
12 modules
on the test bench
Can measure light yield,
energy resolution
Study light guides

## Slide 44
Solenoid Magnetic Field
FCAL lead glass
face
lead glass
downstream end
Longitudinal field at beamline (r = 0)
B Z (Gauss)
Z (cm)
The field was also measure in front of the FCAL (TOF)
- good agreement with the field map
GlueX-doc-4166-v3

## Slide 45
Solenoid Magnetic Field
Insert
Insert
20 cm downstream the FCAL lead block
B Z at Z = 690 cm
B x at Z = 690 cm
B Z at Z = 690 cm
B Z (Gauss)
B X (Gauss)

## Slide 46
Proposed Position of
the FCAL II
Lead Glass L = 45 cm
PWO L = 20 cm
Beam
Z = 625 cm
Z = 670 cm
Lead Glass L = 45 cm
PWO L = 20 cm
Beam
Z = 660 cm
Z (CCAL) = 690 cm
Two FCAL outer rings cannot be calibrated using
photons from the target
Move FCAL downstream the beam by 35 cm
Maximum field at Z = 690 cm 44 Gauss
Bz (r = 0) = 44 Gauss, Bx (r = 40 cm) = 8 Gauss
FCAL modules from the inner part (replaced by PWO)
can be used to instrument one extra outer layer
Z = 635 cm
Z = 655 cm
move downstream by 35 cm
10 cm

## Slide 47
Lead Glass L = 45 cm
PWO L = 20 cm
Solenoid Magnetic Field
Z (cm) | Z (cm) | Bz (r = 0 cm) G | Bx (r = 40 cm ) G
625 | FCAL upstream face | 87 | 20
645 |  | 68 | 15
670 | FCAL downstream face | 54 | 10
690 |  | 44 | 8
Beam
625 cm
670 cm
635 cm
655 cm

## Slide 48
FCAL Acceptance
Move FCAL platform downstream
the beam by 35 cm
FCAL aperture is decreased by
r = 7.5 cm

## Slide 49
PMT Housing: 3 x 3 Prototype
PMT length (no leads): 8.73 cm
LG length: 3.5 cm
Total: 12.23 cm
Flange (+gap): 0.8 cm
12.23 – 0.8 = 11.43 cm
PMT housing length 10.5 cm
can make PMT housing 6 – 7 mm longer
20.6 mm
Average crystal width: 20.55 mm ( sigma  25 m )
ESR thickness: 65 m
20.55 + 0.13 = 20.68 mm (+ Tedlar)
cannot make housing wider
PMT diameter < 19.0 mm
Housing D = 19.9mm
Mu-metal foil: 350 m
50 m
Steel 1020 (not annealed)
(350 m foil welded and
annealed)

## Slide 50
Field Measurement
Longitudinal Field
Transverse Field
Transverse field can be easily shielded
PMT is much less sensitive to the longitudinal field
Expected maximum
transverse field is
8 Gauss
(at Z = 690 cm)

## Slide 51
PMT Response to LED (no Shielding)
Longitudinal Field

## Slide 52
PMT Response to LED (no shielding)
Transverse Field

## Slide 53
LED Setup
Check PMT response for different field values using an LED system

## Slide 54
PMT Response to LED ( with Shielding)
Bz
Bz
Response to the longitudinal field

## Slide 55
PMT Response to LED ( with Shielding )
Rotate prototype by about 10 degrees
- longitudinal and transverse components
Total Field (G)
Total Field (G)
GlueX-doc-4166-v3

## Slide 56
Magnetic shielding (TOSCA Simulation)
1
2
3
4
5
A. Glamazdin
https://halldweb.jlab.org/wiki-private/index.php/Private:Eta_Rare_Decay
Magnetic field shielding simulated
with TOSCA
Field measurements were performed
in Hall D and using Helmholz coils
Realistic 3x3 prototype (PMT housing)
is ready
- will be tested using Helmholz coils
shortly

## Slide 57
19x19 Hybrid FCal with Solenoid
P=5cm, µ-shield 360µm
Sol-1: x=y=3.125cm
Sol-2: x=y=19.925cm
Sol-3: x=y=38.825cm
Sol-4: x=36.725, y=38.825cm
PMT’s face
(3cm guide)
PMT’s face
(4cm guide)
PMT’s face
(5cm guide)
Magnetic field can be reduced below 0.5 Gauss

## Slide 58: Readout Electronics
See talk by Fernando
Use standard readout (fadc250) and trigger electronics. Will require
5 new VSX crates
VXS crates - ordered
FADC 250 - ordered
Procurement of CTP (VTP) is in process
Trigger Interface modules - ordered

## Slide 59: Frame Design
See talk by Tim
Regular meetings with engineering group
https://halldweb.jlab.org/wiki/index.php/FCAL_Insert_Design_Meetings
Working on the frame design
Cooling
Light monitoring system
Electronics patch panels
Stephanie, Keith Harding

## Slide 60: Di-photon Yield
Runs
Runs
Run period: 1 day
Run period: 2 days
Relatively stable
performance
Problematic modules in FCAL
Compute tagged flux for all PrimEx production
runs (CCDB)
Flux normalized yield of di-photons
FCAL XY of clusters

## Slide 61: CCAL Light Monitoring System
Runs
Runs
1st run
last run
Reference PMT receives light from
LED and Alpha source
- good LED stability measured
with respect to Alpha,  0.2 %
A (LED) / A (Alpha)
A (LED)
1st run
last run
LED response in various CCAL
modules
2 %

## Slide 62: PMT Active Base Performance
Observed non – linear response of the calorimeter (couple of percents)
during commissioning run in December 2018
- somewhat worse energy resolution than expected
FADC amplitude as a function of the beam energy
10 GeV : 3200 FADC channels
- PMTs were operated at relatively small HVs (650 – 700 V), HV recommended
by Hamamatsu is about 1 kV
Changed FADC voltage range to 2 V and increase HV by 50 V in January
Study linearity using calibration runs (CCAL in the beam)
Rate < 30 kHz

## Slide 63
570 V
750 V
Signal attenuator (0.1)
Calorimeter CELL = -4, -4
Linearity as a Function of Operational HV
630 V
680 V
Spring, 2019
Fall, 2018
Linearity is improved with the HV increase

## Slide 64: Test Modified Active Base with the Bypassed Amplifier
Install 9 PMT bases with
the bypassed amplifier
two bases with the reduced gain:
gain 3 and 6

## Slide 65
Linearity: ADC Peak Amplitude
Cell 4, 4
Cell 4, 4
Cell 4, 4
CCAL Spring 2019, gain 24
750 V
2 V range
Bypassed amplifier
810 V
0.5 V range
gain 3
850 V
2 V range
gain 6
800 V
2 V range
Cell -5, 4
Cell -5, 5
Bypassed amplifier
960 V
2 V range
Vpeak = 1.7 V (3501 cnt)
Cell 4, 4
Vpeak = 1.6 V (3228 cnt)
Divider limitations
Vpeak = 0.4 V ( 3474 cnt)
Vpeak = 1.2 V ( 2480 cnt)
Vpeak = 1.6 V ( 3359 cnt)

## Slide 66
Hall D Compton Calorimeter
Array of 12x12 PbWO4 crystals
Beam hole: 2 x 2 crystals
Tungsten absorber covers the inner
most layer (taken from HyCal)
Water cooling (minimum 5 C),
nitrogen purge
LED-based gain monitoring system
Positioned on X-Y movable platform
Hamamatsu R4125 PMT
HV dividers with amplifiers, designed by Vlad Popov
V. Berdnikov, T. Horn,
H. Voskonyan, H. Mkrtchyan
Collaboration between Hall D and Hall C (NPS prototype)

## Slide 67
Calorimeter Fabrication
Cooling Plates
PMT housing

## Slide 68: Light Monitoring System
Two reference PMTs are installed to monitor LED stability
Each PMT receives light from the LED fiber and YAP:CE Am light source
LED and Alpha-source triggers are used during data taking (special trigger types)
Reference PMT FADC amplitudes
PMT1
PMT2

## Slide 69: Rate per Counter (PrimEx D Conditions)
Maximum rate per counter
is 250 kHz
(at 30 MeV threshold)
Rates are in a good agreement with the simulation for PrimEx D
250 kHz
FADC scalers

## Slide 70
Linearity: ADC Peak Amplitude
(Bypassed amplifier, 1000 V, 2 V range, Spring 2019)
Cell 4, 4
Better linearity at small amplitudes
Negative non-linearity at large amplitudes