## Slide 1: Polishing and Testing
07/10/14

## Slide 2: Previous Polishing Methods

## Slide 3: Previous Polishing Methods (Cont.)
Scintillator and buffing wheel are kept cool and wet at all times
  Achieved with low buffing speeds (< 1500 rpms)
Polish in 20 minute intervals (~1000 rpms)
  10 min in straight section -> 0.05 μm Buehler Alumina
  10 min in bend and nose -> 0.05 μm Buehler Alumina
Polish both sides of scintillators with equal time and attention

## Slide 4: Paddle 27 Results (Old Polishing Methods)
These data have not been corrected for the trigger (≈ 30 ps reduction)
Paddle 27 is one of the better bars in the lot of 50
Average Time Resolution
  Unpolished: 361 ps
  Trial 1: 332 ps
  Trial 2: 339 ps
  Trail 3: 334 ps
Polishing after 2 trials rid the surface of virtually all fine scratches
Possible to recover ≈ 25 – 40 ps with previous polishing methods

## Slide 5: Current Polishing Methods
Same basic setup as before
Scintillator and buffing wheel are kept cool and wet at all times
  Achieved with low buffing speeds (< 1500 rpms)
Polish in 1 hour intervals (~1000 rpms)
  First polish with diluted 0.3 μm Allied Alumina
    30 min in straight section
    30 min in bend and nose
    If Comet Tails (< 2 mm scratches incurred during the polishing procress) exist then 0.3 μm Alumina is applied to a glass cleaning cloth and polished by hand until the micro-scratches are removed
  Second polish with diluted 0.05 μm Allied Alumina
    30 min in straight section
    30 min in bend and nose

## Slide 6: Paddle 5 Results (New Polishing Methods)
These data have not been corrected for the trigger (≈ 30 ps reduction)
Paddle 5 is one of the worst bars seen to date
  High density scratches in the nose and straight section
  Numerous deep scratches along the entire length of scintillator
Average Time Resolution
  Unpolished: 434 ps
  0.3 μm Alumina: 409 ps
  0.05 μm Alumina: 375 ps
Only scratches that remain on surface are the deep scratches that will most likely be unrecoverable
Possible to recover ≈ 25 – 75 ps with current polishing methods

## Slide 7: Initial Tests of 30 Unpolished Paddles
These data have not been corrected for the trigger (≈ 30 ps reduction)
Large Spread of timing resolutions in the bend is due to varying magnitudes of high density scratches in the bend region
The nose region also suffers from a large spread in timing resolution
  This is a direct result of the bend region
  The nose section also suffers from a varying magnitude of scratch damage
Polishing the worst bar (paddle 5) has made it comparable to the performance of our best bars

## Slide 8: Initial Tests (Cont.)
These data have not been corrected for the trigger (≈ 30 ps reduction)
Top: Max Time Resolution
Middle: Avg Time Resolution
Bottom: Min Time Resolution
These data allow us to easily determine the worst bars in the lot of 50
≈ 15 % (5) of the paddles are “unusable” to date
  These bars will be used for polishing technique development
Once a technique is finalized:
  Polish the good scintillators
  Wrap them in AL Foil
  Fully test each scintillator