## Sheet1
| Consistency | High Voltage | Laser Frequency (Min1 Mean, Min2 Mean) (negative) (mV) | [9 empty] | Standard Deviation (Output 1, Output 2) (mV) | [9 empty] | Fall Time at 5 MHz (Output 1, Output 2) |  | Current at 10 kHz | Distance between | Snapshot taken | Testing done by and date | Notes | Date taken to CUA
PMT Serial # | 800 mV @900 V | V | 10 kHz | 10 kHz | 1 MHz | 1 MHz | 2 MHz | 2 MHz | 5 MHz | 5 MHz | 10 MHz | 10 MHz | 10 kHz | 10 kHz | 1 MHz | 1 MHz | 2 MHz | 2 MHz | 5 MHz | 5 MHz | 10 MHz | 10 MHz | ns | ns | μA | laser/edge PMT
CM9639 |  | 850 V | 339 | 338 | 350 | 348 | 369 | 368 | 391 | 390 | 367 | 361 | [10 empty] | 1.479 | 1.519 | 759.95 | n/a | yes | Natalie on 10/14/13 | no fire resistant | 2013-10-14
|  | 900 V | 450 | 450 | 501 | 503 | 528 | 536 | 585 | 585 | 524 | 523 | [10 empty] | 1.453 | 1.459 | 807.4
|  | 1000 V | 1110 | 1100 | 1150 | 1140 | 1190 | 1180 | 1200 | 1210 | 778 | 781 | [10 empty] | 1.366 | 1.34 | 899.55
|  | 1100 V | 2160 | 2110 | 2170 | 2130 | 2240 | 2200 | 1630 | 1650 | 882 | 895 | [10 empty] | 1.146 | 1.642 | 991.7
CM9594 |  | 850 V | 365 | 368 | 379 | 381 | 399 | 399 | 426 | 425 | 383 | 383 | [10 empty] | 1.51 | 1.512 | 759.5 | n/a | yes | Natalie on 10/14/13 |  | 2013-10-14
|  | 900 V | 550 | 558 | 570 | 575 | 598 | 605 | 640 | 643 | 552 | 550 | [10 empty] | 1.433 | 1.38 | 806
|  | 1000 V | 1190 | 1200 | 1240 | 1250 | 1290 | 1300 | 1280 | 1290 | 796 | 787 | [10 empty] | 1.319 | 1.34 | 899.1
|  | 1100 V | 2290 | 2310 | 2310 | 2320 | 2360 | 2370 | 1750 | 1680 | 880 | 878 | [10 empty] | 1.26 | 1.2 | 991.15
CM9592 |  | 850 V | 336 | 335 | 356 | 354 | 372 | 369 | 395 | 392 | 364 | 360 | [10 empty] | 1.441 | 1.505 | 760 | n/a | yes | Natalie on 10/14/13 |  | 2013-10-14
|  | 900 V | 518 | 515 | 551 | 447 | 578 | 574 | 616 | 612 | 537 | 532 | [10 empty] | 1.405 | 1.431 | 806.8
|  | 1000 V | 1100 | 1090 | 1140 | 1130 | 1180 | 1160 | 1190 | 1180 | 789 | 789 | [10 empty] | 1.31 | 1.361 | 899.5
|  | 1100 V | 1990 | 1960 | 2030 | 2000 | 2100 | 2080 | 1670 | 1680 | 820 | 825 | [10 empty] | 1.25 | 1.25 | 991.55
CM9593 |  | 850 V | 315 | 314 | 325 | 323 | 338 | 337 | 365 | 364 | 333 | 330 | [10 empty] | 1.496 | 1.455 | 758 | n/a | yes | Natalie on 10/14/13 |  | 2013-10-14
|  | 900 V | 525 | 528 | 464 | 468 | 494 | 496 | 532 | 534 | 487 | 484 | [10 empty] | 1.44 | 1.39 | 805
|  | 1000 V | 1040 | 1040 | 1090 | 1090 | 1140 | 1150 | 1150 | 1150 | 787 | 782 | [10 empty] | 1.34 | 1.35 | 898.25
|  | 1100 V | 1980 | 1970 | 2000 | 2000 | 2070 | 2070 | 1650 | 1640 | 898 | 902 | [10 empty] | 1.186 | 1.234 | 990.4
CM9571 |  | 850 V | 356 | 351 | 365 | 363 | 302 | 299 | 360 | 357 | 358 | 352 | [10 empty] | 1.67 | 1.68 | 760 | n/a | no | Franz on 10/24/13 |  | 2013-10-26
|  | 900 V | 403 | 399 | 401 | 398 | 442 | 435 | 538 | 532 | 498 | 494 | [10 empty] | 1.64 | 1.66 | 806
|  | 1000 V | 821 | 812 | 865 | 857 | 955 | 947 | 1110 | 1090 | 739 | 719 | [10 empty] | 1.73 | 1.71 | 907
|  | 1100 V | 1570 | 1550 | 1600 | 1590 | 1680 | 1670 | 1590 | 1520 | 776 | 748 | [10 empty] | 1.84 | 1.85 | 993
CM9791 | yes | 850 V | 708 | 709 | 721 | 722 | 715 | 714 | 707 | 708 | 565 | 568 | 67 | 68 | 59 | 59 | 58 | 57 | 67 | 68 | 58 | 59 | 1.436 | 1.411 | 758.8 | 15.5 inches | yes | Natalie on 11/7/13 | output 1 problem | taken by Natalie on 11/14/13
|  | 900 V | 1090 | 1090 | 1090 | 1090 | 1080 | 1080 | 1020 | 1020 | 785 | 786 | 80 | 80 | 76 | 76 | 72 | 75 | 92 | 93 | 118 | 117 | 1.3 | 1.3 | 806 |  |  |  | (see red)
|  | 1000 V | 2160 | 2150 | 2170 | 2170 | 2150 | 2160 | 1520 | 1530 | 831 | 828 | 162 | 164 | 157 | 155 | 169 | 161 | 144 | 146 | 179 | 180 | 1.274 | 1.079 | 898.25
|  | 1100 V | 3670 | 3670 | 3550 | 3540 | 3340 | 3330 | 1170 | 1540 | 491 | 846 | 197 | 198 | 272 | 264 | 213 | 215 | 125 | 150 | 110 | 141 | 1.29 | 1.216 | 990
|  |  | 10 kHz | 10 kHz | 1 MHz | 1 MHz | 2 MHz | 2 MHz | 5 MHz | 5 MHz | 10 MHz | 10 MHz | 10 kHz | 10 kHz | 1 MHz | 1 MHz | 2 MHz | 2 MHz | 5 MHz | 5 MHz | 10 MHz | 10 MHz | 5 MHz | 5 MHz | 10 kHz
CM9393 | yes | 850 V | 582 | 573 | 586 | 575 | 568 | 567 | 553 | 540 | 478 | 472 | 40.7 | 45.2 | 50.9 | 50.8 | 45.2 | 44.8 | 51.3 | 51.8 | 55.2 | 51.6 | 1.332 | 1.447 | 758.9 | 15.25 inches | yes | Natalie on 11/7/13 |  | taken by Natalie on 11/14/13
|  | 900 V | 915 | 901 | 925 | 913 | 900 | 881 | 836 | 846 | 654 | 642 | 75 | 79.2 | 80.5 | 77.7 | 77.1 | 77.2 | 74.8 | 73.1 | 70.6 | 61.6 | 1.382 | 1.281 | 805.65
|  | 1000 V | 1830 | 1820 | 1820 | 1800 | 1810 | 1790 | 1480 | 1460 | 812 | 816 | 134 | 135 | 116 | 114 | 122 | 124 | 136 | 127 | 147 | 142 | 1.26 | 1.287 | 899.45
|  | 1100 V | 3060 | 3040 | 2980 | 2960 | 2930 | 2900 | 1670 | 1710 | 748 | 842 | 185 | 188 | 162 | 164 | 186 | 189 | 216 | 210 | 172 | 146 | 1.011 | 1.204 | 990.3
CM9635 | yes | 850 V | 684 | 680 | 685 | 680 | 678 | 674 | 636 | 638 | 532 | 533 | 54.9 | 53.1 | 55 | 55.4 | 67.2 | 65 | 62.5 | 60.4 | 50.8 | 52.7 | 1.375 | 1.333 | 760.1 | 15.5 inches | yes | Natalie on 11/7/13 |  | taken by Natalie on 11/14/13
|  | 900 V | 1080 | 1070 | 1050 | 1040 | 1060 | 1050 | 964 | 963 | 665 | 663 | 83.3 | 85.8 | 93.2 | 95.4 | 85.4 | 86.2 | 83.8 | 77.3 | 83 | 84 | 1.409 | 1.337 | 806.95
|  | 1000 V | 2110 | 2090 | 2100 | 2070 | 2020 | 2000 | 1560 | 1570 | 786 | 783 | 126 | 124 | 158 | 151 | 121 | 119 | 161 | 156 | 113 | 112 | 1.257 | 1.125 | 899.55
|  | 1100 V | 3540 | 3500 | 3520 | 3470 | 3230 | 3220 | 1520 | 1600 | 730 | 706 | 196 | 197 | 244 | 231 | 185 | 182 | 238 | 186 | 196 | 201 | 1.153 | 1.192 | 991.45
No Label-1 | yes | 850 V | 658 | 657 | 665 | 661 | 671 | 667 | 626 | 624 | 513 | 510 | 58.2 | 59.4 | 51.1 | 55.3 | 65.2 | 67.4 | 62.4 | 60.7 | 55 | 53.4 | 1.4 | 1.465 | 758.7 | 15.5 inches | yes | Natalie on 11/8/13 |  | taken by Natalie on 11/14/13
|  | 900 V | 1030 | 1030 | 1020 | 1020 | 1030 | 1030 | 934 | 931 | 657 | 658 | 86.3 | 88.2 | 95.3 | 95.8 | 82.2 | 84.5 | 80 | 78.9 | 74 | 74.7 | 1.385 | 1.403 | 805.45
|  | 1000 V | 2060 | 2060 | 2030 | 2030 | 2000 | 2000 | 1550 | 1550 | 787 | 783 | 141 | 137 | 128 | 128 | 121 | 117 | 136 | 137 | 140 | 148 | 1.181 | 1.215 | 898.45
|  | 1100 V | 3350 | 3340 | 3370 | 3380 | 3150 | 3150 | 1640 | 1720 | 785 | 825 | 311 | 310 | 193 | 192 | 200 | 201 | 206 | 204 | 201 | 197 | 1.117 | 1.12 | 990.2
No Label-2 | yes | 850 V | 661 | 599 | 635 | 572 | 672 | 602 | 631 | 567 | 512 | 460 | 48.1 | 43.5 | 50.8 | 45.2 | 82.6 | 72.2 | 57 | 48.1 | 58.6 | 53 | 1.34 | 1.418 | 760.1 | 15.5 inches | yes | Natalie on 11/8/13 | large difference | taken by Natalie on 11/14/13
|  | 900 V | 1030 | 930 | 990 | 886 | 983 | 880 | 923 | 821 | 678 | 612 | 90.6 | 80.8 | 90.5 | 70.8 | 77.5 | 66.3 | 80.5 | 70.9 | 77.9 | 63.5 | 1.367 | 1.333 | 807.5 |  |  |  | between O1 and O2
|  | 1000 V | 2010 | 1810 | 1990 | 1780 | 1970 | 1760 | 1560 | 1390 | 826 | 741 | 131 | 127 | 148 | 126 | 141 | 131 | 124 | 115 | 116 | 108 | 1.291 | 1.244 | 899.95 |  |  |  | no fire resistant
|  | 1100 V | 3440 | 3100 | 3350 | 3000 | 3190 | 2820 | 1500 | 1510 | 735 | 720 | 703 | 604 | 247 | 238 | 186 | 172 | 296 | 150 | 157 | 146 | 1.255 | 1.168 | 991.5
CM9694 | yes | 850 V | 738 | 731 | 726 | 717 | 737 | 731 | 680 | 672 | 557 | 549 | 55 | 51.9 | 55.8 | 53.1 | 66.6 | 65 | 68.3 | 70 | 48.6 | 48 | 1.406 | 1.415 | 760.15 | 15.5 inches | yes | Natalie on 11/8/13 |  | taken by Natalie on 11/14/13
|  | 900 V | 1130 | 1110 | 1150 | 1140 | 1140 | 1130 | 1020 | 1000 | 672 | 673 | 85.1 | 85.9 | 90.4 | 88 | 92.7 | 89.3 | 93.9 | 92.7 | 86.2 | 83 | 1.345 | 1.275 | 807
|  | 1000 V | 2170 | 2150 | 2170 | 2140 | 2100 | 2080 | 1540 | 1550 | 812 | 821 | 133 | 134 | 152 | 148 | 137 | 138 | 128 | 122 | 182 | 179 | 1.121 | 1.202 | 899.9
|  | 1100 V | 3470 | 3440 | 3540 | 3500 | 3290 | 3270 | 1560 | 1660 | 678 | 758 | 236 | 233 | 196 | 195 | 190 | 190 | 211 | 195 | 136 | 137 | 1.213 | 1.106 | 991.55
|  |  | 10 kHz | 10 kHz | 1 MHz | 1 MHz | 2 MHz | 2 MHz | 5 MHz | 5 MHz | 10 MHz | 10 MHz | 10 kHz | 10 kHz | 1 MHz | 1 MHz | 2 MHz | 2 MHz | 5 MHz | 5 MHz | 10 MHz | 10 MHz | 5 MHz | 5 MHz | 10 kHz
CM9756 | yes | 850 V | 717 | 723 | 703 | 705 | 711 | 715 | 658 | 657 | 544 | 544 | 51.1 | 53.7 | 58.9 | 62.3 | 52.3 | 53 | 68 | 69.1 | 55.7 | 55.8 | 1.405 | 1.44 | 759.95 | 15.25 inches | yes | Natalie on 11/8/13 |  | taken by Natalie on 11/14/13
|  | 900 V | 1150 | 1150 | 1100 | 1110 | 1080 | 1090 | 992 | 996 | 670 | 669 | 86.4 | 83.8 | 103 | 105 | 88 | 88.5 | 77.5 | 74.7 | 73.9 | 74.2 | 1.38 | 1.364 | 806.85
|  | 1000 V | 2150 | 2150 | 2120 | 2120 | 2110 | 2100 | 1570 | 1570 | 842 | 837 | 163 | 159 | 174 | 169 | 155 | 151 | 154 | 149 | 138 | 134 | 1.218 | 1.117 | 899.75
|  | 1100 V | 3580 | 2570 | 3560 | 3540 | 3280 | 3270 | 1440 | 1640 | 554 | 792 | 195 | 188 | 208 | 202 | 331 | 300 | 166 | 180 | 991 | 121 | 1.142 | 1.224 | 991.25
CM9634 | yes | 850 V | 736 | 723 | 737 | 726 | 744 | 732 | 697 | 692 | 549 | 543 | 56.6 | 56.4 | 64 | 63 | 61.4 | 60.7 | 62.3 | 58.5 | 64.1 | 63.9 | 1.417 | 1.314 | 759.55 | 15.5 inches | yes | Natalie on 11/11/13 |  | taken by Natalie on 11/14/13
|  | 900 V | 1110 | 1090 | 1120 | 1100 | 1110 | 1090 | 1020 | 1010 | 683 | 687 | 89.1 | 87.1 | 104 | 101 | 91.1 | 95.6 | 95.5 | 91.6 | 77.1 | 76.8 | 1.356 | 1.32 | 806.85
|  | 1000 V | 2240 | 2200 | 2220 | 2180 | 2200 | 2160 | 1570 | 1590 | 826 | 833 | 119 | 118 | 141 | 139 | 135 | 128 | 147 | 141 | 156 | 151 | 1.249 | 1.242 | 899.1
|  | 1100 V | 3650 | 3590 | 3680 | 3620 | 3290 | 3260 | 1490 | 1630 | 675 | 704 | 228 | 226 | 217 | 213 | 164 | 156 | 231 | 180 | 174 | 167 | 1.45 | 1.148 | 990.95
CM9832 | yes | 850 V | 477 | 482 | 467 | 469 | 473 | 471 | 440 | 440 | 395 | 398 | 39 | 38.3 | 39.2 | 38.7 | 44.6 | 43.5 | 41.8 | 44.1 | 53 | 52.3 | 1.47 | 1.408 | 726.75 | 15.5 inches | yes | Natalie on 11/11/13 | lower overall???? | taken by Natalie on 11/14/13
| checked twice | 900 V | 759 | 764 | 761 | 765 | 769 | 774 | 730 | 733 | 588 | 589 | 64.1 | 65.3 | 75.2 | 74 | 63.1 | 62.4 | 73.1 | 72.5 | 64.8 | 63.1 | 1.299 | 1.038 | 773.45 |  |  |  | currents are lower
|  | 1000 V | 1620 | 1640 | 1590 | 1610 | 1620 | 1640 | 1380 | 1390 | 833 | 828 | 117 | 116 | 111 | 113 | 110 | 113 | 121 | 128 | 107 | 112 | 1.281 | 1.319 | 865.9 |  |  |  | too
|  | 1100 V | 2860 | 2920 | 2820 | 2870 | 2690 | 2750 | 1700 | 1760 | 767 | 780 | 199 | 205 | 204 | 214 | 164 | 177 | 184 | 192 | 205 | 167 | 1.188 | 1.143 | 958.9
No label-3 | yes | 850 V | 602 | 600 | 595 | 590 | 600 | 604 | 563 | 560 | 496 | 492 | 50.9 | 50.1 | 60.4 | 58.2 | 55.9 | 54.9 | 50.2 | 49.9 | 46 | 45.8 | 1.404 | 1.384 | 761 | 15.5 inches | yes | Natalie on 11/11/13 |  | taken by Natalie on 11/14/13
|  | 900 V | 965 | 957 | 950 | 941 | 910 | 906 | 865 | 859 | 660 | 658 | 83.3 | 87.2 | 70.8 | 70.2 | 67.5 | 65.6 | 83.4 | 80 | 68.5 | 67.4 | 1.381 | 1.325 | 807.8
|  | 1000 V | 1850 | 1820 | 1840 | 1820 | 1810 | 1790 | 1530 | 1520 | 781 | 791 | 124 | 127 | 135 | 132 | 133 | 312 | 129 | 130 | 153 | 149 | 1.294 | 1.175 | 900.75
|  | 1100 V | 3160 | 3110 | 3150 | 3100 | 3000 | 2960 | 1660 | 1740 | 771 | 865 | 167 | 163 | 224 | 227 | 200 | 201 | 242 | 224 | 226 | 198 | 1.037 | 1.195 | 989.05
|  |  | 10 kHz | 10 kHz | 1 MHz | 1 MHz | 2 MHz | 2 MHz | 5 MHz | 5 MHz | 10 MHz | 10 MHz | 10 kHz | 10 kHz | 1 MHz | 1 MHz | 2 MHz | 2 MHz | 5 MHz | 5 MHz | 10 MHz | 10 MHz | 5 MHz | 5 MHz | 10 kHz
No label-4 | BAD | 850 V | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # |  |  | Natalie on 11/12/13 | BAD, put in box | BAD
|  | 900 V | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # |  |  |  | for Vladimir to look
|  | 1000 V | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # |  |  |  | at later
|  | 1100 V | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | #
CM9759 | BAD | 850 V | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # |  |  | Natalie on 11/12/13 | BAD, put in box | BAD
|  | 900 V | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # |  |  |  | for Vladimir to look
|  | 1000 V | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # |  |  |  | at later
|  | 1100 V | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | #
No label-5 | yes | 850 V | 667 | 671 | 642 | 640 | 668 | 669 | 628 | 629 | 505 | 508 | 50.2 | 50 | 51.5 | 52.5 | 77.7 | 75.7 | 55 | 54.8 | 58 | 58.1 | 1.303 | 1.427 | 760.15 | 15.5 inches | yes | Natalie on 11/12/13 |  | taken by Natalie on 11/14/13
|  | 900 V | 1020 | 1020 | 981 | 974 | 962 | 959 | 903 | 899 | 657 | 661 | 95.9 | 92.5 | 96 | 92.2 | 86.1 | 85 | 82.6 | 79.4 | 99.7 | 97.1 | 1.349 | 1.339 | 807.05
|  | 1000 V | 1980 | 1970 | 2020 | 2000 | 1970 | 1960 | 1550 | 1550 | 834 | 845 | 131 | 129 | 136 | 132 | 144 | 138 | 149 | 149 | 165 | 167 | 1.16 | 1.164 | 899.75
|  | 1100 V | 3340 | 3310 | 3370 | 3350 | 3140 | 3130 | 1490 | 1690 | 746 | 807 | 207 | 200 | 223 | 225 | 190 | 193 | 318 | 195 | 147 | 152 | 1.138 | 1.193 | 991.45
No label-6 | yes | 850 V | 664 | 662 | 654 | 652 | 678 | 673 | 646 | 643 | 525 | 522 | 58.8 | 57.6 | 56.7 | 55.2 | 72.9 | 71.8 | 52.9 | 54.2 | 55.9 | 58.4 | 1.308 | 1.385 | 760.6 | 15.5 inches | yes | Natalie on 11/12/13 |  | taken by Natalie on 11/14/13
|  | 900 V | 1040 | 1040 | 1030 | 1030 | 1020 | 1020 | 948 | 946 | 720 | 718 | 75.8 | 84.4 | 80 | 77.3 | 84.6 | 81.8 | 81.7 | 82.4 | 87.1 | 84.6 | 1.406 | 1.409 | 807.6
|  | 1000 V | 2000 | 1980 | 2000 | 1990 | 1950 | 1920 | 1580 | 1570 | 820 | 827 | 126 | 123 | 134 | 124 | 110 | 113 | 149 | 145 | 120 | 128 | 1.283 | 1.253 | 900.4
|  | 1100 V | 3340 | 3320 | 3310 | 3290 | 3050 | 3020 | 1670 | 1740 | 778 | 751 | 206 | 205 | 192 | 194 | 178 | 176 | 298 | 226 | 170 | 178 | 1.123 | 1.167 | 991.85
CM9626 | yes | 850 V | 716 | 718 | 680 | 683 | 735 | 737 | 674 | 681 | 555 | 559 | 58.5 | 58.1 | 72.2 | 70.2 | 77.3 | 79.5 | 60.1 | 60.2 | 64.6 | 63.8 | 1.307 | 1.407 | 759.95 | 15.5 inches | yes | Natalie on 11/12/13 | pic named 9686 | taken by Natalie on 11/14/13
|  | 900 V | 1110 | 1120 | 1100 | 1100 | 1080 | 1080 | 1020 | 1030 | 717 | 720 | 84 | 87.6 | 101 | 101 | 89 | 88.1 | 92.3 | 90.5 | 93.6 | 92 | 1.36 | 1.355 | 806.8 |  |  |  | by accident
|  | 1000 V | 2170 | 2180 | 2150 | 2150 | 2100 | 2100 | 1610 | 1620 | 832 | 830 | 134 | 132 | 156 | 159 | 148 | 146 | 168 | 174 | 142 | 145 | 1.242 | 1.24 | 899.3
|  | 1100 V | 3590 | 3610 | 3470 | 3480 | 3320 | 3330 | 1630 | 1680 | 692 | 699 | 212 | 210 | 209 | 205 | 187 | 182 | 236 | 197 | 147 | 148 | 1.214 | 1.169 | 991.1
CM9565 | yes | 850 V | 781 | 771 | 773 | 767 | 778 | 771 | 710 | 706 | 582 | 575 | 67.6 | 68.2 | 66.8 | 64 | 71.5 | 70.4 | 73.9 | 70.1 | 74.7 | 70.1 | 1.422 | 1.398 | 758.9 | 15.5 inches | yes | Natalie on 11/12/13 | 1Mhz,1100V | taken by Natalie on 11/14/13
|  | 900 V | 1210 | 1200 | 1190 | 1180 | 1180 | 1170 | 1070 | 1060 | 762 | 760 | 88.8 | 90.6 | 81.6 | 77 | 107 | 105 | 83.5 | 81.9 | 154 | 150 | 1.323 | 1.387 | 806.35 |  |  |  | standard jumped entire time
|  | 1000 V | 2330 | 2280 | 2270 | 2230 | 2220 | 2180 | 1610 | 1610 | 794 | 829 | 140 | 131 | 170 | 162 | 160 | 151 | 206 | 198 | 166 | 158 | 1.2 | 1.266 | 898.55
|  | 1100 V | 3860 | 3780 | 3800 | 3720 | 3490 | 3390 | 895 | 1410 | 474 | 863 | 219 | 217 | 395 | 380 | 201 | 198 | 170 | 155 | 110 | 155 | 1.574 | 1.187 | 990.2
|  |  | 10 kHz | 10 kHz | 1 MHz | 1 MHz | 2 MHz | 2 MHz | 5 MHz | 5 MHz | 10 MHz | 10 MHz | 10 kHz | 10 kHz | 1 MHz | 1 MHz | 2 MHz | 2 MHz | 5 MHz | 5 MHz | 10 MHz | 10 MHz | 5 MHz | 5 MHz | 10 kHz
CM9639 | yes; power setting: 10.90 | 850 V | 642 | 643 | 643 | 643 | 645 | 646 | 642 | 640 | 637 | 633 | 21.2 | 20.8 | 22.3 | 21.4 | 23.8 | 22.5 | 19.2 | 20.1 | 15.5 | 14.6 | 1.469 | 1.285 | 759.7 | 15.5 inches | yes | Nathan on 12/5/13 | retest; no fire-resistant wrap | taken by Franz on 12/11/13
|  | 900 V | 804 | 802 | 813 | 815 | 803 | 799 | 761 | 756 | 680 | 670 | 71.2 | 70.3 | 66.3 | 63.5 | 68.2 | 69.4 | 67.7 | 65.6 | 45.3 | 44.2 | 1.329 | 1.292 | 808.85
|  | 1000 V | 1670 | 1670 | 1700 | 1690 | 1660 | 1650 | 1470 | 1450 | 788 | 770 | 129 | 122 | 135 | 137 | 160 | 161 | 139 | 146 | 91.8 | 90.2 | 1.321 | 1.296 | 899.3
|  | 1100 V | 3000 | 3000 | 2930 | 2930 | 2820 | 2830 | 1540 | 1530 | 870 | 811 | 215 | 214 | 202 | 201 | 173 | 175 | 220 | 216 | 123 | 151 | 1.21 | 1.153 | 991.3
CM9594 | power setting (10.9) fixed by last consistency test | 850 V | 648 | 637 | 645 | 636 | 643 | 637 | 650 | 640 | 640 | 634 | 25.1 | 26.3 | 22.1 | 21.9 | 29.3 | 28.4 | 24.3 | 23.2 | 17.1 | 16.7 | 1.36 | 1.471 | 759.2 | 15.5 inches | yes | Nathan on 12/6/13 | retest | taken by Franz on 12/11/13
|  | 900 V | 850 | 844 | 855 | 845 | 847 | 833 | 800 | 793 | 680 | 670 | 71.8 | 68.1 | 72.2 | 69.5 | 76.3 | 74.4 | 70.1 | 67.2 | 51.1 | 48.2 | 1.36 | 1.288 | 805.65
|  | 1000 V | 1750 | 1720 | 1780 | 1730 | 1750 | 1710 | 1420 | 1420 | 814 | 817 | 154 | 148 | 161 | 155 | 144 | 141 | 143 | 142 | 86.2 | 88.1 | 1.207 | 1.323 | 898.45
|  | 1100 V | 3120 | 3070 | 3030 | 2990 | 2900 | 2860 | 1590 | 1630 | 803 | 842 | 233 | 231 | 259 | 255 | 203 | 206 | 208 | 216 | 134 | 113 | 1.153 | 1.18 | 990.45
CM9592 | power setting (10.9) fixed by last consistency test | 850 V | 471 | 469 | 482 | 486 | 472 | 473 | 463 | 465 | 406 | 405 | 43.6 | 44.9 | 53.2 | 53.7 | 56.2 | 55.5 | 41.3 | 42.5 | 44.6 | 45.3 | 1.362 | 1.379 | 759.65 | 15.5 inches | yes | Nathan on 12/6/13 | retest | taken by Franz on 12/11/13
|  | 900 V | 756 | 757 | 763 | 770 | 757 | 762 | 714 | 725 | 574 | 576 | 70.1 | 70.9 | 73.3 | 75.1 | 73.3 | 74.5 | 75.5 | 74.3 | 71.7 | 71.2 | 1.382 | 1.273 | 805.9
|  | 1000 V | 1560 | 1570 | 1550 | 1560 | 1520 | 1550 | 1350 | 1360 | 757 | 755 | 137 | 133 | 127 | 129 | 118 | 120 | 131 | 134 | 138 | 141 | 1.309 | 1.288 | 898.9
|  | 1100 V | 2700 | 2720 | 2740 | 2760 | 2700 | 2710 | 1690 | 1660 | 835 | 783 | 182 | 187 | 183 | 189 | 215 | 220 | 194 | 197 | 186 | 192 | 1.135 | 1.031 | 990.9
CM9593 | power setting (10.9) fixed by last consistency test | 850 V | 503 | 501 | 509 | 507 | 504 | 500 | 499 | 497 | 426 | 424 | 51.7 | 50.4 | 44.1 | 43.8 | 52.1 | 52.3 | 53.2 | 52.9 | 56.1 | 55.3 | 1.425 | 1.392 | 758.95 | 15.5 inches | yes | Nathan on 12/6/13 | retest | taken by Franz on 12/11/13
|  | 900 V | 785 | 781 | 804 | 803 | 798 | 797 | 780 | 781 | 583 | 585 | 74.6 | 74 | 70.5 | 71.8 | 65.8 | 64.5 | 74.2 | 73.1 | 67.7 | 66 | 1.382 | 1.261 | 805.15
|  | 1000 V | 1630 | 1640 | 1620 | 1610 | 1590 | 1580 | 1350 | 1340 | 785 | 773 | 120 | 121 | 133 | 131 | 144 | 139 | 158 | 155 | 144 | 141 | 1.314 | 1.2 | 898.05
|  | 1100 V | 2800 | 2780 | 2860 | 2850 | 2600 | 2610 | 1640 | 1610 | 857 | 833 | 212 | 214 | 200 | 204 | 197 | 191 | 185 | 183 | 196 | 203 | 1.156 | 1.143 | 990.2
CM9571 | power setting (10.9) fixed by last consistency test | 850 V | 542 | 538 | 530 | 525 | 521 | 512 | 495 | 496 | 436 | 428 | 52.4 | 51.5 | 42.3 | 44.1 | 45 | 44.7 | 51.8 | 50.5 | 45.4 | 45.5 | 1.44 | 1.487 | 759.6 | 15.5 inches | yes | Nathan on 12/6/13 | retest | taken by Franz on 12/11/13
|  | 900 V | 817 | 811 | 815 | 808 | 804 | 797 | 766 | 755 | 586 | 579 | 69.9 | 70.4 | 71.8 | 71.3 | 76.1 | 77.2 | 83.1 | 82.1 | 72.2 | 71.3 | 1.352 | 1.267 | 805.55
|  | 1000 V | 1680 | 1650 | 1630 | 1610 | 1640 | 1610 | 1430 | 1410 | 788 | 768 | 139 | 141 | 128 | 130 | 132 | 127 | 147 | 144 | 102 | 99.8 | 1.293 | 1.269 | 898.4
|  | 1100 V | 2920 | 2860 | 2850 | 2790 | 2710 | 2660 | 1670 | 1570 | 833 | 759 | 175 | 178 | 203 | 207 | 195 | 189 | 176 | 171 | 173 | 182 | 1.169 | 1.188 | 990.65
CM9386 | power setting (10.9) fixed by last consistency test | 850 V | 490 | 487 | 486 | 485 | 500 | 502 | 493 | 487 | 424 | 421 | 49.6 | 50.1 | 57.4 | 56.4 | 44.1 | 45.2 | 50.1 | 49.5 | 48.3 | 49.8 | 1.428 | 1.405 | 759.7 | 15.5 inches | yes | Nathan on 12/6/13 |  | taken by Franz on 12/11/13
|  | 900 V | 742 | 739 | 760 | 755 | 756 | 757 | 723 | 721 | 563 | 561 | 60 | 58.7 | 67.9 | 68 | 68.9 | 69.7 | 67.8 | 68.4 | 79.9 | 77.8 | 1.305 | 1.316 | 806.1
|  | 1000 V | 1520 | 1530 | 1520 | 1520 | 1530 | 1540 | 1360 | 1350 | 777 | 771 | 132 | 135 | 124 | 123 | 118 | 123 | 156 | 158 | 120 | 116 | 1.262 | 1.279 | 899.1
|  | 1100 V | 2710 | 2700 | 2720 | 2730 | 2660 | 2650 | 1630 | 1600 | 789 | 743 | 186 | 193 | 180 | 185 | 209 | 206 | 198 | 196 | 196 | 203 | 1.144 | 1.156 | 991.3
|  |  | 10 kHz | 10 kHz | 1 MHz | 1 MHz | 2 MHz | 2 MHz | 5 MHz | 5 MHz | 10 MHz | 10 MHz | 10 kHz | 10 kHz | 1 MHz | 1 MHz | 2 MHz | 2 MHz | 5 MHz | 5 MHz | 10 MHz | 10 MHz | 5 MHz | 5 MHz | 10 kHz
CM9374 | power setting (10.9) fixed by last consistency test | 850 V | 508 | 510 | 494 | 499 | 483 | 486 | 471 | 475 | 415 | 417 | 51.9 | 52 | 46.1 | 45.2 | 55 | 55.4 | 50.2 | 50.9 | 46.5 | 47.1 | 1.404 | 1.392 | 758.7 | 15.5 inches | yes | Nathan on 12/6/13 |  | taken by Franz on 12/11/13
|  | 900 V | 763 | 773 | 748 | 752 | 743 | 747 | 706 | 715 | 566 | 572 | 65.2 | 66.1 | 75.5 | 77.2 | 69.2 | 71.2 | 73.6 | 72.7 | 63.4 | 63.6 | 1.342 | 1.331 | 804.95
|  | 1000 V | 1560 | 1580 | 1520 | 1530 | 1520 | 1530 | 1340 | 1340 | 777 | 762 | 136 | 141 | 142 | 148 | 122 | 129 | 130 | 129 | 148 | 142 | 1.252 | 1.192 | 898.05
|  | 1100 V | 2670 | 2700 | 2620 | 2640 | 2550 | 2580 | 1600 | 1640 | 787 | 820 | 186 | 182 | 203 | 206 | 187 | 184 | 183 | 191 | 187 | 177 | 1.103 | 1.167 | 989.85
CN7558 | power setting (10.9) fixed by last consistency test | 850 V | 410 | 417 | 399 | 400 | 385 | 394 | 378 | 387 | 339 | 346 | 36.8 | 36.1 | 44.2 | 43 | 41 | 39.1 | 38.2 | 40.1 | 26 | 29.4 | 1.403 | 1.376 | 758.05 | 15.5 inches | yes | Nathan on 12/6/13 |  | taken by Franz on 12/11/13
|  | 900 V | 587 | 598 | 590 | 589 | 586 | 595 | 578 | 589 | 487 | 487 | 62.1 | 60.4 | 51.3 | 52.2 | 54.1 | 53.6 | 54.7 | 53.1 | 66.5 | 64.6 | 1.409 | 1.347 | 804.5
|  | 1000 V | 1230 | 1250 | 1220 | 1230 | 1220 | 1230 | 1110 | 1120 | 733 | 761 | 110 | 114 | 100 | 99.1 | 97 | 96.7 | 120 | 124 | 114 | 116 | 1.372 | 1.342 | 897.45
|  | 1100 V | 2170 | 2210 | 2180 | 2200 | 2090 | 2120 | 1590 | 1640 | 778 | 812 | 171 | 178 | 170 | 167 | 166 | 170 | 156 | 152 | 151 | 159 | 1.194 | 1.171 | 989.55
CM9361 | power setting (10.9) fixed by last consistency test | 850 V | 496 | 493 | 492 | 488 | 493 | 490 | 478 | 475 | 418 | 414 | 50.1 | 50 | 49.2 | 49 | 49 | 49 | 54.5 | 53.9 | 49 | 48.7 | 1.416 | 1.425 | 760.4 | 15.5 inches | yes | Nathan on 12/9/13 |  | taken by Franz on 12/11/13
|  | 900 V | 757 | 752 | 745 | 740 | 744 | 739 | 720 | 715 | 586 | 581 | 72.2 | 72.4 | 73.3 | 72.4 | 74.4 | 73.6 | 75 | 74.5 | 83.1 | 82.7 | 1.377 | 1.366 | 806.95
|  | 1000 V | 1540 | 1530 | 1560 | 1550 | 1550 | 1540 | 1330 | 1320 | 720 | 715 | 133 | 135 | 131 | 133 | 134 | 135 | 165 | 163 | 145 | 144 | 1.333 | 1.266 | 899.5
|  | 1100 V | 2610 | 2600 | 2600 | 2590 | 2590 | 2580 | 1660 | 1650 | 812 | 746 | 210 | 210 | 203 | 205 | 197 | 197 | 205 | 204 | 182 | 204 | 1.174 | 1.137 | 992.05
CM9373 | power setting (10.9) fixed by last consistency test | 850 V | 475 | 473 | 483 | 481 | 477 | 475 | 455 | 452 | 402 | 399 | 47.3 | 46.8 | 48.3 | 49.1 | 49.2 | 49.3 | 49.2 | 48.5 | 49.4 | 49.9 | 1.42 | 1.416 | 759.95 | 15.5 inches | yes | Nathan on 12/9/13 |  | taken by Franz on 12/11/13
|  | 900 V | 718 | 714 | 723 | 720 | 725 | 722 | 683 | 680 | 559 | 558 | 67.1 | 67.5 | 64.1 | 64.2 | 67.1 | 67.3 | 69.5 | 70.2 | 70.1 | 69.3 | 1.335 | 1.366 | 805.95
|  | 1000 V | 1450 | 1440 | 1460 | 1460 | 1440 | 1440 | 1290 | 1290 | 750 | 755 | 115 | 114 | 126 | 126 | 125 | 124 | 119 | 120 | 126 | 123 | 1.312 | 1.3 | 898.15
|  | 1100 V | 2490 | 2470 | 2510 | 2490 | 2440 | 2420 | 1580 | 1600 | 800 | 824 | 171 | 171 | 183 | 181 | 174 | 171 | 159 | 162 | 161 | 164 | 1.085 | 1.169 | 990.85
CM9753 | power setting (10.9) fixed by last consistency test | 850 V | 563 | 566 | 567 | 570 | 555 | 559 | 525 | 530 | 445 | 449 | 52.9 | 53.7 | 52.1 | 52.4 | 52.7 | 52.4 | 51.5 | 52.3 | 54.1 | 53.7 | 1.392 | 1.433 | 759.1 | 15.5 inches | yes | Nathan on 12/9/13 |  | taken by Franz on 12/11/13
|  | 900 V | 833 | 838 | 828 | 831 | 828 | 832 | 781 | 786 | 597 | 599 | 77.8 | 76.3 | 75 | 74.6 | 75.1 | 75.7 | 73.5 | 73.1 | 71.4 | 71.7 | 1.368 | 1.342 | 805.2
|  | 1000 V | 1710 | 1720 | 1680 | 1690 | 1680 | 1690 | 1390 | 1400 | 762 | 750 | 131 | 130 | 136 | 137 | 133 | 132 | 125 | 127 | 136 | 137 | 1.275 | 1.257 | 898.2
|  | 1100 V | 2930 | 2940 | 2940 | 2960 | 2800 | 2820 | 1620 | 1590 | 766 | 704 | 204 | 208 | 207 | 210 | 203 | 205 | 209 | 207 | 176 | 195 | 1.1697 | 1.134 | 990.2
CM9752 | power setting (10.9) fixed by last consistency test | 850 V | 619 | 614 | 592 | 587 | 580 | 577 | 560 | 556 | 465 | 463 | 53.2 | 52.9 | 54.5 | 53.2 | 55.1 | 54.3 | 56.5 | 56.3 | 53.3 | 53.5 | 1.428 | 1.375 | 760.15 | 15.5 inches | yes | Nathan on 12/9/13 | Output 1 signal is less stable at 1100V, 5 MHz | taken by Franz on 12/11/13
|  | 900 V | 894 | 883 | 890 | 880 | 897 | 886 | 836 | 826 | 611 | 609 | 75.3 | 76.1 | 73.5 | 74.3 | 76.2 | 77.1 | 76.2 | 74.9 | 71.2 | 69.4 | 1.389 | 1.355 | 806.25
|  | 1000 V | 1850 | 1820 | 1810 | 1780 | 1790 | 1750 | 1410 | 1400 | 758 | 767 | 144 | 142 | 142 | 140 | 146 | 142 | 147 | 151 | 137 | 135 | 1.243 | 1.232 | 899.2
|  | 1100 V | 3120 | 3050 | 3050 | 2990 | 2910 | 2870 | 1400 | 1560 | 711 | 724 | 210 | 213 | 206 | 205 | 197 | 202 | 278 | 188 | 170 | 169 | 1.158 | 1.105 | 991
|  |  | 10 kHz | 10 kHz | 1 MHz | 1 MHz | 2 MHz | 2 MHz | 5 MHz | 5 MHz | 10 MHz | 10 MHz | 10 kHz | 10 kHz | 1 MHz | 1 MHz | 2 MHz | 2 MHz | 5 MHz | 5 MHz | 10 MHz | 10 MHz | 5 MHz | 5 MHz | 10 kHz
CM9384 | power setting (10.9) fixed by last consistency test | 850 V | 466 | 462 | 465 | 461 | 470 | 465 | 444 | 441 | 397 | 392 | 51.3 | 49.3 | 48.4 | 47.7 | 50.3 | 50.4 | 51.8 | 50.3 | 46.1 | 47 | 1.375 | 1.433 | 759.5 | 15.5 inches | yes | Nathan on 12/9/13 |  | taken by Franz on 12/11/13
|  | 900 V | 725 | 715 | 719 | 707 | 720 | 711 | 690 | 684 | 555 | 549 | 69.4 | 68.8 | 70.8 | 69.4 | 63.5 | 64.4 | 64.6 | 65.2 | 66.1 | 65.3 | 1.356 | 1.347 | 805.7
|  | 1000 V | 1470 | 1440 | 1460 | 1430 | 1500 | 1480 | 1320 | 1310 | 760 | 754 | 123 | 121 | 126 | 128 | 123 | 126 | 129 | 125 | 128 | 125 | 1.261 | 1.242 | 898.75
|  | 1100 V | 2590 | 2530 | 2600 | 2550 | 2570 | 2520 | 1640 | 1580 | 809 | 741 | 202 | 199 | 201 | 198 | 199 | 196 | 182 | 180 | 149 | 163 | 1.057 | 1.123 | 990.7
CM9382 | power setting (10.9) fixed by last consistency test | 850 V | 473 | 469 | 500 | 496 | 496 | 491 | 490 | 486 | 413 | 410 | 49.5 | 48.3 | 48.6 | 49.1 | 48.4 | 48.8 | 48.1 | 47.5 | 47.3 | 46.5 | 1.432 | 1.4 | 760.85 | 15.5 inches | yes | Nathan on 12/10/13 |  | taken by Franz on 12/11/13
|  | 900 V | 766 | 758 | 755 | 750 | 765 | 759 | 740 | 737 | 571 | 570 | 70.6 | 69.5 | 70.3 | 69.8 | 70.1 | 70.7 | 71.5 | 70.4 | 70.4 | 70.3 | 1.349 | 1.372 | 806.75
|  | 1000 V | 1560 | 1550 | 1560 | 1550 | 1540 | 1530 | 1340 | 1340 | 743 | 755 | 126 | 124 | 132 | 135 | 125 | 124 | 115 | 114 | 127 | 126 | 1.281 | 1.276 | 899.6
|  | 1100 V | 2720 | 2710 | 2730 | 2700 | 2660 | 2640 | 1580 | 1660 | 705 | 790 | 185 | 182 | 182 | 179 | 177 | 180 | 183 | 184 | 193 | 172 | 1.146 | 1.118 | 991.65
CM9378 | power setting (10.9) fixed by last consistency test | 850 V | 460 | 461 | 465 | 463 | 459 | 457 | 438 | 437 | 392 | 393 | 45.1 | 44.4 | 47.6 | 46.8 | 47.7 | 47.3 | 44.1 | 44 | 46.1 | 45.5 | 1.477 | 1.442 | 760.15 | 15.5 inches | yes | Nathan on 12/10/13 |  | taken by Franz on 12/11/13
|  | 900 V | 677 | 675 | 695 | 693 | 702 | 701 | 673 | 672 | 553 | 554 | 64.9 | 65.1 | 63.8 | 63.7 | 65.2 | 65 | 60.2 | 59.8 | 60.6 | 60.3 | 1.395 | 1.388 | 806.3
|  | 1000 V | 1410 | 1400 | 1430 | 1420 | 1420 | 1410 | 1280 | 1270 | 762 | 753 | 120 | 118 | 120 | 121 | 118 | 116 | 117 | 116 | 115 | 114 | 1.283 | 1.311 | 899.1
|  | 1100 V | 2450 | 2440 | 2530 | 2510 | 2460 | 2440 | 1660 | 1610 | 813 | 726 | 177 | 176 | 180 | 182 | 183 | 186 | 185 | 187 | 172 | 195 | 1.173 | 1.146 | 991.15
CM9879 | power setting (10.9) fixed by last consistency test | 850 V | 386 | 383 | 386 | 385 | 372 | 373 | 355 | 356 | 320 | 321 | 39.5 | 38.9 | 40.3 | 39.7 | 35 | 34.3 | 34.7 | 34.2 | 35.8 | 35.1 | 1.4467 | 1.508 | 760.2 | 15.5 inches | yes | Nathan on 12/10/13 |  | taken by Franz on 12/11/13
|  | 900 V | 550 | 552 | 541 | 542 | 551 | 552 | 536 | 537 | 458 | 457 | 50.7 | 49.9 | 45.6 | 46.7 | 49.5 | 50.1 | 56.1 | 57.3 | 50.5 | 50.7 | 1.383 | 1.414 | 806.25
|  | 1000 V | 1140 | 1140 | 1130 | 1130 | 1130 | 1130 | 1020 | 1020 | 720 | 721 | 110 | 108 | 108 | 109 | 105 | 104 | 103 | 104 | 101 | 101 | 1.303 | 1.351 | 899.1
|  | 1100 V | 1980 | 1970 | 1960 | 1960 | 1950 | 1940 | 1590 | 1610 | 808 | 820 | 150 | 151 | 152 | 149 | 156 | 155 | 150 | 149 | 133 | 130 | 1.165 | 1.188 | 991.05
No label-7 | power setting (10.9) fixed by last consistency test | 850 V | 495 | 490 | 486 | 482 | 473 | 470 | 455 | 450 | 388 | 386 | 50.1 | 49.7 | 53.3 | 52.9 | 52.3 | 52.5 | 49.5 | 48.2 | 46.2 | 46.2 | 1.4 | 1.425 | 759.6 | 15.5 inches | yes | Nathan on 12/10/13 |  | taken by Franz on 12/11/13
|  | 900 V | 722 | 716 | 712 | 705 | 720 | 714 | 685 | 680 | 550 | 549 | 64.3 | 65.5 | 69 | 69.8 | 67.2 | 67.9 | 70.7 | 69.9 | 70 | 69.3 | 1.37 | 1.406 | 805.9
|  | 1000 V | 1510 | 1490 | 1470 | 1450 | 1480 | 1460 | 1280 | 1270 | 765 | 764 | 121 | 120 | 127 | 126 | 123 | 125 | 120 | 118 | 122 | 121 | 1.332 | 1.25 | 898.75
|  | 1100 V | 2610 | 2580 | 2590 | 2560 | 2510 | 2480 | 1640 | 1650 | 844 | 842 | 172 | 173 | 171 | 172 | 171 | 168 | 173 | 171 | 172 | 174 | 1.123 | 1.15 | 991
No label-8 | power setting (10.9) fixed by last consistency test | 850 V | 502 | 501 | 494 | 494 | 491 | 491 | 463 | 462 | 410 | 411 | 45 | 44.3 | 47.2 | 47.1 | 45.2 | 44.3 | 46.6 | 45.9 | 47 | 47.3 | 1.443 | 1.38 | 760.6 | 15.5 inches | yes | Nathan on 12/10/13 |  | taken by Franz on 12/11/13
|  | 900 V | 737 | 737 | 733 | 731 | 740 | 740 | 726 | 723 | 558 | 556 | 67.1 | 66.5 | 63.4 | 64.5 | 69.1 | 69 | 70.7 | 71.2 | 68.9 | 69.2 | 1.269 | 1.305 | 806.9
|  | 1000 V | 1520 | 1520 | 1510 | 1510 | 1520 | 1520 | 1340 | 1330 | 771 | 755 | 131 | 127 | 131 | 130 | 128 | 130 | 128 | 130 | 123 | 122 | 1.212 | 1.225 | 899.9
|  | 1100 V | 2600 | 2620 | 2660 | 2660 | 2590 | 2580 | 1680 | 1640 | 822 | 743 | 187 | 184 | 180 | 181 | 188 | 191 | 183 | 185 | 178 | 190 | 1.186 | 1.166 | 991.85
|  |  | 10 kHz | 10 kHz | 1 MHz | 1 MHz | 2 MHz | 2 MHz | 5 MHz | 5 MHz | 10 MHz | 10 MHz | 10 kHz | 10 kHz | 1 MHz | 1 MHz | 2 MHz | 2 MHz | 5 MHz | 5 MHz | 10 MHz | 10 MHz | 5 MHz | 5 MHz | 10 kHz
No label-9 | power setting (10.9) fixed by last consistency test | 850 V | 477 | 475 | 490 | 486 | 480 | 477 | 454 | 452 | 400 | 399 | 49.1 | 49.3 | 48.9 | 49.3 | 48.1 | 48.5 | 50.4 | 50.1 | 47.3 | 48.1 | 1.445 | 1.432 | 760.1 | 15.5 inches | yes | Nathan on 12/10/13 |  | taken by Franz on 12/11/13
|  | 900 V | 724 | 720 | 732 | 728 | 741 | 738 | 698 | 694 | 553 | 552 | 69.1 | 68.5 | 69.5 | 68.7 | 69.7 | 70 | 71.2 | 72.1 | 71.2 | 71.7 | 1.35 | 1.361 | 806.3
|  | 1000 V | 1540 | 1520 | 1510 | 1500 | 1510 | 1490 | 1320 | 1320 | 756 | 751 | 129 | 130 | 128 | 127 | 125 | 126 | 126 | 125 | 115 | 116 | 1.308 | 1.272 | 899.15
|  | 1100 V | 2680 | 2660 | 2660 | 2640 | 2600 | 2580 | 1620 | 1640 | 707 | 764 | 185 | 183 | 182 | 181 | 183 | 185 | 173 | 176 | 184 | 173 | 1.121 | 1.175 | 991.1
CM9922 | 13.0 power | 850 V | 392 | 412 | 400 | 436 | 424 | 412 | 400 | 388 | 360 | 380 | 36 | 36.4 | 35.2 | 35 | 38.6 | 37.6 | 34.3 | 34.5 | 40.2 | 40.5 | 1.52 | 1.635 | 760.8 | 15.5 inches | yes | Natalie on 12/17/13 |  | taken by Franz on 12/19/13
|  | 900 V | 608 | 612 | 672 | 692 | 624 | 684 | 640 | 580 | 616 | 516 | 50.4 | 50.8 | 67 | 66.7 | 58.1 | 58.7 | 53.8 | 54.1 | 57.9 | 54.4 | 1.507 | 1.439 | 807.3
|  | 1000 V | 1240 | 1230 | 1220 | 1250 | 1220 | 1210 | 1180 | 1030 | 860 | 810 | 101 | 105 | 97.3 | 94 | 104 | 102 | 96.6 | 3 | 99.7 | 97.1 | 1.305 | 1.32 | 900.25
|  | 1100 V | 2200 | 2450 | 2040 | 2130 | 2320 | 2150 | 1560 | 1690 | 960 | 790 | 145 | 141 | 170 | 166 | 155 | 152 | 155 | 152 | 160 | 143 | 1.216 | 1.246 | 992.65
CM9910 | yes | 850 V | 512 | 476 | 440 | 460 | 352 | 452 | 288 | 412 | 256 | 380 | 35.6 | 40.4 | 47 | 42.8 | 32.2 | 40.5 | 34.2 | 44.3 | 36 | 36.2 | 1.197 | 1.045 | 759.8 | 15.5 inches | yes | Natalie on 12/17/13 | stopped testing because output 1 went CRAZY | BAD
|  | 900 V | 744 | 764 | 648 | 718 | 648 | 660 | 704 | 668 | 584 | 684 | 62.3 | 63.3 | 65.4 | 66 | 52.3 | 58.2 | 53.8 | 56.7 | 58 | 62.5 | 1.441 | 1.446 | 806.75 |  |  |  | see photo taken
|  | 1000 V | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # |  |  |  | changing frequency did not help
|  | 1100 V | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # |  |  |  | box for Vladimir
CM9920 | 13.4 power | 850 V | 496 | 476 | 488 | 452 | 416 | 478 | 424 | 452 | 376 | 372 | 40.1 | 37.5 | 44.6 | 40.8 | 38.7 | 36.5 | 38.8 | 37.7 | 43 | 42.8 | 1.47 | 1.366 | 759.35 | 15.5 inches | yes | Natalie on 12/18/13 |  | taken by Franz on 12/19/13
|  | 900 V | 648 | 668 | 600 | 604 | 736 | 732 | 600 | 644 | 576 | 516 | 60.2 | 61.8 | 58.2 | 58.9 | 57 | 61.1 | 61.4 | 60.9 | 57.8 | 55.9 | 1.381 | 1.391 | 806.05
|  | 1000 V | 1260 | 1430 | 1410 | 1440 | 1560 | 1370 | 1240 | 1290 | 790 | 720 | 111 | 113 | 105 | 109 | 113 | 113 | 126 | 127 | 109 | 111 | 1.303 | 1.241 | 899.05
|  | 1100 V | 2580 | 2270 | 2380 | 2420 | 2380 | 2570 | 1520 | 1550 | 680 | 870 | 171 | 170 | 159 | 158 | 163 | 161 | 160 | 159 | 143 | 144 | 1.173 | 1.205 | 991.25
CM9948 | yes | 850 V | 480 | 532 | 560 | 540 | 496 | 516 | 520 | 484 | 488 | 440 | 55.9 | 58.8 | 49.5 | 50.4 | 44.5 | 44.1 | 40.7 | 39.3 | 45 | 46 | 1.51 | 1.478 | 760.45 | 15.5 inches | yes | Natalie on 12/18/13 |  | taken by Franz on 12/19/13
|  | 900 V | 864 | 852 | 760 | 820 | 800 | 812 | 674 | 724 | 592 | 564 | 75.4 | 74.8 | 66 | 67.8 | 70.5 | 72 | 65.8 | 67.6 | 66.7 | 68.2 | 1.45 | 1.441 | 807.2
|  | 1000 V | 1440 | 1550 | 1700 | 1570 | 1560 | 1700 | 1420 | 1270 | 780 | 810 | 108 | 115 | 114 | 117 | 107 | 113 | 115 | 118 | 120 | 121 | 1.204 | 1.313 | 900.25
|  | 1100 V | 2560 | 2610 | 2500 | 2550 | 2660 | 2330 | 1720 | 1530 | 600 | 670 | 170 | 175 | 177 | 181 | 186 | 185 | 167 | 174 | 148 | 152 | 1.144 | 1.172 | 992.6
|  |  | 10 kHz | 10 kHz | 1 MHz | 1 MHz | 2 MHz | 2 MHz | 5 MHz | 5 MHz | 10 MHz | 10 MHz | 10 kHz | 10 kHz | 1 MHz | 1 MHz | 2 MHz | 2 MHz | 5 MHz | 5 MHz | 10 MHz | 10 MHz | 5 MHz | 5 MHz | 10 kHz
CM9950 | yes | 850 V | 536 | 524 | 536 | 484 | 448 | 436 | 424 | 412 | 384 | 428 | 64.8 | 62.4 | 47.6 | 47.2 | 41.4 | 39.1 | 43.5 | 42.5 | 41.5 | 46.4 | 1.481 | 1.51 | 760.65 | 15.5 inches | yes | Natalie on 12/18/13 |  | taken by Franz on 12/19/13
|  | 900 V | 720 | 764 | 729 | 692 | 648 | 692 | 656 | 676 | 536 | 508 | 67.9 | 68.2 | 57.7 | 60.3 | 55.1 | 54.2 | 60.3 | 61.6 | 71.6 | 70 | 1.44 | 1.491 | 807.2
|  | 1000 V | 1480 | 1430 | 1460 | 1330 | 1360 | 1510 | 1300 | 1270 | 712 | 764 | 103 | 103 | 104 | 105 | 102 | 103 | 115 | 111 | 90.1 | 88.8 | 1.34 | 1.27 | 900.45
|  | 1100 V | 2420 | 2490 | 2400 | 2610 | 2560 | 2270 | 1800 | 1690 | 440 | 790 | 175 | 170 | 174 | 167 | 167 | 162 | 147 | 148 | 186 | 167 | 1.221 | 1.123 | 992.75
CM9952 | yes | 850 V | 440 | 436 | 464 | 428 | 456 | 468 | 464 | 452 | 408 | 412 | 39 | 38 | 37.3 | 35.5 | 40 | 39.7 | 38.4 | 37 | 42.3 | 41.4 | 1.5 | 1.407 | 758.9 | 15.5 inches | yes | Natalie on 12/18/13 |  | taken by Franz on 12/19/13
|  | 900 V | 736 | 716 | 664 | 620 | 720 | 700 | 688 | 660 | 528 | 564 | 56 | 55.9 | 60.9 | 60 | 56.7 | 56.8 | 54.8 | 54.1 | 61.7 | 61.9 | 1.41 | 1.362 | 811.35
|  | 1000 V | 1420 | 1290 | 1520 | 1390 | 1320 | 1270 | 1380 | 1110 | 760 | 790 | 90.3 | 85.9 | 109 | 103 | 110 | 107 | 103 | 97.1 | 140 | 138 | 1.28 | 1.219 | 898.6
|  | 1100 V | 2290 | 2380 | 2280 | 2410 | 2660 | 2410 | 1640 | 1690 | 880 | 770 | 163 | 156 | 171 | 167 | 176 | 178 | 171 | 166 | 213 | 212 | 1.198 | 1.231 | 990.95
CM9914 | 14.0 power | 850 V | 464 | 502 | 444 | 494 | 496 | 458 | 476 | 478 | 368 | 382 | 39.5 | 41.7 | 38.7 | 38.4 | 36.1 | 36.2 | 42.7 | 43.7 | 44.8 | 44.3 | 1.446 | 1.523 | 758.3 | 15.5 inches | yes | Natalie on 12/19/13 |  | taken by Franz on 12/19/13
|  | 900 V | 624 | 660 | 704 | 716 | 712 | 692 | 608 | 620 | 552 | 564 | 68.1 | 69.7 | 68.7 | 68.3 | 56.3 | 57 | 59.8 | 60.6 | 60.1 | 60.2 | 1.317 | 1.374 | 804.7
|  | 1000 V | 1460 | 1430 | 1460 | 1510 | 1560 | 1650 | 1280 | 1430 | 740 | 690 | 105 | 109 | 112 | 117 | 109 | 112 | 108 | 112 | 118 | 116 | 1.298 | 1.24 | 898.1
|  | 1100 V | 2458 | 2537 | 2502 | 2449 | 2378 | 2492 | 1506 | 1668 | 833 | 881 | 65.55 | 63.95 | 64.09 | 60.01 | 62.45 | 66.9 | 112.3 | 102.8 | 154 | 159 | 1.213 | 1.214 | 990.45
CM9949 | 14.0 power | 850 V | 553 | 516 | 531 | 534 | 530 | 547 | 518 | 479 | 409 | 433 | 19 | 19 | 17 | 17 | 19 | 18 | 16 | 16 | 24 | 23 | 1.436 | 1.475 | 758.8 | 15.5 inches | yes | Natalie on 12/20/13
|  | 900 V | 775 | 721 | 794 | 782 | 824 | 819 | 748 | 752 | 617 | 619 | 35 | 35 | 24 | 24 | 25 | 25 | 26 | 26 | 31 | 31 | 1.435 | 1.478 | 805.5
|  | 1000 V | 1680 | 1700 | 1630 | 1620 | 1550 | 1580 | 1340 | 1270 | 712 | 725 | 39 | 39 | 49 | 49 | 41 | 42 | 57 | 58 | 53 | 52 | 1.247 | 1.309 | 898.35
|  | 1100 V | 2800 | 2780 | 2710 | 2760 | 2630 | 2700 | 1610 | 1680 | 827 | 760 | 71 | 75 | 66 | 69 | 68 | 69 | 89 | 87 | 106 | 88 | 1.214 | 1.162 | 990.5
|  |  | 10 kHz | 10 kHz | 1 MHz | 1 MHz | 2 MHz | 2 MHz | 5 MHz | 5 MHz | 10 MHz | 10 MHz | 10 kHz | 10 kHz | 1 MHz | 1 MHz | 2 MHz | 2 MHz | 5 MHz | 5 MHz | 10 MHz | 10 MHz | 5 MHz | 5 MHz | 10 kHz
CM9930 | yes | 850 V | 558 | 524 | 557 | 531 | 564 | 542 | 500 | 506 | 451 | 436 | 21 | 21 | 17 | 17 | 18 | 18 | 22 | 21 | 21 | 21 | 1.288 | 1.472 | 759.65 | 15.5 inches | yes | Natalie on 12/20/13
|  | 900 V | 831 | 867 | 831 | 815 | 837 | 895 | 793 | 804 | 575 | 581 | 24 | 24 | 28 | 28 | 29 | 29 | 31 | 31 | 28 | 28 | 1.423 | 1.403 | 807.45
|  | 1000 V | 1690 | 1670 | 1680 | 1720 | 1650 | 1620 | 1350 | 1430 | 761 | 685 | 37 | 37 | 49 | 49 | 43 | 44 | 48 | 48 | 57 | 57 | 1.247 | 1.273 | 899.3
|  | 1100 V | 2710 | 2780 | 2740 | 2720 | 2680 | 2600 | 1650 | 1550 | 800 | 845 | 66 | 67 | 67 | 69 | 64 | 63 | 83 | 81 | 217 | 79 | 1.269 | 1.171 | 991.4
CM9918 | yes | 850 V | 490 | 479 | 488 | 485 | 499 | 486 | 462 | 488 | 404 | 417 | 18 | 18 | 15 | 15 | 16 | 15 | 16 | 16 | 22 | 21 | 1.477 | 1.439 | 759.7 | 15.5 inches | yes | Natalie on 12/20/13
|  | 900 V | 799 | 780 | 770 | 765 | 767 | 749 | 736 | 718 | 585 | 567 | 23 | 23 | 25 | 25 | 24 | 24 | 26 | 27 | 22 | 23 | 1.415 | 1.4 | 806.55
|  | 1000 V | 1540 | 1510 | 1570 | 1520 | 1570 | 1510 | 1270 | 1390 | 755 | 771 | 40 | 44 | 37 | 39 | 45 | 47 | 56 | 57 | 54 | 53 | 1.295 | 1.319 | 899.4
|  | 1100 V | 2570 | 2670 | 2550 | 2450 | 2460 | 2620 | 1510 | 1650 | 783 | 912 | 58 | 59 | 62 | 62 | 65 | 63 | 77 | 77 | 117 | 83 | 1.23 | 1.139 | 991.65
CM9925 | yes | 850 V | 574 | 554 | 542 | 567 | 549 | 530 | 512 | 534 | 435 | 464 | 16 | 17 | 17 | 18 | 16 | 16 | 19 | 19 | 24 | 24 | 1.462 | 1.556 | 759.75 | 15.5 inches | yes | Natalie on 12/20/13
|  | 900 V | 862 | 834 | 852 | 883 | 893 | 881 | 796 | 776 | 565 | 618 | 27 | 27 | 27 | 28 | 28 | 27 | 27 | 27 | 19 | 19 | 1.444 | 1.408 | 806.7
|  | 1000 V | 1700 | 1670 | 1740 | 1710 | 1710 | 1660 | 1400 | 1410 | 788 | 750 | 44 | 47 | 36 | 37 | 45 | 44 | 59 | 58 | 60 | 58 | 1.239 | 1.25 | 899.45
|  | 1100 V | 2850 | 2920 | 2840 | 2790 | 2680 | 2890 | 1720 | 1620 | 737 | 705 | 76 | 77 | 74 | 77 | 63 | 62 | 70 | 67 | 115 | 73 | 1.248 | 1.146 | 991.65
|  |  | 10 kHz | 10 kHz | 1 MHz | 1 MHz | 2 MHz | 2 MHz | 5 MHz | 5 MHz | 10 MHz | 10 MHz | 10 kHz | 10 kHz | 1 MHz | 1 MHz | 2 MHz | 2 MHz | 5 MHz | 5 MHz | 10 MHz | 10 MHz | 5 MHz | 5 MHz | 10 kHz
CM9921 | yes | 850 V | 516 | 507 | 497 | 487 | 502 | 522 | 459 | 461 | 429 | 402 | 19 | 19 | 16 | 16 | 19 | 18 | 19 | 18 | 24 | 22 | 1.345 | 1.431 | 759.5 | 15.5 inches | yes | Natalie on 12/20/13
|  | 900 V | 792 | 751 | 770 | 782 | 783 | 777 | 739 | 746 | 619 | 635 | 26 | 27 | 23 | 22 | 26 | 25 | 33 | 33 | 31 | 30 | 1.436 | 1.451 | 806.8
|  | 1000 V | 1590 | 1530 | 1560 | 1450 | 1550 | 1490 | 1370 | 1360 | 785 | 701 | 38 | 37 | 46 | 46 | 41 | 42 | 48 | 47 | 51 | 50 | 1.349 | 1.36 | 899.2
47 total? |  | 1100 V | 2650 | 2630 | 2570 | 2520 | 2460 | 2530 | 1560 | 1680 | 778 | 825 | 77 | 75 | 69 | 68 | 64 | 62 | 64 | 63 | 74 | 69 | 1.03 | 1.161 | 991.55
CM9878 | yes; power setting: 13.30 | 850 V | 495 | 492 | 502 | 498 | 492 | 489 | 480 | 474 | 420 | 413 | 15 | 15 | 18 | 18 | 17 | 18 | 16 | 16 | 17 | 17 | 1.503 | 1.496 | 760.05 | 15.5 inches | yes | Nathan on 1/2/14
|  | 900 V | 802 | 795 | 799 | 792 | 777 | 772 | 737 | 732 | 582 | 573 | 27 | 28 | 29 | 28 | 25 | 24 | 28 | 27 | 26 | 25 | 1.466 | 1.417 | 806.45
|  | 1000 V | 1550 | 1550 | 1540 | 1530 | 1530 | 1520 | 1330 | 1310 | 764 | 745 | 43 | 44 | 44 | 43 | 42 | 42 | 42 | 41 | 40 | 39 | 1.34 | 1.348 | 899.4
|  | 1100 V | 2690 | 2670 | 2680 | 2650 | 2600 | 2550 | 1630 | 1600 | 820 | 810 | 63 | 62 | 68 | 69 | 67 | 66 | 64 | 64 | 64 | 65 | 1.217 | 1.176 | 991.4
CM9875 | power setting (13.3) fixed by last consistency test | 850 V | 448 | 446 | 440 | 440 | 427 | 427 | 412 | 410 | 357 | 355 | 15 | 15 | 15 | 15 | 15 | 14 | 15 | 15 | 14 | 15 | 1.492 | 1.467 | 759.65 | 15.5 inches | yes | Nathan on 1/2/14
|  | 900 V | 646 | 644 | 637 | 637 | 631 | 632 | 615 | 614 | 521 | 518 | 21 | 20 | 22 | 21 | 20 | 20 | 21 | 21 | 21 | 21 | 1.381 | 1.35 | 805.8
|  | 1000 V | 1330 | 1320 | 1320 | 1310 | 1310 | 1300 | 1170 | 1160 | 753 | 750 | 41 | 43 | 42 | 41 | 40 | 40 | 44 | 42 | 38 | 39 | 1.328 | 1.313 | 898.75
|  | 1100 V | 2320 | 2290 | 2310 | 2280 | 2260 | 2220 | 1610 | 1620 | 840 | 844 | 63 | 62 | 55 | 55 | 65 | 63 | 58 | 56 | 60 | 60 | 1.189 | 1.2 | 990.7
CM9871 | power setting (13.3) fixed by last consistency test | 850 V | 414 | 413 | 412 | 412 | 409 | 408 | 389 | 385 | 337 | 335 | 14 | 14 | 14 | 14 | 14 | 15 | 15 | 16 | 15 | 14 | 1.53 | 1.493 | 758.85 | 15.5 inches | yes | Nathan on 1/2/14
|  | 900 V | 609 | 610 | 614 | 615 | 605 | 605 | 583 | 581 | 492 | 487 | 21 | 21 | 21 | 21 | 22 | 22 | 24 | 24 | 20 | 20 | 1.411 | 1.361 | 805.1
|  | 1000 V | 1290 | 1290 | 1290 | 1290 | 1270 | 1260 | 1150 | 1150 | 743 | 731 | 41 | 42 | 42 | 42 | 41 | 41 | 40 | 40 | 40 | 39 | 1.311 | 1.344 | 897.95
|  | 1100 V | 2280 | 2280 | 2280 | 2270 | 2250 | 2230 | 1630 | 1580 | 828 | 810 | 64 | 66 | 61 | 62 | 68 | 68 | 60 | 60 | 66 | 66 | 1.162 | 1.197 | 990.1
|  |  | 10 kHz | 10 kHz | 1 MHz | 1 MHz | 2 MHz | 2 MHz | 5 MHz | 5 MHz | 10 MHz | 10 MHz | 10 kHz | 10 kHz | 1 MHz | 1 MHz | 2 MHz | 2 MHz | 5 MHz | 5 MHz | 10 MHz | 10 MHz | 5 MHz | 5 MHz | 10 kHz
CM9876 | power setting (13.3) fixed by last consistency test | 850 V | 538 | 544 | 530 | 537 | 527 | 532 | 500 | 503 | 427 | 428 | 18 | 19 | 17 | 17 | 15 | 15 | 19 | 19 | 19 | 19 | 1.47 | 1.447 | 760.25 | 15.5 inches | yes | Nathan on 1/2/14
|  | 900 V | 779 | 784 | 776 | 781 | 772 | 779 | 735 | 744 | 586 | 587 | 26 | 25 | 23 | 22 | 26 | 25 | 27 | 27 | 25 | 26 | 1.415 | 1.427 | 806.35
|  | 1000 V | 1560 | 1580 | 1540 | 1560 | 1530 | 1540 | 1330 | 1340 | 774 | 760 | 45 | 45 | 44 | 45 | 44 | 45 | 45 | 45 | 44 | 44 | 1.312 | 1.286 | 899.15
|  | 1100 V | 2610 | 2630 | 2580 | 2600 | 2530 | 2540 | 1650 | 1610 | 808 | 744 | 63 | 64 | 73 | 72 | 68 | 69 | 59 | 60 | 60 | 93 | 1.156 | 1.145 | 991.25
CM9870 | power setting (13.3) fixed by last consistency test | 850 V | 440 | 439 | 443 | 442 | 435 | 434 | 418 | 416 | 370 | 367 | 16 | 16 | 15 | 16 | 16 | 16 | 16 | 16 | 16 | 16 | 1.53 | 1.489 | 760.35 | 15.5 inches | yes | Nathan on 1/3/14
|  | 900 V | 666 | 664 | 665 | 664 | 664 | 663 | 637 | 635 | 536 | 532 | 24 | 24 | 23 | 23 | 24 | 24 | 23 | 23 | 21 | 21 | 1.458 | 1.442 | 806.7
|  | 1000 V | 1380 | 1370 | 1370 | 1360 | 1350 | 1340 | 1230 | 1220 | 745 | 742 | 43 | 44 | 41 | 40 | 40 | 40 | 40 | 41 | 40 | 40 | 1.306 | 1.331 | 899.6
|  | 1100 V | 2380 | 2360 | 2360 | 2340 | 2330 | 2300 | 1580 | 1590 | 758 | 808 | 64 | 63 | 65 | 64 | 62 | 63 | 60 | 61 | 91 | 65 | 1.235 | 1.266 | 991.65
CM9884 | power setting (13.3) fixed by last consistency test | 850 V | 432 | 440 | 431 | 437 | 425 | 431 | 400 | 405 | 351 | 353 | 13 | 14 | 15 | 15 | 14 | 14 | 15 | 15 | 14 | 14 | 1.499 | 1.502 | 758.6 | 15.5 inches | yes | Nathan on 1/3/14
|  | 900 V | 632 | 640 | 630 | 639 | 634 | 643 | 608 | 616 | 511 | 513 | 20 | 20 | 21 | 21 | 19 | 19 | 22 | 22 | 22 | 22 | 1.404 | 1.375 | 805
|  | 1000 V | 1320 | 1340 | 1300 | 1320 | 1290 | 1310 | 1170 | 1180 | 736 | 741 | 41 | 42 | 39 | 40 | 38 | 40 | 42 | 43 | 40 | 40 | 1.319 | 1.332 | 897.75
|  | 1100 V | 2300 | 2350 | 2290 | 2330 | 2260 | 2290 | 1570 | 1620 | 814 | 834 | 66 | 68 | 65 | 67 | 55 | 56 | 50 | 50 | 61 | 61 | 1.196 | 1.212 | 990
CM9877 | power setting (13.3) fixed by last consistency test | 850 V | 467 | 466 | 470 | 469 | 460 | 461 | 443 | 440 | 388 | 385 | 17 | 17 | 14 | 14 | 16 | 16 | 17 | 17 | 16 | 16 | 1.504 | 1.487 | 760 | 15.5 inches | yes | Nathan on 1/3/14
|  | 900 V | 694 | 693 | 692 | 692 | 687 | 688 | 658 | 658 | 542 | 540 | 21 | 21 | 24 | 24 | 23 | 23 | 24 | 25 | 22 | 23 | 1.438 | 1.363 | 806
|  | 1000 V | 1400 | 1400 | 1390 | 1390 | 1380 | 1370 | 1240 | 1240 | 754 | 764 | 45 | 45 | 40 | 41 | 44 | 43 | 40 | 41 | 41 | 40 | 1.358 | 1.349 | 898.75
|  | 1100 V | 2400 | 2390 | 2410 | 2400 | 2360 | 2340 | 1600 | 1640 | 755 | 827 | 58 | 58 | 61 | 60 | 60 | 61 | 58 | 59 | 84 | 56 | 1.211 | 1.237 | 990.75
|  |  | 10 kHz | 10 kHz | 1 MHz | 1 MHz | 2 MHz | 2 MHz | 5 MHz | 5 MHz | 10 MHz | 10 MHz | 10 kHz | 10 kHz | 1 MHz | 1 MHz | 2 MHz | 2 MHz | 5 MHz | 5 MHz | 10 MHz | 10 MHz | 5 MHz | 5 MHz | 10 kHz
CM9881 | power setting (13.3) fixed by last consistency test | 850 V | 440 | 437 | 445 | 442 | 442 | 439 | 423 | 419 | 372 | 368 | 17 | 16 | 16 | 15 | 17 | 17 | 16 | 16 | 14 | 14 | 1.476 | 1.492 | 758.25 | 15.5 inches | yes | Nathan on 1/3/14
|  | 900 V | 677 | 672 | 682 | 678 | 675 | 671 | 655 | 650 | 538 | 531 | 21 | 22 | 23 | 23 | 22 | 22 | 23 | 23 | 23 | 22 | 1.427 | 1.395 | 804.5
|  | 1000 V | 1400 | 1390 | 1400 | 1390 | 1390 | 1380 | 1250 | 1240 | 744 | 739 | 40 | 41 | 40 | 40 | 40 | 41 | 44 | 43 | 44 | 43 | 1.334 | 1.351 | 897.2
|  | 1100 V | 2440 | 2430 | 2470 | 2460 | 2410 | 2390 | 1590 | 1590 | 775 | 802 | 63 | 65 | 65 | 65 | 63 | 61 | 63 | 62 | 72 | 62 | 1.186 | 1.199 | 989.2
CM9867 | power setting (13.3) fixed by last consistency test | 850 V | 443 | 443 | 434 | 433 | 423 | 423 | 409 | 408 | 356 | 355 | 16 | 16 | 14 | 14 | 15 | 15 | 16 | 16 | 14 | 14 | 1.524 | 1.509 | 758.75 | 15.5 inches | yes | Nathan on 1/3/14
|  | 900 V | 649 | 650 | 643 | 644 | 640 | 641 | 617 | 618 | 520 | 517 | 22 | 22 | 23 | 23 | 21 | 21 | 24 | 24 | 22 | 23 | 1.383 | 1.409 | 805.05
|  | 1000 V | 1370 | 1370 | 1340 | 1330 | 1330 | 1330 | 1190 | 1190 | 747 | 748 | 40 | 40 | 40 | 40 | 41 | 41 | 43 | 42 | 40 | 40 | 1.299 | 1.318 | 897.9
|  | 1100 V | 2400 | 2390 | 2390 | 2370 | 2340 | 2320 | 1630 | 1650 | 839 | 848 | 67 | 68 | 65 | 64 | 64 | 65 | 58 | 59 | 61 | 60 | 1.159 | 1.216 | 990.05
CM9372 | power setting (13.3) fixed by last consistency test | 850 V | 579 | 577 | 566 | 565 | 562 | 561 | 543 | 541 | 457 | 454 | 19 | 19 | 20 | 20 | 21 | 21 | 20 | 20 | 20 | 20 | 1.466 | 1.47 | 759.75 | 15.5 inches | yes | Nathan on 1/3/14
|  | 900 V | 868 | 862 | 860 | 852 | 862 | 856 | 812 | 809 | 613 | 611 | 27 | 26 | 29 | 28 | 29 | 29 | 25 | 25 | 28 | 28 | 1.402 | 1.427 | 806.05
|  | 1000 V | 1780 | 1760 | 1760 | 1740 | 1740 | 1730 | 1430 | 1430 | 750 | 765 | 50 | 50 | 52 | 52 | 53 | 54 | 47 | 48 | 48 | 49 | 1.319 | 1.286 | 898.95
|  | 1100 V | 3010 | 3000 | 2980 | 2970 | 2870 | 2870 | 1560 | 1620 | 690 | 780 | 80 | 81 | 87 | 88 | 71 | 72 | 72 | 67 | 125 | 68 | 1.2 | 1.189 | 991.05
CM9890 | power setting (13.3) fixed by last consistency test | 850 V | 493 | 482 | 481 | 473 | 475 | 468 | 455 | 451 | 388 | 386 | 17 | 17 | 16 | 16 | 16 | 17 | 17 | 17 | 17 | 17 | 1.488 | 1.507 | 758.05 | 15.5 inches | yes | Nathan on 1/3/14
|  | 900 V | 735 | 721 | 732 | 720 | 726 | 716 | 689 | 686 | 548 | 550 | 25 | 24 | 23 | 22 | 24 | 25 | 23 | 23 | 24 | 24 | 1.427 | 1.415 | 804.6
|  | 1000 V | 1550 | 1520 | 1530 | 1500 | 1510 | 1490 | 1300 | 1300 | 743 | 744 | 47 | 48 | 47 | 46 | 48 | 49 | 48 | 47 | 45 | 45 | 1.332 | 1.318 | 897.45
|  | 1100 V | 2710 | 2660 | 2700 | 2650 | 2600 | 2570 | 1600 | 1630 | 803 | 808 | 69 | 70 | 61 | 62 | 60 | 61 | 64 | 64 | 61 | 62 | 1.154 | 1.166 | 989.75
|  |  | 10 kHz | 10 kHz | 1 MHz | 1 MHz | 2 MHz | 2 MHz | 5 MHz | 5 MHz | 10 MHz | 10 MHz | 10 kHz | 10 kHz | 1 MHz | 1 MHz | 2 MHz | 2 MHz | 5 MHz | 5 MHz | 10 MHz | 10 MHz | 5 MHz | 5 MHz | 10 kHz
CM9888 | power setting (13.3) fixed by last consistency test | 850 V | 482 | 433 | 480 | 430 | 473 | 425 | 452 | 406 | 398 | 357 | 17 | 16 | 15 | 14 | 15 | 14 | 19 | 17 | 17 | 15 | 1.507 | 1.51 | 759.95 | 15.5 inches | yes | Nathan on 1/3/14 | difference between amplitude of signal 1 and 2 is larger than usual
|  | 900 V | 723 | 646 | 727 | 651 | 723 | 644 | 693 | 618 | 557 | 502 | 26 | 24 | 27 | 25 | 27 | 23 | 22 | 19 | 21 | 19 | 1.423 | 1.442 | 806.05
|  | 1000 V | 1500 | 1330 | 1500 | 1340 | 1480 | 1340 | 1290 | 1190 | 745 | 681 | 47 | 41 | 37 | 37 | 41 | 40 | 44 | 42 | 41 | 38 | 1.31 | 1.247 | 898.9
|  | 1100 V | 2600 | 2430 | 2580 | 2390 | 2520 | 2300 | 1570 | 1470 | 752 | 726 | 61 | 58 | 70 | 69 | 60 | 59 | 70 | 65 | 75 | 55 | 1.174 | 1.207 | 990.95
CM9889 | 14.0 power | 850 V | 470 | 475 | 470 | 467 | 454 | 448 | 443 | 441 | 403 | 411 | 17 | 17 | 16 | 16 | 16 | 16 | 15 | 16 | 21 | 21 | 1.58 | 1.423 | 760.1 | 15.5 inches | yes | Natalie on 1/4/14
|  | 900 V | 693 | 699 | 740 | 738 | 708 | 711 | 702 | 715 | 581 | 563 | 28 | 28 | 22 | 23 | 23 | 23 | 25 | 25 | 23 | 24 | 1.501 | 1.426 | 806.6
|  | 1000 V | 1496 | 1483 | 1479 | 1499 | 1460 | 1473 | 1237 | 1291 | 751 | 730 | 37 | 39 | 40 | 43 | 40 | 42 | 45 | 44 | 42 | 41 | 1.289 | 1.314 | 899.9
|  | 1100 V | 2606 | 2457 | 2577 | 2580 | 2410 | 2455 | 1683 | 1769 | 765 | 767 | 67 | 65 | 65 | 65 | 62 | 63 | 72 | 73 | 74 | 58 | 1.159 | 1.22 | 992.2
CM9387 | 12.8 power | 850 V | 566 | 531 | 535 | 550 | 511 | 507 | 497 | 490 | 427 | 441 | 17 | 25 | 17 | 17 | 18 | 17 | 16 | 16 | 23 | 23 | 1.412 | 1.466 | 759.1 | 15.5 inches | yes | Natalie on 1/4/14
|  | 900 V | 849 | 866 | 804 | 796 | 828 | 822 | 762 | 772 | 621 | 630 | 30 | 28 | 28 | 27 | 29 | 28 | 30 | 30 | 35 | 35 | 1.382 | 1.388 | 806.15
|  | 1000 V | 1685 | 1710 | 1770 | 1635 | 1655 | 1630 | 1392 | 1315 | 819 | 803 | 53 | 52 | 55 | 53 | 53 | 52 | 57 | 55 | 46 | 45 | 1.231 | 1.322 | 899.9
|  | 1100 V | 2874 | 2953 | 2879 | 2764 | 2750 | 2755 | 1502 | 1645 | 785 | 811 | 74 | 71 | 78 | 76 | 62 | 61 | 75 | 77 | 114 | 80 | 1.15 | 1.21 | 991.1
CM9743 | yes previous | 850 V | 538 | 546 | 553 | 563 | 544 | 536 | 538 | 508 | 449 | 432 | 31 | 31 | 20 | 21 | 17 | 17 | 23 | 24 | 29 | 29 | 1.461 | 1.451 | 759.75 | 15.5 inches | yes | Natalie on 1/4/14 | fall times @ 1100 V were typos (73,74), so I deleted them (-Nathan)
|  | 900 V | 877 | 867 | 843 | 823 | 834 | 815 | 751 | 760 | 642 | 636 | 27 | 28 | 31 | 31 | 30 | 30 | 31 | 31 | 39 | 40 | 1.421 | 1.403 | 806.45
|  | 1000 V | 1802 | 1769 | 1666 | 1755 | 1757 | 1655 | 1369 | 1335 | 757 | 775 | 54 | 54 | 55 | 54 | 63 | 63 | 54 | 53 | 51 | 50 | 1.247 | 1.241 | 899.6
|  | 1100 V | 2989 | 2926 | 2719 | 2826 | 2815 | 2880 | 1642 | 1553 | 867 | 878 | 83 | 84 | 86 | 84 | 79 | 78 | 74 | 74 | 108 | 93 |  |  | 991.8
|  |  | 10 kHz | 10 kHz | 1 MHz | 1 MHz | 2 MHz | 2 MHz | 5 MHz | 5 MHz | 10 MHz | 10 MHz | 10 kHz | 10 kHz | 1 MHz | 1 MHz | 2 MHz | 2 MHz | 5 MHz | 5 MHz | 10 MHz | 10 MHz | 5 MHz | 5 MHz | 10 kHz
CM9392 | yes previous | 850 V | 512 | 496 | 506 | 498 | 470 | 488 | 487 | 479 | 392 | 393 | 19 | 20 | 18 | 18 | 18 | 19 | 20 | 21 | 21 | 22 | 1.484 | 1.39 | 759.6 | 15.5 inches | yes | Natalie on 1/4/14
|  | 900 V | 747 | 746 | 784 | 779 | 748 | 751 | 732 | 723 | 547 | 561 | 26 | 26 | 30 | 30 | 23 | 23 | 22 | 22 | 27 | 28 | 1.467 | 1.429 | 806.4
|  | 1000 V | 1548 | 1522 | 1509 | 1529 | 1494 | 1546 | 1351 | 1359 | 768 | 771 | 42 | 45 | 39 | 42 | 43 | 47 | 50 | 48 | 40 | 40 | 1.319 | 1.25 | 899.45
|  | 1100 V | 2493 | 2677 | 2712 | 2503 | 2531 | 2471 | 1460 | 1659 | 631 | 657 | 66 | 69 | 83 | 83 | 68 | 70 | 63 | 66 | 90 | 80 | 1.228 | 1.187 | 991.8
CM9371 | yes; power setting: 13.30 | 850 V | 517 | 521 | 529 | 533 | 528 | 531 | 502 | 504 | 437 | 435 | 15 | 15 | 15 | 15 | 17 | 18 | 17 | 17 | 19 | 19 | 1.478 | 1.49 | 759.6 | 15.5 inches | yes | Nathan on 1/6/14
|  | 900 V | 804 | 804 | 807 | 809 | 805 | 807 | 767 | 770 | 597 | 594 | 23 | 23 | 24 | 24 | 27 | 26 | 24 | 24 | 25 | 26 | 1.401 | 1.377 | 805.75
|  | 1000 V | 1610 | 1620 | 1610 | 1620 | 1600 | 1610 | 1380 | 1370 | 774 | 757 | 47 | 47 | 43 | 42 | 40 | 41 | 41 | 40 | 47 | 48 | 1.336 | 1.299 | 898.45
|  | 1100 V | 2830 | 2830 | 2830 | 2830 | 2730 | 2740 | 1660 | 1630 | 822 | 719 | 68 | 67 | 71 | 70 | 61 | 60 | 60 | 60 | 67 | 120 | 1.22 | 1.167 | 990.6
CM9833 | power setting (13.3) fixed by last consistency test | 850 V | 633 | 634 | 612 | 613 | 591 | 592 | 570 | 570 | 478 | 474 | 19 | 19 | 17 | 18 | 19 | 19 | 20 | 20 | 18 | 17 | 1.482 | 1.467 | 760.35 | 15.5 inches | yes | Nathan on 1/6/14
|  | 900 V | 934 | 925 | 910 | 904 | 905 | 899 | 852 | 850 | 627 | 622 | 29 | 29 | 29 | 28 | 31 | 32 | 31 | 30 | 29 | 30 | 1.397 | 1.417 | 806.55
|  | 1000 V | 1890 | 1870 | 1840 | 1830 | 1810 | 1800 | 1470 | 1450 | 766 | 758 | 49 | 50 | 51 | 50 | 46 | 46 | 47 | 46 | 48 | 49 | 1.326 | 1.263 | 899.4
|  | 1100 V | 3170 | 3120 | 3110 | 3070 | 2950 | 2920 | 1590 | 1450 | 747 | 741 | 76 | 74 | 74 | 71 | 72 | 70 | 66 | 172 | 66 | 69 | 1.157 | 1.222 | 991.35
CM9389 | power setting (13.3) fixed by last consistency test | 850 V | 535 | 534 | 538 | 536 | 536 | 535 | 510 | 509 | 435 | 433 | 19 | 19 | 19 | 19 | 19 | 19 | 20 | 19 | 20 | 20 | 1.495 | 1.512 | 760.75 | 15.5 inches | yes | Nathan on 1/6/14
|  | 900 V | 791 | 783 | 798 | 792 | 797 | 791 | 759 | 755 | 595 | 591 | 25 | 25 | 27 | 27 | 27 | 27 | 28 | 27 | 26 | 27 | 1.383 | 1.396 | 807
|  | 1000 V | 1590 | 1580 | 1590 | 1590 | 1570 | 1570 | 1370 | 1360 | 762 | 764 | 48 | 46 | 45 | 45 | 45 | 46 | 43 | 44 | 47 | 47 | 1.344 | 1.289 | 899.8
|  | 1100 V | 2730 | 2700 | 2730 | 2710 | 2640 | 2640 | 1610 | 1640 | 715 | 790 | 64 | 62 | 67 | 66 | 60 | 59 | 60 | 56 | 127 | 61 | 1.212 | 1.151 | 991.85
|  |  | 10 kHz | 10 kHz | 1 MHz | 1 MHz | 2 MHz | 2 MHz | 5 MHz | 5 MHz | 10 MHz | 10 MHz | 10 kHz | 10 kHz | 1 MHz | 1 MHz | 2 MHz | 2 MHz | 5 MHz | 5 MHz | 10 MHz | 10 MHz | 5 MHz | 5 MHz | 10 kHz
CM9391 | power setting (13.3) fixed by last consistency test | 850 V | 588 | 584 | 578 | 576 | 566 | 565 | 542 | 542 | 456 | 457 | 22 | 22 | 19 | 19 | 19 | 19 | 19 | 19 | 20 | 20 | 1.482 | 1.503 | 759.75 | 15.5 inches | yes | Nathan on 1/7/14 | took a long time for amplitude of signal 1 to settle @1100V,10MHz
|  | 900 V | 875 | 868 | 861 | 854 | 862 | 855 | 815 | 811 | 610 | 611 | 28 | 29 | 29 | 28 | 30 | 30 | 29 | 29 | 28 | 29 | 1.441 | 1.39 | 805.95
|  | 1000 V | 1790 | 1760 | 1750 | 1730 | 1700 | 1690 | 1410 | 1400 | 745 | 749 | 50 | 50 | 51 | 50 | 42 | 43 | 48 | 47 | 48 | 48 | 1.327 | 1.255 | 898.85
|  | 1100 V | 2990 | 2940 | 2950 | 2910 | 2810 | 2770 | 1560 | 1590 | 760 | 740 | 84 | 82 | 80 | 80 | 68 | 68 | 61 | 60 | 80 | 70 | 1.222 | 1.184 | 990.95
CM9909 | power setting (13.3) fixed by last consistency test | 850 V | 434 | 429 | 429 | 422 | 423 | 418 | 407 | 401 | 355 | 350 | 15 | 15 | 16 | 15 | 15 | 14 | 16 | 16 | 15 | 15 | 1.458 | 1.481 | 759.05 | 15.5 inches | yes | Nathan on 1/7/14
|  | 900 V | 650 | 642 | 650 | 644 | 640 | 634 | 618 | 611 | 521 | 514 | 21 | 21 | 25 | 25 | 23 | 23 | 25 | 25 | 21 | 21 | 1.405 | 1.365 | 805.45
|  | 1000 V | 1360 | 1340 | 1350 | 1330 | 1340 | 1310 | 1210 | 1180 | 755 | 739 | 41 | 41 | 42 | 41 | 39 | 40 | 42 | 41 | 42 | 42 | 1.317 | 1.305 | 898.25
|  | 1100 V | 2340 | 2310 | 2330 | 2300 | 2270 | 2240 | 1610 | 1570 | 850 | 820 | 55 | 55 | 57 | 58 | 57 | 55 | 54 | 55 | 55 | 58 | 1.214 | 1.197 | 990.35
CM9928 | power setting (13.3) fixed by last consistency test | 850 V | 460 | 456 | 452 | 450 | 444 | 440 | 422 | 417 | 369 | 366 | 16 | 16 | 17 | 17 | 17 | 17 | 18 | 18 | 18 | 18 | 1.517 | 1.463 | 759.9 | 15.5 inches | yes | Nathan on 1/7/14
|  | 900 V | 680 | 674 | 676 | 671 | 671 | 667 | 641 | 637 | 526 | 522 | 22 | 22 | 23 | 23 | 25 | 25 | 24 | 23 | 23 | 22 | 1.442 | 1.416 | 806.1
|  | 1000 V | 1390 | 1380 | 1370 | 1360 | 1370 | 1360 | 1220 | 1200 | 744 | 726 | 42 | 43 | 47 | 46 | 42 | 42 | 42 | 41 | 40 | 40 | 1.344 | 1.311 | 898.85
|  | 1100 V | 2390 | 2340 | 2380 | 2330 | 2330 | 2280 | 1640 | 1560 | 835 | 790 | 60 | 59 | 66 | 68 | 69 | 68 | 67 | 68 | 60 | 61 | 1.192 | 1.159 | 991
CM9892 | power setting (13.3) fixed by last consistency test | 850 V | 480 | 480 | 476 | 476 | 460 | 461 | 440 | 440 | 380 | 381 | 18 | 18 | 17 | 17 | 17 | 17 | 16 | 17 | 17 | 17 | 1.473 | 1.457 | 758.45 | 15.5 inches | yes | Nathan on 1/7/14
|  | 900 V | 702 | 700 | 700 | 699 | 695 | 694 | 665 | 667 | 549 | 548 | 24 | 23 | 24 | 24 | 25 | 25 | 23 | 24 | 24 | 24 | 1.39 | 1.434 | 804.85
|  | 1000 V | 1460 | 1460 | 1450 | 1450 | 1450 | 1440 | 1280 | 1280 | 760 | 762 | 42 | 41 | 44 | 45 | 44 | 44 | 42 | 42 | 45 | 45 | 1.326 | 1.272 | 897.7
|  | 1100 V | 2560 | 2530 | 2550 | 2530 | 2490 | 2480 | 1630 | 1650 | 830 | 827 | 66 | 65 | 71 | 71 | 65 | 65 | 65 | 66 | 70 | 69 | 1.14 | 1.165 | 989.95
|  |  | 10 kHz | 10 kHz | 1 MHz | 1 MHz | 2 MHz | 2 MHz | 5 MHz | 5 MHz | 10 MHz | 10 MHz | 10 kHz | 10 kHz | 1 MHz | 1 MHz | 2 MHz | 2 MHz | 5 MHz | 5 MHz | 10 MHz | 10 MHz | 5 MHz | 5 MHz | 10 kHz
CM9915 | power setting (13.3) fixed by last consistency test | 850 V | 449 | 448 | 448 | 448 | 444 | 443 | 427 | 425 | 372 | 371 | 17 | 17 | 16 | 16 | 16 | 16 | 17 | 17 | 15 | 15 | 1.487 | 1.5 | 758.9 | 15.5 inches | yes | Nathan on 1/7/14
|  | 900 V | 660 | 658 | 666 | 664 | 663 | 661 | 635 | 634 | 535 | 533 | 22 | 22 | 22 | 23 | 23 | 22 | 24 | 24 | 22 | 23 | 1.421 | 1.418 | 805.2
|  | 1000 V | 1360 | 1360 | 1360 | 1360 | 1350 | 1350 | 1220 | 1210 | 747 | 751 | 40 | 40 | 39 | 40 | 40 | 40 | 42 | 42 | 40 | 40 | 1.312 | 1.333 | 897.9
|  | 1100 V | 2370 | 2360 | 2380 | 2370 | 2330 | 2320 | 1630 | 1630 | 812 | 834 | 63 | 62 | 62 | 63 | 62 | 61 | 64 | 64 | 76 | 63 | 1.156 | 1.174 | 989.95
CM9908 | power setting (13.3) fixed by last consistency test | 850 V | 435 | 430 | 435 | 433 | 434 | 431 | 412 | 408 | 364 | 362 | 18 | 18 | 13 | 13 | 15 | 15 | 16 | 16 | 16 | 16 | 1.448 | 1.517 | 760.25 | 15.5 inches | yes | Nathan on 1/8/14
|  | 900 V | 670 | 667 | 663 | 660 | 660 | 659 | 630 | 629 | 528 | 524 | 21 | 21 | 19 | 19 | 22 | 22 | 24 | 23 | 22 | 23 | 1.405 | 1.419 | 806.65
|  | 1000 V | 1390 | 1380 | 1360 | 1360 | 1350 | 1340 | 1210 | 1200 | 750 | 755 | 42 | 42 | 40 | 40 | 42 | 42 | 41 | 41 | 38 | 39 | 1.333 | 1.347 | 899.5
|  | 1100 V | 2390 | 2370 | 2360 | 2340 | 2300 | 2290 | 1600 | 1630 | 770 | 820 | 63 | 60 | 61 | 63 | 64 | 64 | 58 | 59 | 89 | 59 | 1.187 | 1.1239 | 991.45
CM9935 | power setting (13.3) fixed by last consistency test | 850 V | 416 | 418 | 410 | 410 | 406 | 408 | 390 | 390 | 341 | 341 | 14 | 14 | 16 | 16 | 15 | 15 | 14 | 15 | 16 | 16 | 1.51 | 1.486 | 757.75 | 15.5 inches | yes | Nathan on 1/8/14
|  | 900 V | 626 | 628 | 620 | 622 | 618 | 620 | 590 | 593 | 497 | 497 | 23 | 23 | 22 | 22 | 23 | 23 | 21 | 21 | 20 | 20 | 1.427 | 1.447 | 804.2
|  | 1000 V | 1290 | 1310 | 1290 | 1300 | 1270 | 1280 | 1160 | 1160 | 734 | 740 | 40 | 40 | 42 | 43 | 40 | 41 | 40 | 39 | 38 | 38 | 1.35 | 1.345 | 896.95
|  | 1100 V | 2240 | 2260 | 2260 | 2280 | 2220 | 2230 | 1580 | 1600 | 812 | 824 | 60 | 61 | 60 | 59 | 58 | 58 | 54 | 55 | 59 | 58 | 1.214 | 1.266 | 989.05
CM9886 | power setting (13.3) fixed by last consistency test | 850 V | 447 | 456 | 442 | 450 | 431 | 439 | 405 | 411 | 353 | 359 | 15 | 16 | 16 | 17 | 15 | 15 | 16 | 17 | 17 | 17 | 1.457 | 1.463 | 759.6 | 15.5 inches | yes | Nathan on 1/8/14
|  | 900 V | 640 | 651 | 633 | 644 | 639 | 649 | 606 | 618 | 508 | 516 | 21 | 21 | 22 | 22 | 20 | 20 | 20 | 20 | 20 | 20 | 1.414 | 1.417 | 805.9
|  | 1000 V | 1340 | 1360 | 1320 | 1350 | 1310 | 1340 | 1190 | 1210 | 737 | 755 | 39 | 40 | 40 | 42 | 40 | 40 | 41 | 42 | 38 | 39 | 1.333 | 1.315 | 898.7
|  | 1100 V | 2310 | 2370 | 2310 | 2370 | 2250 | 2320 | 1560 | 1650 | 810 | 846 | 63 | 64 | 62 | 63 | 60 | 61 | 60 | 62 | 62 | 63 | 1.164 | 1.123 | 990.8
|  |  | 10 kHz | 10 kHz | 1 MHz | 1 MHz | 2 MHz | 2 MHz | 5 MHz | 5 MHz | 10 MHz | 10 MHz | 10 kHz | 10 kHz | 1 MHz | 1 MHz | 2 MHz | 2 MHz | 5 MHz | 5 MHz | 10 MHz | 10 MHz | 5 MHz | 5 MHz | 10 kHz
CM9901 | power setting (13.3) fixed by last consistency test | 850 V | 417 | 415 | 417 | 415 | 416 | 413 | 398 | 394 | 350 | 347 | 15 | 15 | 14 | 14 | 16 | 16 | 15 | 15 | 16 | 16 | 1.475 | 1.487 | 758.4 | 15.5 inches | yes | Nathan on 1/8/14
|  | 900 V | 649 | 645 | 645 | 641 | 644 | 641 | 614 | 611 | 515 | 513 | 20 | 20 | 22 | 21 | 21 | 21 | 20 | 20 | 23 | 23 | 1.411 | 1.381 | 804.75
|  | 1000 V | 1360 | 1350 | 1350 | 1340 | 1340 | 1330 | 1220 | 1200 | 746 | 747 | 42 | 42 | 41 | 42 | 43 | 43 | 40 | 39 | 41 | 40 | 1.28 | 1.319 | 897.5
|  | 1100 V | 2390 | 2360 | 2390 | 2360 | 2340 | 2320 | 1590 | 1610 | 820 | 821 | 66 | 65 | 64 | 63 | 65 | 64 | 63 | 62 | 63 | 62 | 1.152 | 1.221 | 989.7
CM9887 | power setting (13.3) fixed by last consistency test | 850 V | 383 | 386 | 371 | 374 | 365 | 368 | 348 | 351 | 310 | 312 | 16 | 16 | 14 | 14 | 13 | 13 | 13 | 13 | 15 | 15 | 1.48 | 1.495 | 757.9 | 15.5 inches | yes | Nathan on 1/8/14
|  | 900 V | 571 | 574 | 565 | 568 | 556 | 560 | 539 | 542 | 459 | 462 | 21 | 20 | 20 | 20 | 19 | 19 | 18 | 18 | 20 | 20 | 1.444 | 1.455 | 804.25
|  | 1000 V | 1230 | 1230 | 1200 | 1200 | 1200 | 1200 | 1100 | 1100 | 714 | 724 | 40 | 40 | 39 | 39 | 39 | 40 | 41 | 42 | 38 | 39 | 1.301 | 1.323 | 897.1
|  | 1100 V | 2150 | 2160 | 2140 | 2160 | 2110 | 2130 | 1550 | 1610 | 805 | 834 | 60 | 60 | 56 | 57 | 55 | 56 | 56 | 58 | 53 | 54 | 1.187 | 1.235 | 989.2
CM9897 | power setting (13.3) fixed by last consistency test | 850 V | 490 | 500 | 482 | 492 | 478 | 487 | 450 | 455 | 390 | 391 | 17 | 17 | 17 | 17 | 17 | 18 | 16 | 16 | 18 | 18 | 1.471 | 1.501 | 758.65 | 15.5 inches | yes | Nathan on 1/8/14
|  | 900 V | 715 | 725 | 708 | 718 | 702 | 712 | 670 | 679 | 550 | 552 | 24 | 24 | 25 | 25 | 25 | 26 | 25 | 25 | 23 | 24 | 1.395 | 1.423 | 805
|  | 1000 V | 1480 | 1530 | 1470 | 1510 | 1460 | 1490 | 1280 | 1290 | 753 | 759 | 44 | 45 | 43 | 44 | 44 | 45 | 45 | 45 | 40 | 40 | 1.313 | 1.287 | 897.8
|  | 1100 V | 2560 | 2620 | 2560 | 2610 | 2470 | 2520 | 1620 | 1640 | 817 | 820 | 65 | 64 | 65 | 66 | 68 | 68 | 63 | 63 | 65 | 64 | 1.166 | 1.205 | 990
1 MHz
CM9946 | yes; power setting: 14.60 | 850 V | [23 empty] | 15.5 inches | yes | Nathan on 1/21/14
|  | 900 V |  |  | 789 | 790 | 781 | 781 | 735 | 736 | 570 | 570 |  |  | 25 | 26 | 25 | 25 | 24 | 25 | 24 | 23 | 1.421 | 1.444 | 806.65
|  | 1000 V |  |  | 1550 | 1580 | 1530 | 1550 | 1330 | 1340 | 750 | 756 |  |  | 45 | 44 | 44 | 45 | 40 | 41 | 45 | 44 | 1.291 | 1.335 | 900
|  | 1100 V |  |  | 2630 | 2680 | 2520 | 2570 | 1570 | 1630 | 710 | 775 |  |  | 65 | 64 | 68 | 70 | 65 | 66 | 112 | 67 | 1.201 | 1.188 | 992.5
|  |  | 10 kHz | 10 kHz | 1 MHz | 1 MHz | 2 MHz | 2 MHz | 5 MHz | 5 MHz | 10 MHz | 10 MHz | 10 kHz | 10 kHz | 1 MHz | 1 MHz | 2 MHz | 2 MHz | 5 MHz | 5 MHz | 10 MHz | 10 MHz | 5 MHz | 5 MHz | 1 MHz
CM9229 | power setting (14.6) fixed by last consistency test | 850 V | [23 empty] | 15.5 inches | yes | Nathan on 1/21/14
|  | 900 V |  |  | 1140 | 1130 | 1130 | 1120 | 1020 | 1010 | 630 | 636 |  |  | 31 | 30 | 28 | 28 | 30 | 31 | 25 | 26 | 1.363 | 1.39 | 806.75
|  | 1000 V |  |  | 2240 | 2220 | 2170 | 2150 | 1450 | 1480 | 729 | 736 |  |  | 50 | 52 | 49 | 51 | 44 | 45 | 42 | 41 | 1.271 | 1.256 | 900.1
|  | 1100 V |  |  | 3600 | 3570 | 3250 | 3240 | 1440 | 1530 | 677 | 711 |  |  | 71 | 70 | 74 | 73 | 150 | 67 | 90 | 65 | 1.198 | 1.201 | 993.3
CM9233 | yes; power setting: 12.30 | 850 V | [23 empty] | 15.5 inches | yes | Nathan on 1/23/14
|  | 900 V |  |  | 800 | 797 | 788 | 786 | 740 | 741 | 573 | 575 |  |  | 26 | 25 | 27 | 26 | 26 | 25 | 28 | 27 | 1.431 | 1.415 | 806.1
|  | 1000 V |  |  | 1650 | 1650 | 1620 | 1620 | 1370 | 1370 | 750 | 756 |  |  | 43 | 42 | 45 | 45 | 47 | 48 | 47 | 47 | 1.282 | 1.294 | 899.4
|  | 1100 V |  |  | 2860 | 2860 | 2740 | 2740 | 1610 | 1640 | 800 | 810 |  |  | 68 | 65 | 71 | 70 | 69 | 70 | 72 | 70 | 1.188 | 1.158 | 992.3
CM9229 | power setting (12.3) fixed by last consistency test | 850 V | [23 empty] | 15.5 inches | yes | Nathan on 1/23/14 | retested with lower laser power setting
|  | 900 V |  |  | 730 | 726 | 715 | 710 | 680 | 681 | 534 | 535 |  |  | 27 | 28 | 26 | 25 | 24 | 25 | 21 | 22 | 1.4 | 1.41 | 806
|  | 1000 V |  |  | 1520 | 1510 | 1490 | 1480 | 1290 | 1280 | 725 | 731 |  |  | 48 | 48 | 47 | 48 | 50 | 51 | 51 | 50 | 1.318 | 1.288 | 899.7
|  | 1100 V |  |  | 2630 | 2610 | 2530 | 2510 | 1590 | 1620 | 780 | 800 |  |  | 74 | 75 | 70 | 68 | 68 | 66 | 71 | 69 | 1.207 | 1.171 | 992.6
CM9225 | yes; power setting: 11.50 | 850 V | [23 empty] | 15.5 inches | yes | Nathan on 1/23/14
|  | 900 V |  |  | 805 | 800 | 785 | 780 | 744 | 741 | 579 | 571 |  |  | 28 | 27 | 28 | 28 | 27 | 27 | 30 | 30 | 1.405 | 1.397 | 807.3
|  | 1000 V |  |  | 1600 | 1600 | 1580 | 1580 | 1350 | 1330 | 760 | 745 |  |  | 45 | 46 | 47 | 48 | 51 | 52 | 50 | 49 | 1.275 | 1.287 | 900.25
|  | 1100 V |  |  | 2720 | 2710 | 2610 | 2600 | 1660 | 1630 | 800 | 730 |  |  | 79 | 80 | 75 | 77 | 76 | 77 | 70 | 120 | 1.208 | 1.187 | 993.3
|  |  | 10 kHz | 10 kHz | 1 MHz | 1 MHz | 2 MHz | 2 MHz | 5 MHz | 5 MHz | 10 MHz | 10 MHz | 10 kHz | 10 kHz | 1 MHz | 1 MHz | 2 MHz | 2 MHz | 5 MHz | 5 MHz | 10 MHz | 10 MHz | 5 MHz | 5 MHz | 1 MHz
CM9231 | power setting (11.5) fixed by last consistency test | 850 V | [23 empty] | 15.5 inches | yes | Nathan on 1/23/14
|  | 900 V |  |  | 680 | 676 | 655 | 650 | 628 | 623 | 515 | 508 |  |  | 29 | 28 | 30 | 29 | 24 | 25 | 25 | 27 | 1.437 | 1.393 | 804
|  | 1000 V |  |  | 1420 | 1410 | 1400 | 1390 | 1240 | 1220 | 730 | 721 |  |  | 49 | 50 | 51 | 50 | 46 | 45 | 46 | 44 | 1.276 | 1.308 | 897.15
|  | 1100 V |  |  | 2530 | 2510 | 2440 | 2410 | 1620 | 1600 | 800 | 790 |  |  | 70 | 69 | 70 | 72 | 70 | 68 | 67 | 68 | 1.15 | 1.237 | 990.3
CM9232 | power setting (11.5) fixed by last consistency test | 850 V | [23 empty] | 15.5 inches | yes | Nathan on 1/23/14
|  | 900 V |  |  | 690 | 691 | 680 | 681 | 647 | 648 | 535 | 533 |  |  | 24 | 24 | 27 | 26 | 27 | 28 | 27 | 26 | 1.403 | 1.425 | 806.3
|  | 1000 V |  |  | 1430 | 1430 | 1410 | 1410 | 1240 | 1240 | 740 | 745 |  |  | 48 | 47 | 50 | 50 | 48 | 49 | 46 | 47 | 1.33 | 1.314 | 899.25
|  | 1100 V |  |  | 2500 | 2500 | 2410 | 2410 | 1600 | 1620 | 773 | 810 |  |  | 70 | 73 | 75 | 74 | 73 | 72 | 75 | 77 | 1.197 | 1.163 | 992
CM9942 | yes; power setting: 13.40 | 850 V | [23 empty] | 15.5 inches | yes | Nathan on 1/23/14
|  | 900 V |  |  | 785 | 788 | 775 | 777 | 728 | 733 | 574 | 575 |  |  | 27 | 26 | 26 | 25 | 23 | 25 | 23 | 23 | 1.38 | 1.399 | 805.8
|  | 1000 V |  |  | 1580 | 1600 | 1540 | 1570 | 1350 | 1350 | 760 | 744 |  |  | 48 | 48 | 50 | 51 | 44 | 46 | 43 | 47 | 1.337 | 1.27 | 898.95
|  | 1100 V |  |  | 2700 | 2730 | 2570 | 2610 | 1630 | 1600 | 855 | 851 |  |  | 70 | 68 | 64 | 66 | 68 | 70 | 64 | 67 | 1.201 | 1.159 | 991.6
CM9938 | power setting (13.4) fixed by last consistency test | 850 V | [23 empty] | 15.5 inches | yes | Nathan on 1/23/14
|  | 900 V |  |  | 740 | 737 | 730 | 728 | 682 | 684 | 550 | 548 |  |  | 28 | 27 | 24 | 23 | 22 | 22 | 23 | 22 | 1.397 | 1.427 | 807.7
|  | 1000 V |  |  | 1490 | 1490 | 1460 | 1450 | 1290 | 1280 | 747 | 753 |  |  | 46 | 47 | 45 | 46 | 46 | 46 | 49 | 50 | 1.344 | 1.304 | 900.8
|  | 1100 V |  |  | 2530 | 2500 | 2420 | 2400 | 1610 | 1620 | 831 | 824 |  |  | 69 | 70 | 55 | 55 | 54 | 55 | 56 | 58 | 1.171 | 1.203 | 993.55
|  |  | 10 kHz | 10 kHz | 1 MHz | 1 MHz | 2 MHz | 2 MHz | 5 MHz | 5 MHz | 10 MHz | 10 MHz | 10 kHz | 10 kHz | 1 MHz | 1 MHz | 2 MHz | 2 MHz | 5 MHz | 5 MHz | 10 MHz | 10 MHz | 5 MHz | 5 MHz | 1 MHz
CM9936 | power setting (13.4) fixed by last consistency test | 850 V | [23 empty] | 15.5 inches | yes | Nathan on 1/24/14
|  | 900 V |  |  | 690 | 699 | 672 | 680 | 645 | 653 | 523 | 525 |  |  | 23 | 23 | 24 | 23 | 22 | 23 | 26 | 25 | 1.417 | 1.404 | 806.3
|  | 1000 V |  |  | 1400 | 1420 | 1360 | 1380 | 1220 | 1230 | 755 | 744 |  |  | 45 | 44 | 47 | 48 | 44 | 45 | 43 | 43 | 1.345 | 1.317 | 899.3
|  | 1100 V |  |  | 2360 | 2400 | 2300 | 2350 | 1610 | 1600 | 830 | 811 |  |  | 68 | 70 | 66 | 69 | 68 | 67 | 66 | 67 | 1.236 | 1.16 | 991.95
CM9941 | power setting (13.4) fixed by last consistency test | 850 V | [23 empty] | 15.5 inches | yes | Nathan on 1/24/14
|  | 900 V |  |  | 646 | 651 | 630 | 634 | 591 | 596 | 500 | 501 |  |  | 24 | 24 | 24 | 24 | 23 | 22 | 21 | 22 | 1.422 | 1.41 | 806.35
|  | 1000 V |  |  | 1310 | 1320 | 1290 | 1300 | 1170 | 1170 | 752 | 756 |  |  | 44 | 45 | 43 | 42 | 46 | 46 | 43 | 43 | 1.344 | 1.318 | 899.5
|  | 1100 V |  |  | 2290 | 2310 | 2250 | 2260 | 1630 | 1650 | 831 | 846 |  |  | 68 | 71 | 68 | 70 | 69 | 68 | 64 | 66 | 1.221 | 1.18 | 992.2
CM9236 | yes; power setting: 12.30 | 850 V | [23 empty] | 15.5 inches | yes | Nathan on 1/24/14
|  | 900 V |  |  | 785 | 783 | 762 | 763 | 711 | 712 | 580 | 577 |  |  | 27 | 28 | 28 | 27 | 25 | 25 | 28 | 27 | 1.375 | 1.43 | 807.3
|  | 1000 V |  |  | 1580 | 1580 | 1570 | 1570 | 1350 | 1340 | 760 | 764 |  |  | 49 | 50 | 50 | 51 | 52 | 51 | 46 | 47 | 1.321 | 1.291 | 900.2
|  | 1100 V |  |  | 2720 | 2690 | 2620 | 2600 | 1650 | 1660 | 710 | 780 |  |  | 70 | 67 | 72 | 71 | 75 | 76 | 125 | 67 | 1.17 | 1.218 | 993.1
CM9241 | yes; power setting: 11.60 | 850 V | [23 empty] | 15.5 inches | yes | Nathan on 1/24/14
|  | 900 V |  |  | 800 | 802 | 788 | 791 | 746 | 747 | 588 | 587 |  |  | 28 | 27 | 28 | 28 | 28 | 27 | 26 | 27 | 1.402 | 1.393 | 806.3
|  | 1000 V |  |  | 1650 | 1660 | 1610 | 1620 | 1380 | 1370 | 755 | 756 |  |  | 48 | 49 | 49 | 48 | 45 | 45 | 46 | 48 | 1.319 | 1.32 | 899.4
|  | 1100 V |  |  | 2810 | 2800 | 2670 | 2650 | 1630 | 1640 | 711 | 765 |  |  | 76 | 77 | 77 | 78 | 74 | 73 | 108 | 73 | 1.169 | 1.221 | 992
|  |  | 10 kHz | 10 kHz | 1 MHz | 1 MHz | 2 MHz | 2 MHz | 5 MHz | 5 MHz | 10 MHz | 10 MHz | 10 kHz | 10 kHz | 1 MHz | 1 MHz | 2 MHz | 2 MHz | 5 MHz | 5 MHz | 10 MHz | 10 MHz | 5 MHz | 5 MHz | 1 MHz
CM9234 | power setting (11.6) fixed by last consistency test | 850 V | [23 empty] | 15.5 inches | yes | Nathan on 1/24/14
|  | 900 V |  |  | 721 | 712 | 708 | 700 | 666 | 659 | 536 | 530 |  |  | 27 | 26 | 32 | 31 | 28 | 27 | 28 | 28 | 1.4 | 1.382 | 805.3
|  | 1000 V |  |  | 1460 | 1440 | 1440 | 1410 | 1270 | 1240 | 755 | 742 |  |  | 48 | 49 | 48 | 48 | 45 | 46 | 49 | 48 | 1.36 | 1.335 | 898.6
|  | 1100 V |  |  | 2540 | 2490 | 2450 | 2390 | 1630 | 1590 | 830 | 817 |  |  | 75 | 74 | 83 | 85 | 83 | 83 | 72 | 71 | 1.198 | 1.213 | 991.7
CM9235 | power setting (11.6) fixed by last consistency test | 850 V | [23 empty] | 15.5 inches | yes | Nathan on 1/24/14
|  | 900 V |  |  | 752 | 749 | 740 | 739 | 690 | 688 | 550 | 546 |  |  | 30 | 29 | 26 | 28 | 29 | 29 | 29 | 28 | 1.401 | 1.377 | 806.75
|  | 1000 V |  |  | 1540 | 1530 | 1520 | 1510 | 1310 | 1300 | 760 | 750 |  |  | 48 | 49 | 47 | 48 | 44 | 45 | 48 | 49 | 1.311 | 1.3 | 899.7
|  | 1100 V |  |  | 2650 | 2640 | 2540 | 2530 | 1670 | 1650 | 800 | 720 |  |  | 83 | 79 | 77 | 78 | 81 | 78 | 70 | 126 | 1.199 | 1.163 | 992.4
CM9381 | power setting (11.6) fixed by last consistency test | 850 V | [23 empty] | 15.5 inches | yes | Nathan on 1/24/14
|  | 900 V |  |  | 671 | 676 | 663 | 665 | 640 | 641 | 533 | 530 |  |  | 25 | 28 | 30 | 30 | 26 | 27 | 24 | 25 | 1.4 | 1.359 | 807.3
|  | 1000 V |  |  | 1380 | 1390 | 1370 | 1370 | 1230 | 1230 | 762 | 744 |  |  | 48 | 49 | 45 | 47 | 48 | 47 | 49 | 47 | 1.344 | 1.271 | 900.3
|  | 1100 V |  |  | 2430 | 2460 | 2370 | 2390 | 1670 | 1630 | 820 | 760 |  |  | 75 | 77 | 75 | 75 | 76 | 78 | 68 | 110 | 1.147 | 1.176 | 993.1
CM9445 | power setting (11.6) fixed by last consistency test | 850 V | [23 empty] | 15.5 inches | yes | Nathan on 1/24/14
|  | 900 V |  |  | 685 | 687 | 677 | 683 | 646 | 649 | 527 | 526 |  |  | 26 | 27 | 27 | 27 | 24 | 25 | 23 | 24 | 1.39 | 1.397 | 807.15
|  | 1000 V |  |  | 1400 | 1400 | 1390 | 1390 | 1250 | 1250 | 755 | 762 |  |  | 50 | 51 | 45 | 46 | 44 | 43 | 49 | 48 | 1.308 | 1.29 | 899.75
|  | 1100 V |  |  | 2470 | 2470 | 2400 | 2390 | 1640 | 1680 | 745 | 811 |  |  | 76 | 78 | 80 | 78 | 81 | 80 | 105 | 68 | 1.2 | 1.236 | 992.7
|  |  | 10 kHz | 10 kHz | 1 MHz | 1 MHz | 2 MHz | 2 MHz | 5 MHz | 5 MHz | 10 MHz | 10 MHz | 10 kHz | 10 kHz | 1 MHz | 1 MHz | 2 MHz | 2 MHz | 5 MHz | 5 MHz | 10 MHz | 10 MHz | 5 MHz | 5 MHz | 1 MHz
CM9446 | power setting (11.6) fixed by last consistency test | 850 V | [23 empty] | 15.5 inches | yes | Nathan on 1/27/14
|  | 900 V |  |  | 720 | 697 | 690 | 671 | 650 | 644 | 526 | 524 |  |  | 27 | 27 | 28 | 28 | 28 | 29 | 27 | 28 | 1.415 | 1.4218 | 805.9
|  | 1000 V |  |  | 1470 | 1430 | 1450 | 1420 | 1260 | 1240 | 738 | 742 |  |  | 46 | 47 | 47 | 47 | 49 | 48 | 50 | 50 | 1.344 | 1.318 | 899
|  | 1100 V |  |  | 2590 | 2520 | 2470 | 2410 | 1640 | 1670 | 810 | 827 |  |  | 70 | 70 | 70 | 68 | 74 | 72 | 70 | 69 | 1.205 | 1.224 | 991.8
CM9444 | power setting (11.6) fixed by last consistency test | 850 V | [23 empty] | 15.5 inches | yes | Nathan on 1/27/14
|  | 900 V |  |  | 755 | 752 | 730 | 726 | 700 | 697 | 550 | 547 |  |  | 25 | 24 | 25 | 27 | 29 | 29 | 27 | 28 | 1.419 | 1.405 | 808.1
|  | 1000 V |  |  | 1510 | 1500 | 1490 | 1480 | 1300 | 1280 | 750 | 740 |  |  | 46 | 47 | 48 | 49 | 49 | 50 | 50 | 49 | 1.288 | 1.298 | 901.05
|  | 1100 V |  |  | 2550 | 2530 | 2500 | 2460 | 1650 | 1610 | 780 | 720 |  |  | 71 | 72 | 75 | 74 | 70 | 70 | 70 | 110 | 1.202 | 1.158 | 993.75
CM9441 | power setting (11.6) fixed by last consistency test | 850 V | [23 empty] | 15.5 inches | yes | Nathan on 1/27/14
|  | 900 V |  |  | 701 | 704 | 690 | 692 | 660 | 663 | 545 | 546 |  |  | 27 | 28 | 28 | 28 | 27 | 27 | 26 | 27 | 1.424 | 1.418 | 806.35
|  | 1000 V |  |  | 1460 | 1470 | 1430 | 1440 | 1280 | 1270 | 760 | 758 |  |  | 46 | 48 | 47 | 47 | 51 | 52 | 50 | 49 | 1.345 | 1.328 | 899.4
|  | 1100 V |  |  | 2530 | 2530 | 2440 | 2440 | 1640 | 1640 | 800 | 823 |  |  | 68 | 70 | 70 | 71 | 68 | 69 | 72 | 70 | 1.165 | 1.17 | 992.25
CM9250 | power setting (11.6) fixed by last consistency test | 850 V | [23 empty] | 15.5 inches | yes | Nathan on 1/27/14
|  | 900 V |  |  | 670 | 669 | 650 | 652 | 614 | 615 | 506 | 504 |  |  | 26 | 26 | 27 | 26 | 22 | 23 | 25 | 25 | 1.423 | 1.409 | 806.55
|  | 1000 V |  |  | 1360 | 1350 | 1370 | 1360 | 1220 | 1220 | 756 | 742 |  |  | 49 | 49 | 48 | 49 | 48 | 47 | 46 | 47 | 1.282 | 1.294 | 899.6
|  | 1100 V |  |  | 2420 | 2390 | 2350 | 2320 | 1630 | 1560 | 800 | 710 |  |  | 67 | 65 | 68 | 69 | 74 | 72 | 70 | 108 | 1.198 | 1.181 | 992.2
|  |  | 10 kHz | 10 kHz | 1 MHz | 1 MHz | 2 MHz | 2 MHz | 5 MHz | 5 MHz | 10 MHz | 10 MHz | 10 kHz | 10 kHz | 1 MHz | 1 MHz | 2 MHz | 2 MHz | 5 MHz | 5 MHz | 10 MHz | 10 MHz | 5 MHz | 5 MHz | 1 MHz
CM9243 | power setting (11.6) fixed by last consistency test | 850 V | [23 empty] | 15.5 inches | yes | Nathan on 1/27/14
|  | 900 V |  |  | 720 | 730 | 700 | 710 | 650 | 655 | 539 | 533 |  |  | 25 | 25 | 26 | 25 | 27 | 26 | 29 | 29 | 1.403 | 1.377 | 806.1
|  | 1000 V |  |  | 1450 | 1480 | 1420 | 1440 | 1250 | 1250 | 746 | 728 |  |  | 48 | 49 | 46 | 47 | 52 | 52 | 45 | 46 | 1.33 | 1.272 | 899.15
|  | 1100 V |  |  | 2510 | 2550 | 2430 | 2460 | 1630 | 1590 | 838 | 823 |  |  | 70 | 72 | 72 | 74 | 73 | 72 | 69 | 70 | 1.21 | 1.196 | 991.9
CM9242 | power setting (11.6) fixed by last consistency test | 850 V | [23 empty] | 15.5 inches | yes | Nathan on 1/27/14
|  | 900 V |  |  | 610 | 606 | 593 | 590 | 570 | 563 | 471 | 470 |  |  | 22 | 23 | 22 | 23 | 23 | 24 | 24 | 24 | 1.433 | 1.351 | 805.15
|  | 1000 V |  |  | 1310 | 1300 | 1290 | 1270 | 1160 | 1150 | 725 | 733 |  |  | 50 | 51 | 51 | 51 | 49 | 48 | 49 | 50 | 1.291 | 1.326 | 898.4
|  | 1100 V |  |  | 2340 | 2300 | 2280 | 2250 | 1600 | 1630 | 830 | 844 |  |  | 68 | 69 | 70 | 71 | 71 | 72 | 68 | 70 | 1.209 | 1.229 | 991.3
CM9383 | power setting (11.6) fixed by last consistency test | 850 V | [23 empty] | 15.5 inches | yes | Nathan on 1/28/14
|  | 900 V |  |  | 630 | 636 | 627 | 630 | 600 | 603 | 510 | 508 |  |  | 26 | 27 | 23 | 24 | 25 | 26 | 26 | 26 | 1.412 | 1.43 | 806.5
|  | 1000 V |  |  | 1340 | 1340 | 1320 | 1320 | 1180 | 1170 | 742 | 733 |  |  | 50 | 52 | 48 | 49 | 50 | 51 | 49 | 50 | 1.329 | 1.285 | 899.55
|  | 1100 V |  |  | 2340 | 2360 | 2260 | 2270 | 1630 | 1610 | 820 | 782 |  |  | 71 | 72 | 67 | 69 | 69 | 70 | 68 | 70 | 1.179 | 1.189 | 992.1
CM9385 | power setting (11.6) fixed by last consistency test | 850 V | [23 empty] | 15.5 inches | yes | Nathan on 1/28/14
|  | 900 V |  |  | 705 | 711 | 690 | 695 | 665 | 670 | 542 | 545 |  |  | 27 | 28 | 26 | 25 | 25 | 26 | 28 | 27 | 1.395 | 1.408 | 807.05
|  | 1000 V |  |  | 1430 | 1430 | 1420 | 1420 | 1280 | 1280 | 765 | 770 |  |  | 50 | 49 | 49 | 50 | 50 | 50 | 48 | 50 | 1.301 | 1.289 | 899.8
|  | 1100 V |  |  | 2520 | 2510 | 2440 | 2430 | 1660 | 1680 | 800 | 842 |  |  | 70 | 71 | 69 | 70 | 72 | 73 | 90 | 60 | 1.18 | 1.208 | 992.65
|  |  | 10 kHz | 10 kHz | 1 MHz | 1 MHz | 2 MHz | 2 MHz | 5 MHz | 5 MHz | 10 MHz | 10 MHz | 10 kHz | 10 kHz | 1 MHz | 1 MHz | 2 MHz | 2 MHz | 5 MHz | 5 MHz | 10 MHz | 10 MHz | 5 MHz | 5 MHz | 1 MHz
CM9245 | power setting (11.6) fixed by last consistency test | 850 V | [23 empty] | 15.5 inches | yes | Nathan on 1/28/14
|  | 900 V |  |  | 750 | 752 | 724 | 726 | 690 | 695 | 557 | 555 |  |  | 25 | 25 | 27 | 28 | 27 | 26 | 26 | 26 | 1.37 | 1.4 | 807.45
|  | 1000 V |  |  | 1480 | 1480 | 1460 | 1460 | 1300 | 1290 | 770 | 760 |  |  | 51 | 50 | 52 | 52 | 51 | 50 | 50 | 51 | 1.26 | 1.295 | 900.1
|  | 1100 V |  |  | 2560 | 2610 | 2470 | 2430 | 1660 | 1590 | 830 | 710 |  |  | 70 | 72 | 69 | 70 | 70 | 70 | 70 | 130 | 1.199 | 1.156 | 992.9
CM9244 | power setting (11.6) fixed by last consistency test | 850 V | [23 empty] | 15.5 inches | yes | Nathan on 1/28/14
|  | 900 V |  |  | 640 | 640 | 622 | 619 | 599 | 600 | 495 | 493 |  |  | 26 | 26 | 26 | 25 | 27 | 26 | 25 | 25 | 1.44 | 1.393 | 806.15
|  | 1000 V |  |  | 1310 | 1300 | 1300 | 1290 | 1160 | 1160 | 740 | 730 |  |  | 49 | 50 | 51 | 52 | 47 | 48 | 47 | 49 | 1.304 | 1.324 | 899
|  | 1100 V |  |  | 2310 | 2330 | 2250 | 2270 | 1570 | 1580 | 815 | 820 |  |  | 68 | 70 | 69 | 70 | 71 | 70 | 67 | 69 | 1.169 | 1.228 | 991.9
CM9248 | power setting (11.6) fixed by last consistency test | 850 V | [23 empty] | 15.5 inches | yes | Nathan on 1/28/14
|  | 900 V |  |  | 675 | 671 | 655 | 652 | 626 | 623 | 515 | 513 |  |  | 27 | 28 | 25 | 25 | 24 | 24 | 25 | 25 | 1.453 | 1.423 | 806.4
|  | 1000 V |  |  | 1360 | 1350 | 1340 | 1330 | 1200 | 1190 | 750 | 730 |  |  | 48 | 48 | 45 | 46 | 50 | 50 | 47 | 49 | 1.358 | 1.337 | 899.25
|  | 1100 V |  |  | 2400 | 2350 | 2300 | 2250 | 1610 | 1520 | 800 | 680 |  |  | 70 | 70 | 73 | 72 | 75 | 74 | 60 | 95 | 1.221 | 1.148 | 991.9
CM9463 | power setting (11.6) fixed by last consistency test | 850 V | [23 empty] | 15.5 inches | yes | Nathan on 1/28/14
|  | 900 V |  |  | 720 | 722 | 715 | 717 | 686 | 686 | 544 | 542 |  |  | 28 | 27 | 26 | 26 | 24 | 26 | 24 | 25 | 1.398 | 1.405 | 807.5
|  | 1000 V |  |  | 1480 | 1480 | 1460 | 1460 | 1250 | 1250 | 750 | 752 |  |  | 49 | 51 | 46 | 48 | 50 | 50 | 50 | 50 | 1.329 | 1.267 | 900.4
|  | 1100 V |  |  | 2560 | 2560 | 2450 | 2450 | 1630 | 1650 | 700 | 770 |  |  | 71 | 72 | 69 | 70 | 67 | 69 | 120 | 67 | 1.153 | 1.228 | 993.1
|  |  | 10 kHz | 10 kHz | 1 MHz | 1 MHz | 2 MHz | 2 MHz | 5 MHz | 5 MHz | 10 MHz | 10 MHz | 10 kHz | 10 kHz | 1 MHz | 1 MHz | 2 MHz | 2 MHz | 5 MHz | 5 MHz | 10 MHz | 10 MHz | 5 MHz | 5 MHz | 1 MHz
CM9471 | power setting (11.6) fixed by last consistency test | 850 V | [23 empty] | 15.5 inches | yes | Nathan on 1/30/14
|  | 900 V |  |  | 620 | 629 | 614 | 621 | 590 | 598 | 500 | 502 |  |  | 27 | 28 | 26 | 26 | 25 | 25 | 24 | 25 | 1.409 | 1.357 | 806.55
|  | 1000 V |  |  | 1330 | 1350 | 1320 | 1340 | 1200 | 1200 | 750 | 740 |  |  | 46 | 48 | 50 | 49 | 52 | 53 | 45 | 46 | 1.339 | 1.32 | 899.8
|  | 1100 V |  |  | 2310 | 2330 | 2270 | 2290 | 1620 | 1580 | 820 | 790 |  |  | 71 | 72 | 68 | 70 | 70 | 70 | 72 | 75 | 1.245 | 1.214 | 992.5
CM9413 | power setting (11.6) fixed by last consistency test | 850 V | [23 empty] | 15.5 inches | yes | Nathan on 1/30/14
|  | 900 V |  |  | 681 | 682 | 661 | 662 | 628 | 628 | 520 | 514 |  |  | 27 | 26 | 26 | 25 | 29 | 29 | 28 | 27 | 1.427 | 1.383 | 805.1
|  | 1000 V |  |  | 1440 | 1450 | 1420 | 1420 | 1250 | 1230 | 751 | 733 |  |  | 48 | 52 | 50 | 52 | 50 | 50 | 49 | 50 | 1.327 | 1.27 | 898.3
|  | 1100 V |  |  | 2490 | 2490 | 2420 | 2420 | 1650 | 1610 | 820 | 810 |  |  | 73 | 75 | 69 | 70 | 71 | 72 | 70 | 72 | 1.25 | 1.189 | 991.05
CM9414 | power setting (11.6) fixed by last consistency test | 850 V | [23 empty] | 15.5 inches | yes | Nathan on 1/30/14
|  | 900 V |  |  | 730 | 732 | 713 | 717 | 678 | 682 | 554 | 552 |  |  | 27 | 28 | 27 | 27 | 25 | 25 | 28 | 28 | 1.411 | 1.386 | 807.6
|  | 1000 V |  |  | 1480 | 1490 | 1470 | 1480 | 1300 | 1290 | 762 | 758 |  |  | 52 | 50 | 50 | 52 | 50 | 50 | 47 | 49 | 1.357 | 1.311 | 900.55
|  | 1100 V |  |  | 2550 | 2560 | 2490 | 2490 | 1660 | 1640 | 800 | 770 |  |  | 72 | 74 | 72 | 73 | 71 | 72 | 80 | 110 | 1.157 | 1.23 | 993.3
CM9451 | power setting (11.6) fixed by last consistency test | 850 V | [23 empty] | 15.5 inches | yes | Nathan on 1/30/14
|  | 900 V |  |  | 695 | 700 | 668 | 671 | 643 | 649 | 522 | 527 |  |  | 28 | 27 | 28 | 27 | 26 | 27 | 27 | 27 | 1.434 | 1.398 | 808.1
|  | 1000 V |  |  | 1400 | 1410 | 1380 | 1390 | 1230 | 1230 | 745 | 748 |  |  | 50 | 52 | 49 | 50 | 48 | 49 | 49 | 50 | 1.345 | 1.323 | 901.05
|  | 1100 V |  |  | 2460 | 2490 | 2360 | 2390 | 1630 | 1650 | 780 | 730 |  |  | 69 | 70 | 70 | 72 | 68 | 70 | 80 | 120 | 1.221 | 1.173 | 994.1
|  |  | 10 kHz | 10 kHz | 1 MHz | 1 MHz | 2 MHz | 2 MHz | 5 MHz | 5 MHz | 10 MHz | 10 MHz | 10 kHz | 10 kHz | 1 MHz | 1 MHz | 2 MHz | 2 MHz | 5 MHz | 5 MHz | 10 MHz | 10 MHz | 5 MHz | 5 MHz | 1 MHz
CM9457 | power setting (11.6) fixed by last consistency test | 850 V | [23 empty] | 15.5 inches | yes | Nathan on 1/30/14
|  | 900 V |  |  | 710 | 708 | 698 | 695 | 660 | 662 | 541 | 539 |  |  | 26 | 26 | 25 | 25 | 28 | 29 | 25 | 27 | 1.399 | 1.443 | 807.5
|  | 1000 V |  |  | 1460 | 1450 | 1450 | 1440 | 1290 | 1280 | 732 | 740 |  |  | 48 | 47 | 51 | 50 | 51 | 48 | 46 | 48 | 1.342 | 1.296 | 900.6
|  | 1100 V |  |  | 2530 | 2530 | 2470 | 2460 | 1600 | 1630 | 710 | 790 |  |  | 74 | 75 | 72 | 70 | 70 | 69 | 125 | 72 | 1.132 | 1.153 | 993.3
CM9456 | power setting (11.6) fixed by last consistency test | 850 V | [23 empty] | 15.5 inches | yes | Nathan on 1/31/14
|  | 900 V |  |  | 683 | 681 | 670 | 668 | 635 | 638 | 530 | 528 |  |  | 25 | 24 | 27 | 28 | 27 | 28 | 26 | 27 | 1.41 | 1.39 | 806.05
|  | 1000 V |  |  | 1400 | 1400 | 1390 | 1390 | 1240 | 1230 | 740 | 750 |  |  | 52 | 50 | 51 | 52 | 48 | 49 | 46 | 48 | 1.28 | 1.319 | 898.85
|  | 1100 V |  |  | 2450 | 2440 | 2370 | 2360 | 1600 | 1630 | 812 | 838 |  |  | 70 | 72 | 70 | 68 | 72 | 74 | 70 | 71 | 1.144 | 1.184 | 991.5
CM9464 | power setting (11.6) fixed by last consistency test | 850 V | [23 empty] | 15.5 inches | yes | Nathan on 1/31/14
|  | 900 V |  |  | 694 | 690 | 674 | 675 | 632 | 636 | 520 | 522 |  |  | 25 | 26 | 25 | 25 | 26 | 27 | 26 | 27 | 1.425 | 1.456 | 806.75
|  | 1000 V |  |  | 1430 | 1430 | 1380 | 1380 | 1230 | 1230 | 730 | 732 |  |  | 47 | 49 | 50 | 51 | 51 | 50 | 49 | 49 | 1.344 | 1.315 | 899.9
|  | 1100 V |  |  | 2440 | 2420 | 2360 | 2340 | 1610 | 1620 | 813 | 818 |  |  | 70 | 73 | 70 | 72 | 68 | 70 | 69 | 70 | 1.178 | 1.206 | 992.6
CM9317 | yes; power setting: 11.60 | 850 V | [23 empty] | 15.5 inches | yes | Nathan on 3/14/14 | Changed scopes; scope set to average of 4 mode
|  | 900 V |  |  | 796 | 794 | 789 | 787 | 740 | 742 | 541 | 547 |  |  | 25 | 25 | 24 | 25 | 28 | 28 | 25 | 26 | 1.442 | 1.446 | 806
|  | 1000 V |  |  | 1650 | 1640 | 1610 | 1610 | 1250 | 1270 | 670 | 690 |  |  | 45 | 45 | 47 | 46 | 44 | 45 | 42 | 43 | 1.247 | 1.232 | 899.2
|  | 1100 V |  |  | 2740 | 2740 | 2580 | 2590 | 1350 | 1420 | 610 | 680 |  |  | 70 | 70 | 67 | 68 | 70 | 68 | 85 | 60 | 1.015 | 1.175 | 992.25
|  |  | 10 kHz | 10 kHz | 1 MHz | 1 MHz | 2 MHz | 2 MHz | 5 MHz | 5 MHz | 10 MHz | 10 MHz | 10 kHz | 10 kHz | 1 MHz | 1 MHz | 2 MHz | 2 MHz | 5 MHz | 5 MHz | 10 MHz | 10 MHz | 5 MHz | 5 MHz | 1 MHz
CM9317 | yes; power setting: 11.60 | 850 V | [23 empty] | 15.5 inches | yes | Nathan on 3/14/14 | retest with scope set to sample mode
|  | 900 V |  |  | 795 | 791 | 784 | 785 | 737 | 738 | 535 | 542 |  |  | 72 | 73 | 65 | 67 | 66 | 68 | 65 | 66 | 1.501 | 1.506 | 805.6
|  | 1000 V |  |  | 1630 | 1630 | 1600 | 1590 | 1260 | 1280 | 672 | 692 |  |  | 130 | 130 | 130 | 130 | 120 | 120 | 110 | 110 | 1.237 | 1.229 | 899.05
|  | 1100 V |  |  | 2770 | 2760 | 2580 | 2590 | 1350 | 1420 | 610 | 680 |  |  | 200 | 200 | 190 | 200 | 180 | 180 | 160 | 160 | 1.15 | 1.222 | 992.3
CM9302 | yes; power setting: 10.80 | 850 V | [23 empty] | 15.5 inches | yes | Nathan on 3/14/14
|  | 900 V |  |  | 802 | 801 | 798 | 796 | 752 | 752 | 553 | 558 |  |  | 28 | 28 | 27 | 29 | 24 | 24 | 26 | 27 | 1.452 | 1.448 | 806
|  | 1000 V |  |  | 1653 | 1650 | 1630 | 1630 | 1260 | 1280 | 666 | 683 |  |  | 45 | 47 | 48 | 48 | 45 | 46 | 45 | 44 | 1.272 | 1.276 | 899.4
|  | 1100 V |  |  | 2790 | 2790 | 2630 | 2630 | 1270 | 1370 | 590 | 660 |  |  | 72 | 71 | 69 | 70 | 76 | 72 | 110 | 70 | 1.255 | 1.315 | 992.55
CM9478 | yes; power setting: 11.60 | 850 V | [23 empty] | 15.5 inches | yes | Nathan on 3/14/14 | changed order of diffuser/attenuator
|  | 900 V |  |  | 783 | 787 | 774 | 778 | 744 | 748 | 560 | 567 |  |  | 27 | 28 | 24 | 24 | 29 | 28 | 22 | 24 | 1.18 | 1.198 | 806.75 |  |  |  | switched back to original order after this test
|  | 1000 V |  |  | 1640 | 1640 | 1610 | 1620 | 1350 | 1360 | 720 | 750 |  |  | 47 | 47 | 45 | 46 | 49 | 50 | 47 | 48 | 1.049 | 1.031 | 900
|  | 1100 V |  |  | 2720 | 2770 | 2590 | 2650 | 1430 | 1590 | 670 | 817 |  |  | 75 | 74 | 70 | 70 | 68 | 70 | 160 | 69 | 0.769 | 0.902 | 992.95
CM9888 | power setting: 11.60 | 850 V | [23 empty] | 15.5 inches | yes | Nathan on 3/17/14 | retest
|  | 900 V |  |  | 694 | 706 | 675 | 687 | 640 | 650 | 508 | 514 |  |  | 25 | 25 | 23 | 25 | 26 | 26 | 23 | 24 | 1.491 | 1.5 | 806.6
|  | 1000 V |  |  | 1450 | 1460 | 1410 | 1420 | 1200 | 1200 | 684 | 680 |  |  | 45 | 44 | 43 | 44 | 42 | 43 | 42 | 42 | 1.284 | 1.303 | 899.75
|  | 1100 V |  |  | 2470 | 2490 | 2350 | 2370 | 1380 | 1350 | 690 | 620 |  |  | 73 | 72 | 68 | 67 | 55 | 55 | 59 | 100 | 1.23 | 1.223 | 992.65
|  |  | 10 kHz | 10 kHz | 1 MHz | 1 MHz | 2 MHz | 2 MHz | 5 MHz | 5 MHz | 10 MHz | 10 MHz | 10 kHz | 10 kHz | 1 MHz | 1 MHz | 2 MHz | 2 MHz | 5 MHz | 5 MHz | 10 MHz | 10 MHz | 5 MHz | 5 MHz | 1 MHz
CM9478 | power setting: 11.00 | 850 V | [23 empty] | 15.5 inches | yes | Nathan on 3/17/14 | retest
|  | 900 V |  |  | 725 | 732 | 722 | 729 | 692 | 698 | 520 | 530 |  |  | 25 | 25 | 28 | 28 | 24 | 24 | 27 | 27 | 1.491 | 1.523 | 806.7
|  | 1000 V |  |  | 1510 | 1520 | 1500 | 1510 | 1220 | 1240 | 668 | 698 |  |  | 45 | 44 | 46 | 46 | 42 | 43 | 43 | 44 | 1.323 | 1.282 | 899.95
|  | 1100 V |  |  | 2570 | 2610 | 2460 | 2500 | 1320 | 1460 | 578 | 700 |  |  | 70 | 72 | 61 | 63 | 72 | 70 | 116 | 60 | 1.229 | 1.24 | 992.8
CM9485 | power setting: 10.80 | 850 V | [23 empty] | 15.5 inches | yes | Nathan on 3/17/14
|  | 900 V |  |  | 766 | 758 | 772 | 765 | 730 | 723 | 538 | 542 |  |  | 26 | 26 | 26 | 25 | 26 | 27 | 28 | 29 | 1.492 | 1.476 | 807.3
|  | 1000 V |  |  | 1570 | 1540 | 1560 | 1530 | 1260 | 1260 | 664 | 687 |  |  | 45 | 44 | 50 | 50 | 45 | 46 | 44 | 45 | 1.341 | 1.27 | 900.5
|  | 1100 V |  |  | 2660 | 2630 | 2550 | 2520 | 1240 | 1410 | 630 | 680 |  |  | 70 | 73 | 74 | 72 | 175 | 70 | 70 | 55 | 1.403 | 1.371 | 993.45
CM9493 | power setting: 11.20 | 850 V | [23 empty] | 15.5 inches | yes | Nathan on 3/18/14
|  | 900 V |  |  | 761 | 757 | 755 | 753 | 708 | 708 | 530 | 534 |  |  | 25 | 26 | 23 | 24 | 25 | 25 | 22 | 23 | 1.441 | 1.43 | 807.7
|  | 1000 V |  |  | 1530 | 1530 | 1500 | 1500 | 1240 | 1240 | 676 | 697 |  |  | 41 | 42 | 45 | 44 | 45 | 45 | 40 | 39 | 1.3 | 1.23 | 900.7
|  | 1100 V |  |  | 2600 | 2600 | 2460 | 2460 | 1280 | 1450 | 630 | 700 |  |  | 70 | 71 | 64 | 65 | 120 | 70 | 82 | 63 | 1.373 | 1.278 | 993.9
CM9501 | power setting: 11.20 | 850 V | [23 empty] | 15.5 inches | yes | Nathan on 3/18/14
|  | 900 V |  |  | 791 | 794 | 782 | 784 | 732 | 734 | 542 | 540 |  |  | 29 | 30 | 25 | 26 | 25 | 25 | 23 | 23 | 1.453 | 1.412 | 807.15
|  | 1000 V |  |  | 1600 | 1600 | 1570 | 1570 | 1270 | 1260 | 682 | 675 |  |  | 43 | 44 | 45 | 44 | 42 | 43 | 45 | 44 | 1.262 | 1.236 | 900.1
|  | 1100 V |  |  | 2690 | 2690 | 2550 | 2550 | 1360 | 1220 | 672 | 653 |  |  | 71 | 72 | 71 | 70 | 73 | 170 | 68 | 65 | 1.333 | 1.387 | 992.95
|  |  | 10 kHz | 10 kHz | 1 MHz | 1 MHz | 2 MHz | 2 MHz | 5 MHz | 5 MHz | 10 MHz | 10 MHz | 10 kHz | 10 kHz | 1 MHz | 1 MHz | 2 MHz | 2 MHz | 5 MHz | 5 MHz | 10 MHz | 10 MHz | 5 MHz | 5 MHz | 1 MHz
CM9502 | power setting: 11.20 | 850 V | [23 empty] | 15.5 inches | yes | Nathan on 3/18/14
|  | 900 V |  |  | 797 | 795 | 786 | 784 | 735 | 735 | 533 | 537 |  |  | 25 | 25 | 27 | 27 | 23 | 24 | 23 | 22 | 1.48 | 1.49 | 806.15
|  | 1000 V |  |  | 1590 | 1580 | 1560 | 1550 | 1230 | 1240 | 656 | 675 |  |  | 44 | 43 | 43 | 42 | 42 | 42 | 41 | 42 | 1.32 | 1.307 | 899.6
|  | 1100 V |  |  | 2630 | 2620 | 2470 | 2460 | 1330 | 1380 | 583 | 675 |  |  | 69 | 68 | 65 | 65 | 70 | 70 | 84 | 68 | 1.201 | 1.289 | 992.6
CM9555 | power setting: 11.00 | 850 V | [23 empty] | 15.5 inches | yes | Nathan on 3/18/14
|  | 900 V |  |  | 794 | 799 | 808 | 812 | 758 | 763 | 550 | 558 |  |  | 27 | 27 | 29 | 30 | 30 | 30 | 25 | 25 | 1.408 | 1.415 | 807.25
|  | 1000 V |  |  | 1640 | 1640 | 1610 | 1610 | 1250 | 1270 | 665 | 690 |  |  | 45 | 46 | 44 | 44 | 45 | 46 | 46 | 48 | 1.275 | 1.244 | 900.25
|  | 1100 V |  |  | 2750 | 2740 | 2570 | 2570 | 1180 | 1350 | 595 | 667 |  |  | 69 | 68 | 72 | 70 | 150 | 70 | 70 | 60 | 1.414 | 1.385 | 993.25
CM9353 | power setting: 11.00 | 850 V | [23 empty] | 15.5 inches | yes | Nathan on 3/18/14
|  | 900 V |  |  | 770 | 768 | 768 | 764 | 718 | 715 | 540 | 538 |  |  | 25 | 25 | 24 | 24 | 25 | 25 | 26 | 25 | 1.4 | 1.395 | 806.45
|  | 1000 V |  |  | 1550 | 1540 | 1530 | 1520 | 1250 | 1230 | 685 | 681 |  |  | 45 | 45 | 42 | 43 | 46 | 45 | 43 | 42 | 1.21 | 1.223 | 899.35
|  | 1100 V |  |  | 2590 | 2580 | 2480 | 2460 | 1460 | 1390 | 695 | 630 |  |  | 70 | 71 | 70 | 69 | 68 | 73 | 70 | 112 | 1.15 | 1.232 | 991.95
CM9352 | power setting: 11.00 | 850 V | [23 empty] | 15.5 inches | yes | Nathan on 3/18/14
|  | 900 V |  |  | 785 | 777 | 775 | 769 | 715 | 709 | 533 | 530 |  |  | 27 | 27 | 28 | 27 | 25 | 24 | 25 | 25 | 1.437 | 1.4 | 805.3
|  | 1000 V |  |  | 1570 | 1550 | 1550 | 1530 | 1260 | 1250 | 680 | 675 |  |  | 45 | 44 | 48 | 48 | 45 | 45 | 43 | 45 | 1.227 | 1.168 | 898.55
|  | 1100 V |  |  | 2620 | 2600 | 2480 | 2470 | 1420 | 1420 | 674 | 692 |  |  | 71 | 70 | 67 | 67 | 65 | 65 | 70 | 72 | 1.011 | 1.048 | 991.45
|  |  | 10 kHz | 10 kHz | 1 MHz | 1 MHz | 2 MHz | 2 MHz | 5 MHz | 5 MHz | 10 MHz | 10 MHz | 10 kHz | 10 kHz | 1 MHz | 1 MHz | 2 MHz | 2 MHz | 5 MHz | 5 MHz | 10 MHz | 10 MHz | 5 MHz | 5 MHz | 1 MHz
CM9259 | power setting: 11.00 | 850 V | [23 empty] | 15.5 inches | yes | Nathan on 3/19/14
|  | 900 V |  |  | 765 | 768 | 756 | 761 | 714 | 716 | 535 | 534 |  |  | 27 | 27 | 26 | 27 | 24 | 25 | 23 | 24 | 1.494 | 1.486 | 806.95
|  | 1000 V |  |  | 1570 | 1570 | 1540 | 1540 | 1250 | 1240 | 684 | 681 |  |  | 43 | 42 | 45 | 44 | 45 | 45 | 43 | 42 | 1.266 | 1.325 | 900.05
|  | 1100 V |  |  | 2660 | 2660 | 2520 | 2510 | 1420 | 1380 | 675 | 625 |  |  | 70 | 70 | 70 | 69 | 66 | 67 | 68 | 105 | 1.19 | 1.221 | 993
CM9255 | power setting: 10.80 | 850 V | [23 empty] | 15.5 inches | yes | Nathan on 3/19/14
|  | 900 V |  |  | 761 | 760 | 750 | 750 | 703 | 703 | 533 | 533 |  |  | 28 | 27 | 25 | 26 | 26 | 25 | 26 | 26 | 1.53 | 1.529 | 806.3
|  | 1000 V |  |  | 1540 | 1530 | 1530 | 1520 | 1240 | 1230 | 681 | 675 |  |  | 47 | 48 | 51 | 51 | 45 | 46 | 43 | 45 | 1.344 | 1.367 | 899.45
|  | 1100 V |  |  | 2590 | 2580 | 2460 | 2460 | 1440 | 1390 | 695 | 631 |  |  | 73 | 72 | 70 | 69 | 68 | 70 | 66 | 98 | 1.296 | 1.396 | 992.3
CM9252 | power setting: 10.80 | 850 V | [23 empty] | 15.5 inches | yes | Nathan on 3/19/14
|  | 900 V |  |  | 760 | 761 | 745 | 747 | 702 | 707 | 525 | 534 |  |  | 27 | 28 | 25 | 25 | 26 | 27 | 26 | 25 | 1.465 | 1.445 | 805.3
|  | 1000 V |  |  | 1550 | 1550 | 1520 | 1520 | 1240 | 1250 | 667 | 684 |  |  | 44 | 44 | 44 | 44 | 47 | 47 | 44 | 44 | 1.311 | 1.233 | 898.1
|  | 1100 V |  |  | 2600 | 2600 | 2460 | 2470 | 1360 | 1430 | 610 | 684 |  |  | 70 | 70 | 68 | 70 | 66 | 68 | 97 | 66 | 1.244 | 1.218 | 990.75
CM9253 | power setting: 10.80 | 850 V | [23 empty] | 15.5 inches | yes | Nathan on 3/19/14
|  | 900 V |  |  | 772 | 767 | 754 | 750 | 705 | 706 | 525 | 531 |  |  | 22 | 23 | 28 | 27 | 26 | 26 | 24 | 25 | 1.545 | 1.503 | 806.05
|  | 1000 V |  |  | 1570 | 1560 | 1540 | 1530 | 1240 | 1250 | 672 | 684 |  |  | 45 | 44 | 47 | 46 | 46 | 45 | 43 | 44 | 1.378 | 1.285 | 899.2
|  | 1100 V |  |  | 2650 | 2640 | 2490 | 2490 | 1400 | 1440 | 627 | 681 |  |  | 70 | 68 | 72 | 70 | 67 | 67 | 93 | 65 | 1.274 | 1.201 | 992
|  |  | 10 kHz | 10 kHz | 1 MHz | 1 MHz | 2 MHz | 2 MHz | 5 MHz | 5 MHz | 10 MHz | 10 MHz | 10 kHz | 10 kHz | 1 MHz | 1 MHz | 2 MHz | 2 MHz | 5 MHz | 5 MHz | 10 MHz | 10 MHz | 5 MHz | 5 MHz | 1 MHz
CM9551 | power setting: 10.80 | 850 V | [23 empty] | 15.5 inches | yes | Nathan on 3/20/14
|  | 900 V |  |  | 715 | 730 | 710 | 726 | 670 | 685 | 522 | 529 |  |  | 24 | 24 | 25 | 25 | 26 | 28 | 23 | 23 | 1.47 | 1.5 | 806.2
|  | 1000 V |  |  | 1480 | 1500 | 1450 | 1480 | 1220 | 1230 | 685 | 677 |  |  | 42 | 42 | 45 | 45 | 41 | 42 | 40 | 40 | 1.282 | 1.318 | 899.3
|  | 1100 V |  |  | 2530 | 2580 | 2370 | 2410 | 1470 | 1420 | 713 | 655 |  |  | 68 | 70 | 69 | 71 | 70 | 68 | 70 | 92 | 1.197 | 1.171 | 992.2
CM9511 | power setting: 10.80 | 850 V | [23 empty] | 15.5 inches | yes | Nathan on 3/20/14
|  | 900 V |  |  | 727 | 733 | 721 | 724 | 688 | 692 | 518 | 523 |  |  | 25 | 25 | 25 | 25 | 26 | 25 | 22 | 23 | 1.529 | 1.485 | 807.5
|  | 1000 V |  |  | 1490 | 1500 | 1470 | 1480 | 1220 | 1220 | 657 | 672 |  |  | 43 | 44 | 45 | 45 | 48 | 50 | 43 | 45 | 1.351 | 1.272 | 900.55
|  | 1100 V |  |  | 2550 | 2570 | 2420 | 2440 | 1330 | 1400 | 612 | 672 |  |  | 70 | 71 | 65 | 67 | 70 | 68 | 108 | 70 | 1.3 | 1.203 | 993.6
CM9504 | power setting: 10.80 | 850 V | [23 empty] | 15.5 inches | yes | Nathan on 3/20/14 | large difference between amplitudes at 1100V
|  | 900 V |  |  | 751 | 744 | 736 | 730 | 688 | 684 | 528 | 525 |  |  | 28 | 27 | 25 | 26 | 25 | 25 | 26 | 26 | 1.584 | 1.6 | 806.6
|  | 1000 V |  |  | 1520 | 1480 | 1500 | 1470 | 1240 | 1220 | 698 | 672 |  |  | 43 | 44 | 45 | 45 | 44 | 43 | 40 | 41 | 1.37 | 1.434 | 899.75
|  | 1100 V |  |  | 2570 | 2500 | 2450 | 2370 | 1470 | 1300 | 693 | 577 |  |  | 71 | 70 | 75 | 72 | 68 | 90 | 62 | 118 | 1.382 | 1.49 | 992.6
CM9415 | power setting: 10.80 | 850 V | [23 empty] | 15.5 inches | yes | Nathan on 3/20/14
|  | 900 V |  |  | 740 | 747 | 736 | 745 | 702 | 709 | 537 | 539 |  |  | 25 | 25 | 25 | 26 | 27 | 26 | 25 | 25 | 1.47 | 1.431 | 807.55
|  | 1000 V |  |  | 1510 | 1520 | 1490 | 1500 | 1250 | 1240 | 684 | 678 |  |  | 47 | 47 | 48 | 48 | 44 | 45 | 45 | 45 | 1.28 | 1.265 | 900.65
|  | 1100 V |  |  | 2540 | 2570 | 2440 | 2470 | 1450 | 1380 | 709 | 647 |  |  | 73 | 75 | 73 | 72 | 70 | 74 | 60 | 87 | 1.248 | 1.284 | 993.55
|  |  | 10 kHz | 10 kHz | 1 MHz | 1 MHz | 2 MHz | 2 MHz | 5 MHz | 5 MHz | 10 MHz | 10 MHz | 10 kHz | 10 kHz | 1 MHz | 1 MHz | 2 MHz | 2 MHz | 5 MHz | 5 MHz | 10 MHz | 10 MHz | 5 MHz | 5 MHz | 1 MHz
CM9418 | power setting: 10.80 | 850 V | [23 empty] | 15.5 inches | yes | Nathan on 3/20/14
|  | 900 V |  |  | 710 | 708 | 698 | 695 | 655 | 657 | 514 | 521 |  |  | 25 | 25 | 23 | 22 | 25 | 25 | 23 | 22 | 1.51 | 1.503 | 806.55
|  | 1000 V |  |  | 1440 | 1430 | 1410 | 1400 | 1190 | 1200 | 665 | 683 |  |  | 46 | 45 | 44 | 43 | 46 | 45 | 42 | 43 | 1.336 | 1.304 | 899.6
|  | 1100 V |  |  | 2440 | 2420 | 2340 | 2330 | 1410 | 1470 | 627 | 700 |  |  | 70 | 68 | 68 | 71 | 67 | 64 | 110 | 70 | 1.227 | 1.221 | 992.45
CM9417 | power setting: 11.00 | 850 V | [23 empty] | 15.5 inches | yes | Nathan on 3/21/14
|  | 900 V |  |  | 722 | 725 | 712 | 716 | 678 | 683 | 523 | 526 |  |  | 25 | 25 | 24 | 24 | 25 | 25 | 23 | 24 | 1.507 | 1.52 | 805.5
|  | 1000 V |  |  | 1530 | 1530 | 1490 | 1490 | 1240 | 1240 | 687 | 680 |  |  | 48 | 47 | 47 | 46 | 48 | 48 | 44 | 46 | 1.312 | 1.318 | 898.7
|  | 1100 V |  |  | 2580 | 2580 | 2440 | 2450 | 1420 | 1380 | 701 | 625 |  |  | 72 | 70 | 70 | 71 | 65 | 67 | 67 | 96 | 1.32 | 1.3 | 991.55
CM9433 | power setting: 10.80 | 850 V | [23 empty] | 15.5 inches | yes | Nathan on 3/21/14
|  | 900 V |  |  | 770 | 773 | 760 | 764 | 712 | 717 | 530 | 533 |  |  | 25 | 24 | 26 | 25 | 27 | 27 | 25 | 25 | 1.492 | 1.462 | 806.2
|  | 1000 V |  |  | 1570 | 1570 | 1540 | 1540 | 1240 | 1240 | 690 | 684 |  |  | 46 | 46 | 49 | 47 | 49 | 50 | 45 | 44 | 1.322 | 1.273 | 899.4
|  | 1100 V |  |  | 2680 | 2690 | 2530 | 2550 | 1410 | 1300 | 655 | 580 |  |  | 74 | 76 | 68 | 70 | 74 | 96 | 70 | 110 | 1.356 | 1.382 | 992.3
CM9432 | power setting: 10.80 | 850 V | [23 empty] | 15.5 inches | yes | Nathan on 3/21/14
|  | 900 V |  |  | 792 | 799 | 777 | 782 | 730 | 734 | 540 | 544 |  |  | 25 | 25 | 26 | 27 | 28 | 27 | 26 | 26 | 1.445 | 1.464 | 807.9
|  | 1000 V |  |  | 1580 | 1590 | 1560 | 1560 | 1260 | 1250 | 677 | 673 |  |  | 45 | 45 | 46 | 45 | 47 | 48 | 40 | 42 | 1.238 | 1.287 | 901
|  | 1100 V |  |  | 2680 | 2690 | 2550 | 2560 | 1420 | 1350 | 692 | 636 |  |  | 71 | 72 | 69 | 69 | 70 | 72 | 70 | 93 | 1.221 | 1.303 | 993.95
|  |  | 10 kHz | 10 kHz | 1 MHz | 1 MHz | 2 MHz | 2 MHz | 5 MHz | 5 MHz | 10 MHz | 10 MHz | 10 kHz | 10 kHz | 1 MHz | 1 MHz | 2 MHz | 2 MHz | 5 MHz | 5 MHz | 10 MHz | 10 MHz | 5 MHz | 5 MHz | 1 MHz
CM9437 | power setting: 10.80 | 850 V | [23 empty] | 15.5 inches | yes | Nathan on 3/21/14
|  | 900 V |  |  | 720 | 742 | 706 | 725 | 670 | 680 | 520 | 524 |  |  | 24 | 24 | 27 | 27 | 25 | 25 | 23 | 23 | 1.58 | 1.52 | 804.9
|  | 1000 V |  |  | 1480 | 1530 | 1460 | 1490 | 1220 | 1220 | 680 | 686 |  |  | 47 | 48 | 50 | 51 | 44 | 45 | 41 | 40 | 1.38 | 1.319 | 898.1
|  | 1100 V |  |  | 2520 | 2590 | 2400 | 2440 | 1410 | 1440 | 672 | 711 |  |  | 75 | 75 | 69 | 70 | 68 | 65 | 72 | 68 | 1.353 | 1.278 | 991.05
CM9379 | power setting: 10.80 | 850 V | [23 empty] | 15.5 inches | yes | Nathan on 3/21/14
|  | 900 V |  |  | 735 | 750 | 725 | 740 | 685 | 698 | 517 | 525 |  |  | 27 | 27 | 29 | 30 | 28 | 28 | 24 | 26 | 1.521 | 1.525 | 805.5
|  | 1000 V |  |  | 1500 | 1530 | 1450 | 1480 | 1200 | 1220 | 650 | 679 |  |  | 47 | 48 | 48 | 50 | 44 | 45 | 45 | 46 | 1.38 | 1.324 | 899
|  | 1100 V |  |  | 2470 | 2540 | 2350 | 2410 | 1330 | 1440 | 674 | 721 |  |  | 75 | 76 | 69 | 68 | 68 | 70 | 70 | 70 | 1.39 | 1.293 | 992.1
CM9509 | power setting: 10.80 | 850 V | [23 empty] | 15.5 inches | yes | Nathan on 3/25/14
|  | 900 V |  |  | 722 | 724 | 713 | 716 | 677 | 681 | 517 | 525 |  |  | 25 | 24 | 23 | 24 | 26 | 26 | 25 | 27 | 1.48 | 1.544 | 806.7
|  | 1000 V |  |  | 1470 | 1470 | 1440 | 1440 | 1210 | 1220 | 667 | 688 |  |  | 45 | 45 | 44 | 44 | 45 | 44 | 42 | 43 | 1.305 | 1.35 | 899.7
|  | 1100 V |  |  | 2490 | 2480 | 2380 | 2380 | 1400 | 1460 | 633 | 707 |  |  | 69 | 70 | 67 | 65 | 64 | 66 | 94 | 62 | 1.212 | 1.293 | 992.5
CM9512 | power setting: 11.00 | 850 V | [23 empty] | 15.5 inches | yes | Nathan on 3/25/14
|  | 900 V |  |  | 738 | 742 | 722 | 727 | 676 | 683 | 515 | 525 |  |  | 28 | 27 | 25 | 25 | 27 | 27 | 24 | 23 | 1.513 | 1.49 | 806.7
|  | 1000 V |  |  | 1490 | 1490 | 1450 | 1450 | 1200 | 1210 | 665 | 690 |  |  | 44 | 43 | 45 | 44 | 44 | 43 | 43 | 45 | 1.355 | 1.303 | 899.9
|  | 1100 V |  |  | 2500 | 2490 | 2370 | 2360 | 1360 | 1440 | 602 | 699 |  |  | 70 | 68 | 71 | 72 | 65 | 64 | 95 | 64 | 1.353 | 1.284 | 992.75
|  |  | 10 kHz | 10 kHz | 1 MHz | 1 MHz | 2 MHz | 2 MHz | 5 MHz | 5 MHz | 10 MHz | 10 MHz | 10 kHz | 10 kHz | 1 MHz | 1 MHz | 2 MHz | 2 MHz | 5 MHz | 5 MHz | 10 MHz | 10 MHz | 5 MHz | 5 MHz | 1 MHz
CM9518 | power setting: 11.00 | 850 V | [23 empty] | 15.5 inches | yes | Nathan on 3/25/14
|  | 900 V |  |  | 720 | 715 | 700 | 701 | 652 | 655 | 495 | 499 |  |  | 26 | 25 | 25 | 25 | 24 | 23 | 22 | 23 | 1.526 | 1.549 | 803.95
|  | 1000 V |  |  | 1480 | 1470 | 1440 | 1440 | 1190 | 1200 | 675 | 667 |  |  | 45 | 44 | 45 | 46 | 47 | 47 | 44 | 45 | 1.308 | 1.31 | 897.3
|  | 1100 V |  |  | 2500 | 2490 | 2340 | 2350 | 1450 | 1410 | 722 | 701 |  |  | 69 | 68 | 68 | 67 | 66 | 68 | 69 | 68 | 1.183 | 1.117 | 990.4
CM9558 | power setting: 10.80 | 850 V | [23 empty] | 15.5 inches | yes | Nathan on 3/25/14
|  | 900 V |  |  | 783 | 787 | 772 | 776 | 733 | 737 | 545 | 544 |  |  | 28 | 28 | 26 | 26 | 25 | 25 | 28 | 27 | 1.493 | 1.461 | 807.4
|  | 1000 V |  |  | 1610 | 1610 | 1590 | 1590 | 1270 | 1270 | 682 | 670 |  |  | 46 | 46 | 44 | 44 | 46 | 45 | 45 | 45 | 1.321 | 1.286 | 900.5
|  | 1100 V |  |  | 2720 | 2750 | 2560 | 2580 | 1370 | 1220 | 648 | 624 |  |  | 70 | 69 | 70 | 70 | 70 | 186 | 65 | 70 | 1.378 | 1.4 | 993.55
CM9560 | power setting: 10.80 | 850 V | [23 empty] | 15.5 inches | yes | Nathan on 3/25/14
|  | 900 V |  |  | 700 | 704 | 695 | 697 | 665 | 667 | 515 | 520 |  |  | 25 | 25 | 26 | 25 | 24 | 24 | 25 | 25 | 1.448 | 1.51 | 806.3
|  | 1000 V |  |  | 1450 | 1460 | 1430 | 1430 | 1200 | 1210 | 675 | 700 |  |  | 46 | 46 | 44 | 45 | 46 | 47 | 43 | 44 | 1.275 | 1.28 | 899.4
|  | 1100 V |  |  | 2490 | 2520 | 2370 | 2410 | 1340 | 1460 | 576 | 690 |  |  | 68 | 70 | 71 | 72 | 70 | 71 | 102 | 66 | 1.189 | 1.219 | 992.3
CM9516 | power setting: 11.00 | 850 V | [23 empty] | 15.5 inches | yes | Nathan on 3/26/14
|  | 900 V |  |  | 712 | 722 | 702 | 711 | 655 | 662 | 509 | 514 |  |  | 25 | 25 | 28 | 28 | 27 | 27 | 25 | 25 | 1.512 | 1.521 | 806.1
|  | 1000 V |  |  | 1460 | 1490 | 1440 | 1470 | 1190 | 1200 | 652 | 680 |  |  | 45 | 45 | 45 | 46 | 47 | 48 | 44 | 44 | 1.331 | 1.293 | 899.2
|  | 1100 V |  |  | 2480 | 2540 | 2360 | 2410 | 1330 | 1450 | 555 | 688 |  |  | 67 | 68 | 72 | 70 | 68 | 65 | 98 | 70 | 1.305 | 1.285 | 992.1
|  |  | 10 kHz | 10 kHz | 1 MHz | 1 MHz | 2 MHz | 2 MHz | 5 MHz | 5 MHz | 10 MHz | 10 MHz | 10 kHz | 10 kHz | 1 MHz | 1 MHz | 2 MHz | 2 MHz | 5 MHz | 5 MHz | 10 MHz | 10 MHz | 5 MHz | 5 MHz | 1 MHz
CM9856 | power setting: 11.20 | 850 V | [23 empty] | 15.5 inches | yes | Nathan on 3/26/14
|  | 900 V |  |  | 720 | 715 | 714 | 707 | 677 | 671 | 523 | 518 |  |  | 24 | 23 | 25 | 24 | 22 | 23 | 26 | 26 | 1.48 | 1.452 | 807.45
|  | 1000 V |  |  | 1480 | 1460 | 1460 | 1440 | 1230 | 1210 | 680 | 677 |  |  | 44 | 45 | 44 | 42 | 43 | 42 | 40 | 40 | 1.275 | 1.253 | 900.6
|  | 1100 V |  |  | 2560 | 2530 | 2420 | 2390 | 1420 | 1410 | 689 | 643 |  |  | 69 | 67 | 64 | 63 | 64 | 65 | 61 | 85 | 1.198 | 1.126 | 993.7
CM9872 | power setting: 11.20 | 850 V | [23 empty] | 15.5 inches | yes | Nathan on 3/26/14
|  | 900 V |  |  | 725 | 729 | 712 | 713 | 667 | 668 | 519 | 521 |  |  | 25 | 25 | 25 | 25 | 24 | 25 | 23 | 24 | 1.448 | 1.465 | 807.75
|  | 1000 V |  |  | 1450 | 1450 | 1430 | 1430 | 1210 | 1200 | 678 | 672 |  |  | 46 | 45 | 40 | 40 | 41 | 42 | 40 | 42 | 1.26 | 1.258 | 900.6
|  | 1100 V |  |  | 2450 | 2450 | 2340 | 2340 | 1410 | 1390 | 670 | 624 |  |  | 67 | 68 | 70 | 68 | 65 | 68 | 63 | 93 | 1.21 | 1.24 | 993.6
CM9847 | power setting: 11.20 | 850 V | [23 empty] | 15.5 inches | yes | Nathan on 3/26/14
|  | 900 V |  |  | 743 | 745 | 734 | 735 | 698 | 698 | 534 | 536 |  |  | 25 | 25 | 25 | 25 | 26 | 27 | 25 | 26 | 1.473 | 1.44 | 807.95
|  | 1000 V |  |  | 1520 | 1520 | 1500 | 1500 | 1240 | 1240 | 674 | 684 |  |  | 45 | 45 | 45 | 45 | 43 | 44 | 43 | 43 | 1.3 | 1.22 | 901.05
|  | 1100 V |  |  | 2590 | 2610 | 2450 | 2470 | 1340 | 1440 | 610 | 684 |  |  | 70 | 70 | 66 | 67 | 68 | 63 | 101 | 70 | 1.329 | 1.177 | 994.2
CM9849 | power setting: 11.20 | 850 V | [23 empty] | 15.5 inches | yes | Nathan on 3/26/14
|  | 900 V |  |  | 687 | 690 | 670 | 673 | 632 | 633 | 501 | 503 |  |  | 25 | 25 | 26 | 25 | 23 | 23 | 25 | 25 | 1.582 | 1.574 | 807.95
|  | 1000 V |  |  | 1380 | 1380 | 1370 | 1370 | 1180 | 1170 | 680 | 672 |  |  | 45 | 44 | 42 | 43 | 43 | 42 | 34 | 35 | 1.385 | 1.39 | 900.85
|  | 1100 V |  |  | 2390 | 2380 | 2280 | 2270 | 1410 | 1380 | 661 | 602 |  |  | 70 | 71 | 65 | 68 | 66 | 65 | 69 | 105 | 1.345 | 1.386 | 993.6
|  |  | 10 kHz | 10 kHz | 1 MHz | 1 MHz | 2 MHz | 2 MHz | 5 MHz | 5 MHz | 10 MHz | 10 MHz | 10 kHz | 10 kHz | 1 MHz | 1 MHz | 2 MHz | 2 MHz | 5 MHz | 5 MHz | 10 MHz | 10 MHz | 5 MHz | 5 MHz | 1 MHz
CM9854 | power setting: 11.20 | 850 V | [23 empty] | 15.5 inches | yes | Nathan on 3/26/14
|  | 900 V |  |  | 676 | 684 | 670 | 676 | 635 | 642 | 502 | 504 |  |  | 25 | 25 | 25 | 24 | 22 | 23 | 24 | 24 | 1.633 | 1.645 | 808.1
|  | 1000 V |  |  | 1430 | 1440 | 1400 | 1410 | 1190 | 1190 | 688 | 670 |  |  | 46 | 47 | 43 | 44 | 42 | 43 | 40 | 41 | 1.411 | 1.415 | 901.1
|  | 1100 V |  |  | 2450 | 2480 | 2340 | 2360 | 1430 | 1390 | 677 | 618 |  |  | 67 | 68 | 63 | 64 | 67 | 70 | 68 | 97 | 1.382 | 1.419 | 994
CM9855 | power setting: 11.20 | 850 V | [23 empty] | 15.5 inches | yes | Nathan on 3/26/14
|  | 900 V |  |  | 698 | 709 | 693 | 702 | 662 | 668 | 510 | 516 |  |  | 24 | 25 | 24 | 24 | 25 | 25 | 26 | 25 | 1.495 | 1.517 | 807.2
|  | 1000 V |  |  | 1430 | 1440 | 1410 | 1420 | 1200 | 1210 | 683 | 692 |  |  | 44 | 45 | 44 | 44 | 44 | 43 | 42 | 43 | 1.335 | 1.306 | 900.1
|  | 1100 V |  |  | 2440 | 2470 | 2330 | 2360 | 1420 | 1450 | 641 | 703 |  |  | 70 | 71 | 68 | 69 | 66 | 65 | 85 | 66 | 1.297 | 1.244 | 992.8
CM9858 | power setting: 11.60 | 850 V | [23 empty] | 15.5 inches | yes | Nathan on 3/27/14
|  | 900 V |  |  | 734 | 737 | 716 | 718 | 680 | 682 | 520 | 523 |  |  | 24 | 24 | 24 | 25 | 26 | 25 | 23 | 24 | 1.501 | 1.525 | 806.95
|  | 1000 V |  |  | 1500 | 1510 | 1480 | 1480 | 1210 | 1210 | 663 | 674 |  |  | 42 | 41 | 44 | 44 | 42 | 42 | 43 | 42 | 1.332 | 1.338 | 900.05
|  | 1100 V |  |  | 2600 | 2610 | 2450 | 2460 | 1360 | 1370 | 593 | 652 |  |  | 67 | 67 | 67 | 70 | 65 | 64 | 93 | 61 | 1.333 | 1.359 | 993.35
CM9850 | power setting: 11.20 | 850 V | [23 empty] | 15.5 inches | yes | Nathan on 3/27/14
|  | 900 V |  |  | 760 | 752 | 740 | 733 | 699 | 694 | 525 | 526 |  |  | 26 | 26 | 27 | 28 | 25 | 24 | 22 | 21 | 1.476 | 1.424 | 807.9
|  | 1000 V |  |  | 1530 | 1510 | 1510 | 1490 | 1240 | 1230 | 665 | 685 |  |  | 43 | 42 | 43 | 43 | 45 | 44 | 42 | 43 | 1.306 | 1.222 | 900.5
|  | 1100 V |  |  | 2610 | 2590 | 2470 | 2460 | 1290 | 1390 | 606 | 676 |  |  | 67 | 68 | 70 | 69 | 90 | 57 | 82 | 61 | 1.399 | 1.259 | 993.45
|  |  | 10 kHz | 10 kHz | 1 MHz | 1 MHz | 2 MHz | 2 MHz | 5 MHz | 5 MHz | 10 MHz | 10 MHz | 10 kHz | 10 kHz | 1 MHz | 1 MHz | 2 MHz | 2 MHz | 5 MHz | 5 MHz | 10 MHz | 10 MHz | 5 MHz | 5 MHz | 1 MHz
CM9873 | power setting: 11.20 | 850 V | [23 empty] | 15.5 inches | yes | Nathan on 3/27/14
|  | 900 V |  |  | 716 | 717 | 704 | 706 | 667 | 667 | 510 | 510 |  |  | 24 | 24 | 27 | 26 | 22 | 23 | 23 | 22 | 1.556 | 1.516 | 806.25
|  | 1000 V |  |  | 1460 | 1460 | 1440 | 1440 | 1210 | 1200 | 679 | 677 |  |  | 45 | 46 | 43 | 43 | 42 | 41 | 43 | 43 | 1.351 | 1.34 | 898.95
|  | 1100 V |  |  | 2490 | 2490 | 2370 | 2370 | 1410 | 1400 | 684 | 656 |  |  | 68 | 69 | 68 | 67 | 64 | 65 | 64 | 67 | 1.286 | 1.251 | 992
CM9801 | power setting: 11.20 | 850 V | [23 empty] | 15.5 inches | yes | Nathan on 3/28/14
|  | 900 V |  |  | 709 | 718 | 700 | 707 | 665 | 672 | 513 | 516 |  |  | 25 | 26 | 26 | 26 | 26 | 26 | 22 | 23 | 1.506 | 1.491 | 806.6
|  | 1000 V |  |  | 1460 | 1470 | 1420 | 1430 | 1190 | 1190 | 678 | 675 |  |  | 41 | 42 | 43 | 43 | 41 | 43 | 44 | 45 | 1.3 | 1.302 | 899.45
|  | 1100 V |  |  | 2470 | 2490 | 2350 | 2360 | 1410 | 1380 | 686 | 634 |  |  | 69 | 67 | 65 | 67 | 64 | 66 | 62 | 86 | 1.201 | 1.212 | 992.55
CM9768 | power setting: 11.60 | 850 V | [23 empty] | 15.5 inches | yes | Nathan on 4/1/14
|  | 900 V |  |  | 702 | 694 | 683 | 678 | 640 | 634 | 502 | 496 |  |  | 24 | 23 | 24 | 24 | 22 | 23 | 20 | 21 | 1.625 | 1.57 | 807.4
|  | 1000 V |  |  | 1400 | 1380 | 1380 | 1360 | 1180 | 1160 | 683 | 674 |  |  | 43 | 43 | 43 | 44 | 40 | 41 | 39 | 37 | 1.452 | 1.386 | 900.2
|  | 1100 V |  |  | 2400 | 2370 | 2290 | 2260 | 1400 | 1380 | 674 | 621 |  |  | 69 | 66 | 67 | 65 | 63 | 62 | 73 | 115 | 1.48 | 1.425 | 993.05
CM9769 | power setting: 11.60 | 850 V | [23 empty] | 15.5 inches | yes | Nathan on 4/1/14
|  | 900 V |  |  | 693 | 705 | 683 | 695 | 638 | 648 | 500 | 509 |  |  | 23 | 23 | 21 | 22 | 25 | 25 | 22 | 22 | 1.487 | 1.506 | 808.2
|  | 1000 V |  |  | 1410 | 1430 | 1390 | 1410 | 1180 | 1190 | 659 | 681 |  |  | 42 | 42 | 45 | 45 | 44 | 44 | 40 | 40 | 1.286 | 1.294 | 901.15
|  | 1100 V |  |  | 2430 | 2460 | 2300 | 2330 | 1370 | 1410 | 603 | 681 |  |  | 68 | 68 | 69 | 70 | 63 | 65 | 103 | 66 | 1.133 | 1.199 | 994.1
|  |  | 10 kHz | 10 kHz | 1 MHz | 1 MHz | 2 MHz | 2 MHz | 5 MHz | 5 MHz | 10 MHz | 10 MHz | 10 kHz | 10 kHz | 1 MHz | 1 MHz | 2 MHz | 2 MHz | 5 MHz | 5 MHz | 10 MHz | 10 MHz | 5 MHz | 5 MHz | 1 MHz
CM9939 | power setting: 11.20 | 850 V | [23 empty] | 15.5 inches | yes | Nathan on 4/1/14
|  | 900 V |  |  | 746 | 749 | 725 | 726 | 683 | 685 | 532 | 530 |  |  | 25 | 25 | 23 | 24 | 23 | 23 | 23 | 23 | 1.508 | 1.44 | 807.9
|  | 1000 V |  |  | 1470 | 1470 | 1450 | 1450 | 1230 | 1220 | 686 | 672 |  |  | 45 | 45 | 43 | 44 | 41 | 41 | 40 | 39 | 1.31 | 1.267 | 900.8
|  | 1100 V |  |  | 2460 | 2470 | 2350 | 2360 | 1420 | 1370 | 695 | 614 |  |  | 70 | 70 | 69 | 70 | 61 | 60 | 63 | 91 | 1.26 | 1.232 | 993.65
CM9947 | BAD | 850 V | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | 15.5 inches | no | Nathan on 4/1/14 | No signals
|  | 900 V | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | #
|  | 1000 V | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | #
|  | 1100 V | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | #
CM9775 | power setting: 11.80 | 850 V | [23 empty] | 15.5 inches | yes | Nathan on 4/2/14
|  | 900 V |  |  | 710 | 721 | 700 | 711 | 665 | 672 | 515 | 520 |  |  | 22 | 22 | 21 | 21 | 23 | 23 | 24 | 24 | 1.57 | 1.552 | 805.6
|  | 1000 V |  |  | 1430 | 1450 | 1410 | 1430 | 1190 | 1200 | 663 | 684 |  |  | 48 | 49 | 41 | 42 | 42 | 41 | 42 | 43 | 1.394 | 1.362 | 898.5
|  | 1100 V |  |  | 2460 | 2490 | 2350 | 2380 | 1360 | 1430 | 626 | 698 |  |  | 64 | 65 | 66 | 66 | 64 | 65 | 89 | 62 | 1.351 | 1.331 | 991.45
CM9776 | power setting: 11.80 | 850 V | [23 empty] | 15.5 inches | yes | Nathan on 4/2/14
|  | 900 V |  |  | 767 | 771 | 755 | 759 | 703 | 707 | 535 | 536 |  |  | 22 | 22 | 22 | 23 | 21 | 21 | 23 | 23 | 1.504 | 1.546 | 807.5
|  | 1000 V |  |  | 1520 | 1520 | 1500 | 1500 | 1250 | 1250 | 685 | 677 |  |  | 44 | 44 | 41 | 40 | 40 | 40 | 41 | 42 | 1.301 | 1.376 | 900.55
|  | 1100 V |  |  | 2570 | 2570 | 2430 | 2430 | 1410 | 1370 | 670 | 623 |  |  | 70 | 71 | 65 | 67 | 65 | 69 | 61 | 99 | 1.222 | 1.387 | 993.4
|  |  | 10 kHz | 10 kHz | 1 MHz | 1 MHz | 2 MHz | 2 MHz | 5 MHz | 5 MHz | 10 MHz | 10 MHz | 10 kHz | 10 kHz | 1 MHz | 1 MHz | 2 MHz | 2 MHz | 5 MHz | 5 MHz | 10 MHz | 10 MHz | 5 MHz | 5 MHz | 1 MHz
CM9778 | power setting: 12.40 | 850 V | [23 empty] | 15.5 inches | yes | Nathan on 4/3/14
|  | 900 V |  |  | 699 | 696 | 690 | 687 | 650 | 646 | 512 | 508 |  |  | 23 | 22 | 20 | 20 | 21 | 22 | 20 | 20 | 1.521 | 1.506 | 806.85
|  | 1000 V |  |  | 1440 | 1430 | 1420 | 1400 | 1200 | 1180 | 680 | 665 |  |  | 44 | 43 | 42 | 41 | 38 | 39 | 34 | 35 | 1.337 | 1.352 | 900
|  | 1100 V |  |  | 2430 | 2400 | 2310 | 2280 | 1410 | 1350 | 660 | 583 |  |  | 64 | 64 | 60 | 60 | 56 | 55 | 72 | 115 | 1.32 | 1.375 | 992.8
CM9780 | power setting: 11.80 | 850 V | [23 empty] | 15.5 inches | yes | Nathan on 4/3/14
|  | 900 V |  |  | 726 | 732 | 704 | 709 | 661 | 666 | 514 | 522 |  |  | 22 | 22 | 22 | 23 | 20 | 20 | 21 | 21 | 1.533 | 1.545 | 805.85
|  | 1000 V |  |  | 1410 | 1420 | 1380 | 1390 | 1190 | 1200 | 667 | 685 |  |  | 40 | 41 | 42 | 42 | 37 | 38 | 37 | 39 | 1.348 | 1.341 | 898.6
|  | 1100 V |  |  | 2410 | 2420 | 2270 | 2280 | 1370 | 1440 | 632 | 707 |  |  | 60 | 61 | 66 | 64 | 55 | 56 | 77 | 59 | 1.333 | 1.301 | 991.5
CM9787 | power setting: 11.80 | 850 V | [23 empty] | 15.5 inches | yes | Nathan on 4/3/14
|  | 900 V |  |  | 745 | 740 | 732 | 728 | 686 | 684 | 525 | 525 |  |  | 26 | 25 | 26 | 26 | 23 | 22 | 22 | 23 | 1.558 | 1.525 | 806.95
|  | 1000 V |  |  | 1500 | 1470 | 1480 | 1460 | 1230 | 1210 | 687 | 672 |  |  | 42 | 41 | 44 | 43 | 43 | 43 | 36 | 35 | 1.361 | 1.377 | 900.05
|  | 1100 V |  |  | 2520 | 2470 | 2370 | 2320 | 1430 | 1340 | 685 | 587 |  |  | 66 | 63 | 65 | 67 | 60 | 60 | 61 | 93 | 1.301 | 1.368 | 993
CM9782 | power setting: 12.40 | 850 V | [23 empty] | 15.5 inches | yes | Nathan on 4/4/14
|  | 900 V |  |  | 732 | 725 | 720 | 711 | 675 | 671 | 521 | 520 |  |  | 22 | 21 | 20 | 20 | 23 | 22 | 21 | 21 | 1.531 | 1.508 | 806.7
|  | 1000 V |  |  | 1480 | 1460 | 1450 | 1430 | 1220 | 1210 | 687 | 680 |  |  | 45 | 45 | 42 | 42 | 37 | 36 | 41 | 40 | 1.31 | 1.32 | 899.85
|  | 1100 V |  |  | 2510 | 2480 | 2360 | 2340 | 1400 | 1370 | 689 | 620 |  |  | 65 | 63 | 63 | 62 | 55 | 55 | 59 | 87 | 1.288 | 1.296 | 992.85
|  |  | 10 kHz | 10 kHz | 1 MHz | 1 MHz | 2 MHz | 2 MHz | 5 MHz | 5 MHz | 10 MHz | 10 MHz | 10 kHz | 10 kHz | 1 MHz | 1 MHz | 2 MHz | 2 MHz | 5 MHz | 5 MHz | 10 MHz | 10 MHz | 5 MHz | 5 MHz | 1 MHz
CM9798 | power setting: 11.80 | 850 V | [23 empty] | 15.5 inches | yes | Nathan on 4/4/14
|  | 900 V |  |  | 726 | 731 | 716 | 721 | 672 | 678 | 524 | 532 |  |  | 22 | 22 | 23 | 23 | 21 | 21 | 20 | 20 | 1.401 | 1.438 | 806.9
|  | 1000 V |  |  | 1480 | 1490 | 1460 | 1460 | 1220 | 1230 | 661 | 682 |  |  | 43 | 42 | 46 | 47 | 41 | 41 | 37 | 36 | 1.235 | 1.247 | 899.9
|  | 1100 V |  |  | 2520 | 2520 | 2420 | 2420 | 1270 | 1400 | 573 | 646 |  |  | 63 | 60 | 61 | 60 | 110 | 64 | 97 | 61 | 1.263 | 1.244 | 992.6
CM9796 | power setting: 11.80 | 850 V | [23 empty] | 15.5 inches | yes | Nathan on 4/4/14
|  | 900 V |  |  | 738 | 742 | 725 | 726 | 680 | 686 | 518 | 525 |  |  | 22 | 22 | 22 | 22 | 20 | 20 | 22 | 22 | 1.464 | 1.466 | 807.35
|  | 1000 V |  |  | 1480 | 1480 | 1450 | 1450 | 1200 | 1200 | 660 | 672 |  |  | 42 | 42 | 42 | 42 | 41 | 41 | 37 | 37 | 1.287 | 1.271 | 900.35
|  | 1100 V |  |  | 2470 | 2470 | 2340 | 2340 | 1360 | 1380 | 586 | 644 |  |  | 70 | 70 | 59 | 60 | 55 | 56 | 95 | 64 | 1.22 | 1.225 | 993
CM9795 | power setting: 11.80 | 850 V | [23 empty] | 15.5 inches | yes | Nathan on 4/4/14
|  | 900 V |  |  | 764 | 754 | 750 | 742 | 705 | 700 | 520 | 521 |  |  | 23 | 23 | 25 | 25 | 25 | 24 | 26 | 26 | 1.523 | 1.499 | 804.8
|  | 1000 V |  |  | 1580 | 1550 | 1540 | 1510 | 1230 | 1230 | 667 | 675 |  |  | 44 | 44 | 41 | 40 | 45 | 45 | 40 | 40 | 1.338 | 1.271 | 898.05
|  | 1100 V |  |  | 2640 | 2600 | 2470 | 2460 | 1340 | 1380 | 664 | 680 |  |  | 64 | 63 | 66 | 65 | 54 | 55 | 61 | 60 | 1.268 | 1.222 | 991.1
CM9799 | power setting: 11.40 | 850 V | [23 empty] | 15.5 inches | yes | Nathan on 4/4/14
|  | 900 V |  |  | 749 | 750 | 730 | 732 | 688 | 689 | 530 | 526 |  |  | 23 | 23 | 24 | 23 | 22 | 22 | 23 | 23 | 1.494 | 1.455 | 807.25
|  | 1000 V |  |  | 1520 | 1510 | 1500 | 1490 | 1230 | 1220 | 682 | 665 |  |  | 42 | 41 | 43 | 43 | 40 | 39 | 40 | 40 | 1.297 | 1.293 | 900.2
|  | 1100 V |  |  | 2560 | 2550 | 2450 | 2440 | 1370 | 1190 | 652 | 584 |  |  | 70 | 71 | 67 | 65 | 61 | 155 | 63 | 98 | 1.315 | 1.395 | 993
|  |  | 10 kHz | 10 kHz | 1 MHz | 1 MHz | 2 MHz | 2 MHz | 5 MHz | 5 MHz | 10 MHz | 10 MHz | 10 kHz | 10 kHz | 1 MHz | 1 MHz | 2 MHz | 2 MHz | 5 MHz | 5 MHz | 10 MHz | 10 MHz | 5 MHz | 5 MHz | 1 MHz
CM9893 | power setting: 11.40 | 850 V | [23 empty] | 15.5 inches | yes | Nathan on 4/4/14
|  | 900 V |  |  | 710 | 713 | 690 | 694 | 655 | 658 | 505 | 510 |  |  | 23 | 24 | 23 | 23 | 25 | 25 | 23 | 23 | 1.532 | 1.508 | 806.35
|  | 1000 V |  |  | 1440 | 1440 | 1420 | 1420 | 1210 | 1210 | 670 | 673 |  |  | 44 | 44 | 45 | 44 | 43 | 42 | 42 | 43 | 1.362 | 1.303 | 899.45
|  | 1100 V |  |  | 2450 | 2450 | 2340 | 2350 | 1340 | 1370 | 590 | 659 |  |  | 67 | 65 | 64 | 65 | 67 | 68 | 102 | 62 | 1.285 | 1.222 | 992.4
CM9904 | power setting: 11.40 | 850 V | [23 empty] | 15.5 inches | yes | Nathan on 4/4/14
|  | 900 V |  |  | 794 | 799 | 772 | 777 | 720 | 724 | 535 | 535 |  |  | 27 | 26 | 25 | 26 | 23 | 23 | 22 | 22 | 1.48 | 1.506 | 806.4
|  | 1000 V |  |  | 1570 | 1570 | 1540 | 1540 | 1260 | 1240 | 683 | 667 |  |  | 42 | 42 | 41 | 40 | 43 | 43 | 42 | 42 | 1.284 | 1.382 | 899.45
|  | 1100 V |  |  | 2590 | 2570 | 2480 | 2450 | 1410 | 1320 | 658 | 572 |  |  | 68 | 67 | 62 | 63 | 55 | 55 | 55 | 85 | 1.206 | 1.338 | 992.4
CM9905 | power setting: 11.40 | 850 V | [23 empty] | 15.5 inches | yes | Nathan on 4/4/14
|  | 900 V |  |  | 784 | 783 | 770 | 770 | 722 | 723 | 539 | 543 |  |  | 26 | 26 | 24 | 24 | 25 | 25 | 23 | 23 | 1.466 | 1.505 | 807.65
|  | 1000 V |  |  | 1540 | 1540 | 1520 | 1520 | 1240 | 1240 | 672 | 680 |  |  | 44 | 44 | 40 | 42 | 44 | 45 | 40 | 40 | 1.318 | 1.311 | 900.4
|  | 1100 V |  |  | 2620 | 2620 | 2480 | 2480 | 1210 | 1370 | 591 | 658 |  |  | 70 | 70 | 67 | 66 | 155 | 72 | 82 | 65 | 1.459 | 1.385 | 993.3
CM9926 | power setting: 11.40 | 850 V | [23 empty] | 15.5 inches | yes | Nathan on 4/7/14
|  | 900 V |  |  | 794 | 802 | 790 | 797 | 740 | 745 | 538 | 547 |  |  | 25 | 26 | 22 | 22 | 25 | 26 | 24 | 24 | 1.511 | 1.506 | 807.5
|  | 1000 V |  |  | 1580 | 1600 | 1550 | 1570 | 1240 | 1260 | 675 | 706 |  |  | 42 | 42 | 45 | 46 | 45 | 45 | 40 | 41 | 1.391 | 1.333 | 900.5
|  | 1100 V |  |  | 2570 | 2630 | 2430 | 2490 | 1250 | 1430 | 555 | 673 |  |  | 71 | 70 | 64 | 66 | 76 | 61 | 104 | 68 | 1.455 | 1.363 | 993.4
|  |  | 10 kHz | 10 kHz | 1 MHz | 1 MHz | 2 MHz | 2 MHz | 5 MHz | 5 MHz | 10 MHz | 10 MHz | 10 kHz | 10 kHz | 1 MHz | 1 MHz | 2 MHz | 2 MHz | 5 MHz | 5 MHz | 10 MHz | 10 MHz | 5 MHz | 5 MHz | 1 MHz
CM9945 | power setting: 11.40 | 850 V | [23 empty] | 15.5 inches | yes | Nathan on 4/7/14
|  | 900 V |  |  | 782 | 778 | 767 | 763 | 715 | 713 | 533 | 530 |  |  | 27 | 27 | 28 | 28 | 22 | 23 | 23 | 22 | 1.53 | 1.522 | 806.9
|  | 1000 V |  |  | 1540 | 1520 | 1500 | 1490 | 1230 | 1220 | 685 | 667 |  |  | 48 | 48 | 42 | 41 | 40 | 40 | 41 | 40 | 1.356 | 1.371 | 899.7
|  | 1100 V |  |  | 2560 | 2540 | 2420 | 2410 | 1400 | 1330 | 672 | 591 |  |  | 67 | 68 | 65 | 67 | 70 | 69 | 63 | 96 | 1.318 | 1.365 | 992.7
CM9819 | power setting: 12.10 | 850 V | [23 empty] | 15.5 inches | yes | Nathan on 4/8/14
|  | 900 V |  |  | 705 | 707 | 690 | 692 | 650 | 653 | 511 | 516 |  |  | 21 | 21 | 24 | 24 | 21 | 21 | 21 | 21 | 1.42 | 1.436 | 807.6
|  | 1000 V |  |  | 1440 | 1440 | 1420 | 1420 | 1200 | 1200 | 671 | 689 |  |  | 40 | 40 | 45 | 45 | 40 | 40 | 39 | 38 | 1.241 | 1.223 | 900.55
|  | 1100 V |  |  | 2440 | 2440 | 2320 | 2320 | 1360 | 1420 | 603 | 686 |  |  | 60 | 62 | 65 | 64 | 60 | 61 | 93 | 60 | 1.076 | 1.137 | 993.35
CM9820 | power setting: 12.40 | 850 V | [23 empty] | 15.5 inches | yes | Nathan on 4/8/14
|  | 900 V |  |  | 700 | 691 | 690 | 685 | 647 | 647 | 502 | 506 |  |  | 21 | 20 | 23 | 23 | 23 | 22 | 21 | 21 | 1.555 | 1.625 | 805.55
|  | 1000 V |  |  | 1460 | 1430 | 1420 | 1400 | 1190 | 1190 | 682 | 671 |  |  | 45 | 44 | 44 | 42 | 41 | 40 | 40 | 39 | 1.356 | 1.45 | 898.8
|  | 1100 V |  |  | 2450 | 2420 | 2330 | 2310 | 1410 | 1350 | 691 | 615 |  |  | 60 | 61 | 61 | 60 | 55 | 57 | 56 | 88 | 1.313 | 1.425 | 991.65
CM9829 | power setting: 11.80 | 850 V | [23 empty] | 15.5 inches | yes | Nathan on 4/9/14
|  | 900 V |  |  | 722 | 712 | 715 | 707 | 675 | 667 | 527 | 525 |  |  | 25 | 25 | 21 | 22 | 25 | 24 | 25 | 24 | 1.534 | 1.49 | 807.35
|  | 1000 V |  |  | 1490 | 1470 | 1470 | 1450 | 1230 | 1220 | 677 | 689 |  |  | 45 | 44 | 42 | 43 | 44 | 44 | 40 | 40 | 1.361 | 1.285 | 900.6
|  | 1100 V |  |  | 2540 | 2520 | 2420 | 2400 | 1390 | 1400 | 598 | 654 |  |  | 68 | 68 | 70 | 70 | 57 | 55 | 91 | 62 | 1.272 | 1.228 | 993.6
|  |  | 10 kHz | 10 kHz | 1 MHz | 1 MHz | 2 MHz | 2 MHz | 5 MHz | 5 MHz | 10 MHz | 10 MHz | 10 kHz | 10 kHz | 1 MHz | 1 MHz | 2 MHz | 2 MHz | 5 MHz | 5 MHz | 10 MHz | 10 MHz | 5 MHz | 5 MHz | 1 MHz
CM9852 | power setting: 12.10 | 850 V | [23 empty] | 15.5 inches | yes | Nathan on 4/9/14 | current is lower than usual and has larger fluctuations
|  | 900 V |  |  | 716 | 710 | 692 | 687 | 647 | 642 | 501 | 504 |  |  | 23 | 23 | 21 | 21 | 21 | 20 | 20 | 20 | 1.52 | 1.501 | 802.8
|  | 1000 V |  |  | 1430 | 1420 | 1410 | 1390 | 1190 | 1190 | 662 | 685 |  |  | 40 | 41 | 42 | 41 | 42 | 42 | 39 | 40 | 1.346 | 1.296 | 893.65
|  | 1100 V |  |  | 2470 | 2450 | 2330 | 2330 | 1260 | 1410 | 561 | 662 |  |  | 61 | 62 | 58 | 57 | 76 | 55 | 103 | 67 | 1.434 | 1.306 | 985.95
CM9857 | power setting: 12.10 | 850 V | [23 empty] | 15.5 inches | yes | Nathan on 4/9/14
|  | 900 V |  |  | 708 | 710 | 700 | 701 | 662 | 663 | 510 | 511 |  |  | 25 | 26 | 23 | 23 | 20 | 20 | 21 | 20 | 1.516 | 1.518 | 807.4
|  | 1000 V |  |  | 1430 | 1430 | 1390 | 1390 | 1190 | 1190 | 671 | 676 |  |  | 44 | 45 | 40 | 40 | 41 | 40 | 35 | 36 | 1.365 | 1.333 | 900.25
|  | 1100 V |  |  | 2390 | 2410 | 2270 | 2290 | 1380 | 1410 | 601 | 668 |  |  | 61 | 61 | 60 | 60 | 55 | 55 | 93 | 58 | 1.286 | 1.263 | 993.05
CM9862 | power setting: 12.10 | 850 V | [23 empty] | 15.5 inches | yes | Nathan on 4/9/14
|  | 900 V |  |  | 697 | 692 | 674 | 668 | 635 | 631 | 501 | 495 |  |  | 23 | 22 | 25 | 24 | 20 | 20 | 19 | 20 | 1.515 | 1.523 | 807.4
|  | 1000 V |  |  | 1360 | 1350 | 1340 | 1330 | 1150 | 1130 | 682 | 661 |  |  | 41 | 40 | 35 | 36 | 38 | 39 | 37 | 37 | 1.332 | 1.345 | 900.4
|  | 1100 V |  |  | 2280 | 2260 | 2160 | 2140 | 1410 | 1360 | 662 | 607 |  |  | 61 | 60 | 60 | 62 | 61 | 60 | 62 | 103 | 1.286 | 1.329 | 993.1
CM9863 | power setting: 12.10 | 850 V | [23 empty] | 15.5 inches | yes | Nathan on 4/9/14
|  | 900 V |  |  | 702 | 705 | 685 | 688 | 638 | 639 | 503 | 505 |  |  | 21 | 21 | 21 | 20 | 19 | 20 | 23 | 23 | 1.533 | 1.517 | 805.95
|  | 1000 V |  |  | 1420 | 1420 | 1390 | 1390 | 1170 | 1170 | 685 | 681 |  |  | 45 | 45 | 42 | 41 | 35 | 37 | 39 | 40 | 1.341 | 1.326 | 899.15
|  | 1100 V |  |  | 2350 | 2350 | 2220 | 2220 | 1410 | 1360 | 705 | 691 |  |  | 65 | 62 | 55 | 55 | 55 | 57 | 52 | 53 | 1.281 | 1.308 | 991.95
|  |  | 10 kHz | 10 kHz | 1 MHz | 1 MHz | 2 MHz | 2 MHz | 5 MHz | 5 MHz | 10 MHz | 10 MHz | 10 kHz | 10 kHz | 1 MHz | 1 MHz | 2 MHz | 2 MHz | 5 MHz | 5 MHz | 10 MHz | 10 MHz | 5 MHz | 5 MHz | 1 MHz
CM9864 | power setting: 12.40 | 850 V | [23 empty] | 15.5 inches | yes | Nathan on 4/10/14
|  | 900 V |  |  | 693 | 695 | 675 | 678 | 637 | 639 | 501 | 507 |  |  | 20 | 20 | 22 | 21 | 21 | 21 | 20 | 20 | 1.451 | 1.486 | 807.1
|  | 1000 V |  |  | 1400 | 1400 | 1380 | 1380 | 1180 | 1180 | 662 | 694 |  |  | 40 | 40 | 42 | 42 | 38 | 37 | 35 | 36 | 1.278 | 1.26 | 900.25
|  | 1100 V |  |  | 2350 | 2360 | 2250 | 2270 | 1330 | 1450 | 654 | 731 |  |  | 58 | 58 | 60 | 61 | 52 | 55 | 62 | 61 | 1.22 | 1.207 | 993.2
CM9821 | power setting: 12.40 | 850 V | [23 empty] | 15.5 inches | yes | Nathan on 4/10/14
|  | 900 V |  |  | 768 | 770 | 747 | 749 | 708 | 710 | 531 | 535 |  |  | 22 | 22 | 21 | 21 | 22 | 22 | 20 | 19 | 1.531 | 1.511 | 806.9
|  | 1000 V |  |  | 1560 | 1560 | 1520 | 1520 | 1250 | 1250 | 670 | 676 |  |  | 43 | 44 | 44 | 44 | 40 | 40 | 37 | 38 | 1.351 | 1.262 | 900
|  | 1100 V |  |  | 2600 | 2630 | 2460 | 2480 | 1400 | 1410 | 623 | 671 |  |  | 56 | 55 | 64 | 65 | 60 | 60 | 87 | 56 | 1.273 | 1.175 | 992.75
CM9817 | power setting: 12.40 | 850 V | [23 empty] | 15.5 inches | yes | Nathan on 4/10/14
|  | 900 V |  |  | 725 | 734 | 707 | 714 | 671 | 674 | 513 | 512 |  |  | 20 | 21 | 22 | 22 | 20 | 20 | 18 | 19 | 1.528 | 1.506 | 806.3
|  | 1000 V |  |  | 1500 | 1510 | 1470 | 1480 | 1210 | 1210 | 657 | 663 |  |  | 42 | 43 | 36 | 35 | 38 | 40 | 38 | 39 | 1.356 | 1.308 | 899.6
|  | 1100 V |  |  | 2540 | 2580 | 2400 | 2430 | 1370 | 1400 | 610 | 670 |  |  | 61 | 60 | 59 | 58 | 51 | 50 | 87 | 55 | 1.303 | 1.233 | 992.6
CM9812 | power setting: 12.40 | 850 V | [23 empty] | 15.5 inches | yes | Nathan on 4/10/14
|  | 900 V |  |  | 770 | 770 | 759 | 757 | 718 | 718 | 533 | 533 |  |  | 24 | 24 | 22 | 22 | 22 | 23 | 20 | 20 | 1.532 | 1.569 | 807.9
|  | 1000 V |  |  | 1560 | 1550 | 1530 | 1530 | 1240 | 1240 | 678 | 684 |  |  | 44 | 44 | 40 | 40 | 35 | 35 | 36 | 38 | 1.388 | 1.39 | 900.9
|  | 1100 V |  |  | 2620 | 2630 | 2440 | 2450 | 1330 | 1390 | 606 | 661 |  |  | 61 | 59 | 56 | 55 | 85 | 61 | 93 | 61 | 1.432 | 1.385 | 993.75
|  |  | 10 kHz | 10 kHz | 1 MHz | 1 MHz | 2 MHz | 2 MHz | 5 MHz | 5 MHz | 10 MHz | 10 MHz | 10 kHz | 10 kHz | 1 MHz | 1 MHz | 2 MHz | 2 MHz | 5 MHz | 5 MHz | 10 MHz | 10 MHz | 5 MHz | 5 MHz | 1 MHz
CM9838 | power setting: 12.40 | 850 V | [23 empty] | 15.5 inches | yes | Nathan on 4/10/14
|  | 900 V |  |  | 770 | 774 | 750 | 750 | 708 | 710 | 533 | 532 |  |  | 20 | 20 | 23 | 23 | 22 | 22 | 22 | 22 | 1.482 | 1.486 | 806.65
|  | 1000 V |  |  | 1550 | 1540 | 1530 | 1520 | 1250 | 1240 | 690 | 680 |  |  | 41 | 40 | 41 | 40 | 39 | 38 | 41 | 41 | 1.299 | 1.324 | 899.7
|  | 1100 V |  |  | 2650 | 2630 | 2490 | 2470 | 1430 | 1390 | 671 | 615 |  |  | 64 | 62 | 57 | 56 | 60 | 59 | 63 | 96 | 1.221 | 1.256 | 992.45
CM9818 | power setting: 12.40 | 850 V | [23 empty] | 15.5 inches | yes | Nathan on 4/11/14
|  | 900 V |  |  | 764 | 772 | 747 | 752 | 703 | 705 | 535 | 533 |  |  | 22 | 22 | 24 | 24 | 22 | 22 | 21 | 21 | 1.521 | 1.493 | 806.55
|  | 1000 V |  |  | 1570 | 1590 | 1530 | 1540 | 1260 | 1250 | 686 | 680 |  |  | 40 | 40 | 42 | 42 | 38 | 39 | 35 | 34 | 1.315 | 1.316 | 899.8
|  | 1100 V |  |  | 2620 | 2640 | 2470 | 2470 | 1410 | 1380 | 662 | 624 |  |  | 65 | 63 | 60 | 62 | 55 | 55 | 51 | 83 | 1.254 | 1.237 | 992.65
CM9811 | power setting: 11.80 | 850 V | [23 empty] | 15.5 inches | yes | Nathan on 4/11/14
|  | 900 V |  |  | 708 | 710 | 705 | 709 | 664 | 666 | 523 | 525 |  |  | 24 | 24 | 21 | 21 | 23 | 22 | 24 | 25 | 1.491 | 1.454 | 806.75
|  | 1000 V |  |  | 1460 | 1460 | 1430 | 1430 | 1210 | 1210 | 675 | 677 |  |  | 40 | 40 | 41 | 40 | 39 | 38 | 38 | 37 | 1.295 | 1.274 | 899.85
|  | 1100 V |  |  | 2480 | 2490 | 2350 | 2360 | 1380 | 1390 | 652 | 599 |  |  | 70 | 71 | 60 | 61 | 57 | 58 | 67 | 100 | 1.218 | 1.204 | 992.7
CM9802 | power setting: 11.80 | 850 V | [23 empty] | 15.5 inches | yes | Nathan on 4/11/14
|  | 900 V |  |  | 790 | 785 | 779 | 772 | 730 | 724 | 541 | 542 |  |  | 25 | 25 | 23 | 23 | 25 | 25 | 24 | 23 | 1.481 | 1.499 | 806.7
|  | 1000 V |  |  | 1590 | 1570 | 1570 | 1550 | 1240 | 1240 | 663 | 684 |  |  | 44 | 43 | 45 | 45 | 41 | 41 | 37 | 36 | 1.313 | 1.29 | 899.75
|  | 1100 V |  |  | 2640 | 2620 | 2500 | 2480 | 1360 | 1430 | 599 | 676 |  |  | 70 | 67 | 62 | 62 | 65 | 62 | 102 | 64 | 1.169 | 1.236 | 992.55
|  |  | 10 kHz | 10 kHz | 1 MHz | 1 MHz | 2 MHz | 2 MHz | 5 MHz | 5 MHz | 10 MHz | 10 MHz | 10 kHz | 10 kHz | 1 MHz | 1 MHz | 2 MHz | 2 MHz | 5 MHz | 5 MHz | 10 MHz | 10 MHz | 5 MHz | 5 MHz | 1 MHz
CM9804 | power setting: 12.40 | 850 V | [23 empty] | 15.5 inches | yes | Nathan on 4/11/14
|  | 900 V |  |  | 775 | 778 | 769 | 770 | 721 | 722 | 536 | 533 |  |  | 22 | 23 | 22 | 22 | 23 | 23 | 24 | 23 | 1.519 | 1.472 | 806.45
|  | 1000 V |  |  | 1550 | 1550 | 1520 | 1520 | 1250 | 1240 | 683 | 668 |  |  | 44 | 44 | 37 | 38 | 43 | 43 | 40 | 39 | 1.323 | 1.292 | 899.6
|  | 1100 V |  |  | 2590 | 2610 | 2460 | 2470 | 1410 | 1330 | 664 | 603 |  |  | 62 | 62 | 63 | 64 | 59 | 60 | 56 | 98 | 1.253 | 1.225 | 992.65
CM9805 | power setting: 11.80 | 850 V | [23 empty] | 15.5 inches | yes | Nathan on 4/11/14
|  | 900 V |  |  | 748 | 748 | 741 | 741 | 704 | 703 | 535 | 530 |  |  | 25 | 25 | 22 | 22 | 25 | 25 | 25 | 25 | 1.575 | 1.548 | 805.25
|  | 1000 V |  |  | 1570 | 1560 | 1550 | 1540 | 1260 | 1240 | 690 | 675 |  |  | 44 | 44 | 40 | 41 | 42 | 42 | 40 | 38 | 1.358 | 1.369 | 898.4
|  | 1100 V |  |  | 2670 | 2660 | 2520 | 2500 | 1410 | 1340 | 659 | 585 |  |  | 66 | 66 | 65 | 67 | 62 | 60 | 57 | 94 | 1.328 | 1.381 | 991.45
CM9903 | power setting: 11.80 | 850 V | [23 empty] | 15.5 inches | yes | Nathan on 4/11/14
|  | 900 V |  |  | 780 | 790 | 755 | 763 | 707 | 715 | 530 | 532 |  |  | 25 | 25 | 25 | 25 | 24 | 25 | 25 | 25 | 1.472 | 1.463 | 807.8
|  | 1000 V |  |  | 1520 | 1540 | 1510 | 1520 | 1240 | 1240 | 685 | 670 |  |  | 44 | 44 | 42 | 42 | 45 | 45 | 37 | 37 | 1.275 | 1.313 | 900.7
|  | 1100 V |  |  | 2570 | 2590 | 2440 | 2450 | 1410 | 1290 | 674 | 605 |  |  | 70 | 70 | 66 | 65 | 62 | 87 | 55 | 93 | 1.291 | 1.341 | 993.6
CN6899 | power setting: 11.40 | 850 V | [23 empty] | 15.5 inches | yes | Nathan on 4/11/14
|  | 900 V |  |  | 745 | 745 | 738 | 738 | 692 | 691 | 534 | 530 |  |  | 22 | 22 | 23 | 23 | 23 | 23 | 20 | 20 | 1.53 | 1.539 | 807.35
|  | 1000 V |  |  | 1490 | 1480 | 1460 | 1450 | 1230 | 1220 | 680 | 662 |  |  | 43 | 42 | 40 | 40 | 42 | 42 | 41 | 40 | 1.32 | 1.341 | 900.3
|  | 1100 V |  |  | 2480 | 2470 | 2380 | 2370 | 1430 | 1360 | 671 | 601 |  |  | 61 | 60 | 62 | 63 | 64 | 66 | 65 | 105 | 1.275 | 1.343 | 993.2
|  |  | 10 kHz | 10 kHz | 1 MHz | 1 MHz | 2 MHz | 2 MHz | 5 MHz | 5 MHz | 10 MHz | 10 MHz | 10 kHz | 10 kHz | 1 MHz | 1 MHz | 2 MHz | 2 MHz | 5 MHz | 5 MHz | 10 MHz | 10 MHz | 5 MHz | 5 MHz | 1 MHz
CN6896 | power setting: 11.40 | 850 V | [23 empty] | 15.5 inches | yes | Nathan on 4/11/14
|  | 900 V |  |  | 728 | 730 | 720 | 723 | 678 | 681 | 520 | 518 |  |  | 22 | 22 | 21 | 21 | 24 | 24 | 22 | 22 | 1.599 | 1.574 | 807.3
|  | 1000 V |  |  | 1440 | 1440 | 1420 | 1420 | 1210 | 1210 | 699 | 681 |  |  | 44 | 43 | 44 | 44 | 40 | 40 | 42 | 43 | 1.401 | 1.382 | 900.3
|  | 1100 V |  |  | 2460 | 2470 | 2350 | 2350 | 1440 | 1380 | 672 | 596 |  |  | 70 | 70 | 62 | 60 | 56 | 55 | 72 | 113 | 1.339 | 1.366 | 993.1
CN6898 | power setting: 11.60 | 850 V | [23 empty] | 15.5 inches | yes | Nathan on 4/15/14
|  | 900 V |  |  | 738 | 730 | 725 | 718 | 690 | 686 | 527 | 528 |  |  | 23 | 23 | 20 | 20 | 22 | 21 | 23 | 23 | 1.612 | 1.594 | 806.75
|  | 1000 V |  |  | 1500 | 1480 | 1470 | 1450 | 1220 | 1220 | 668 | 685 |  |  | 40 | 41 | 42 | 41 | 43 | 42 | 42 | 42 | 1.437 | 1.398 | 899.9
|  | 1100 V |  |  | 2530 | 2500 | 2400 | 2380 | 1370 | 1420 | 577 | 650 |  |  | 60 | 61 | 68 | 67 | 67 | 65 | 109 | 69 | 1.431 | 1.351 | 992.75
CM9750 | power setting: 11.40 | 850 V | [23 empty] | 15.5 inches | yes | Nathan on 4/16/14
|  | 900 V |  |  | 792 | 789 | 771 | 768 | 724 | 722 | 535 | 539 |  |  | 25 | 25 | 24 | 24 | 25 | 24 | 21 | 21 | 1.441 | 1.479 | 807.8
|  | 1000 V |  |  | 1590 | 1580 | 1560 | 1550 | 1250 | 1250 | 676 | 693 |  |  | 40 | 40 | 45 | 44 | 40 | 40 | 37 | 38 | 1.285 | 1.302 | 900.95
|  | 1100 V |  |  | 2670 | 2660 | 2500 | 2500 | 1260 | 1430 | 585 | 654 |  |  | 65 | 66 | 64 | 65 | 128 | 61 | 95 | 58 | 1.356 | 1.325 | 993.9
CM9754 | power setting: 11.20 | 850 V | [23 empty] | 15.5 inches | yes | Nathan on 4/16/14
|  | 900 V |  |  | 756 | 752 | 738 | 735 | 695 | 694 | 523 | 528 |  |  | 25 | 25 | 25 | 25 | 25 | 24 | 21 | 21 | 1.488 | 1.479 | 807
|  | 1000 V |  |  | 1500 | 1490 | 1480 | 1470 | 1220 | 1220 | 664 | 684 |  |  | 42 | 43 | 40 | 40 | 42 | 41 | 39 | 40 | 1.33 | 1.276 | 900.05
|  | 1100 V |  |  | 2540 | 2540 | 2430 | 2430 | 1310 | 1390 | 609 | 675 |  |  | 70 | 68 | 67 | 68 | 83 | 62 | 95 | 63 | 1.395 | 1.279 | 992.8
|  |  | 10 kHz | 10 kHz | 1 MHz | 1 MHz | 2 MHz | 2 MHz | 5 MHz | 5 MHz | 10 MHz | 10 MHz | 10 kHz | 10 kHz | 1 MHz | 1 MHz | 2 MHz | 2 MHz | 5 MHz | 5 MHz | 10 MHz | 10 MHz | 5 MHz | 5 MHz | 1 MHz
CM9760 | power setting: 11.20 | 850 V | [23 empty] | 15.5 inches | yes | Nathan on 4/16/14
|  | 900 V |  |  | 780 | 790 | 766 | 775 | 717 | 727 | 535 | 543 |  |  | 24 | 25 | 27 | 27 | 24 | 25 | 25 | 25 | 1.511 | 1.542 | 807.05
|  | 1000 V |  |  | 1570 | 1580 | 1540 | 1550 | 1240 | 1250 | 686 | 693 |  |  | 41 | 41 | 44 | 43 | 43 | 43 | 40 | 40 | 1.36 | 1.373 | 900.1
|  | 1100 V |  |  | 2680 | 2680 | 2500 | 2510 | 1250 | 1390 | 620 | 663 |  |  | 69 | 66 | 62 | 64 | 153 | 61 | 95 | 56 | 1.5 | 1.422 | 992.9
CM9764 | power setting: 11.20 | 850 V | [23 empty] | 15.5 inches | yes | Nathan on 4/16/14
|  | 900 V |  |  | 734 | 737 | 722 | 725 | 682 | 686 | 525 | 531 |  |  | 25 | 25 | 23 | 23 | 24 | 25 | 22 | 22 | 1.488 | 1.536 | 807.8
|  | 1000 V |  |  | 1500 | 1500 | 1470 | 1470 | 1230 | 1240 | 680 | 695 |  |  | 43 | 43 | 41 | 41 | 45 | 45 | 39 | 40 | 1.312 | 1.335 | 900.85
|  | 1100 V |  |  | 2560 | 2550 | 2420 | 2420 | 1390 | 1450 | 604 | 681 |  |  | 65 | 64 | 64 | 66 | 63 | 61 | 105 | 64 | 1.222 | 1.296 | 993.85
CM9763 | power setting: 11.20 | 850 V | [23 empty] | 15.5 inches | yes | Nathan on 4/16/14
|  | 900 V |  |  | 706 | 710 | 686 | 691 | 648 | 653 | 502 | 507 |  |  | 25 | 25 | 22 | 23 | 25 | 25 | 20 | 20 | 1.545 | 1.567 | 806.7
|  | 1000 V |  |  | 1410 | 1420 | 1400 | 1400 | 1180 | 1180 | 671 | 689 |  |  | 43 | 43 | 44 | 43 | 41 | 42 | 40 | 40 | 1.361 | 1.375 | 899.7
|  | 1100 V |  |  | 2390 | 2410 | 2290 | 2310 | 1390 | 1430 | 623 | 681 |  |  | 64 | 64 | 62 | 61 | 60 | 59 | 92 | 61 | 1.349 | 1.39 | 992.5
CM9765 | power setting: 11.20 | 850 V | [23 empty] | 15.5 inches | yes | Nathan on 4/17/14
|  | 900 V |  |  | 735 | 736 | 729 | 729 | 688 | 689 | 519 | 520 |  |  | 25 | 25 | 24 | 24 | 22 | 21 | 25 | 25 | 1.452 | 1.47 | 806.8
|  | 1000 V |  |  | 1510 | 1510 | 1480 | 1480 | 1220 | 1220 | 655 | 674 |  |  | 42 | 42 | 44 | 43 | 39 | 40 | 40 | 40 | 1.268 | 1.261 | 899.9
|  | 1100 V |  |  | 2560 | 2560 | 2420 | 2420 | 1350 | 1410 | 589 | 662 |  |  | 65 | 64 | 62 | 61 | 62 | 64 | 88 | 63 | 1.081 | 1.198 | 992.8
|  |  | 10 kHz | 10 kHz | 1 MHz | 1 MHz | 2 MHz | 2 MHz | 5 MHz | 5 MHz | 10 MHz | 10 MHz | 10 kHz | 10 kHz | 1 MHz | 1 MHz | 2 MHz | 2 MHz | 5 MHz | 5 MHz | 10 MHz | 10 MHz | 5 MHz | 5 MHz | 1 MHz
CM9762 | power setting: 11.20 | 850 V | [23 empty] | 15.5 inches | yes | Nathan on 4/17/14
|  | 900 V |  |  | 706 | 707 | 692 | 693 | 645 | 647 | 503 | 510 |  |  | 21 | 21 | 24 | 24 | 25 | 25 | 21 | 21 | 1.655 | 1.64 | 805.9
|  | 1000 V |  |  | 1430 | 1430 | 1420 | 1420 | 1180 | 1190 | 655 | 687 |  |  | 43 | 43 | 40 | 41 | 38 | 39 | 37 | 38 | 1.473 | 1.416 | 898.85
|  | 1100 V |  |  | 2440 | 2470 | 2330 | 2370 | 1320 | 1460 | 579 | 688 |  |  | 64 | 66 | 62 | 64 | 61 | 61 | 101 | 63 | 1.467 | 1.43 | 991.6
CM9794 | power setting: 11.40 | 850 V | [23 empty] | 15.5 inches | yes | Nathan on 4/17/14
|  | 900 V |  |  | 742 | 744 | 732 | 736 | 690 | 693 | 520 | 524 |  |  | 23 | 23 | 24 | 24 | 25 | 25 | 20 | 20 | 1.552 | 1.555 | 805.45
|  | 1000 V |  |  | 1540 | 1540 | 1510 | 1510 | 1220 | 1220 | 661 | 672 |  |  | 44 | 43 | 44 | 44 | 41 | 41 | 40 | 41 | 1.335 | 1.322 | 898.8
|  | 1100 V |  |  | 2620 | 2630 | 2450 | 2470 | 1340 | 1390 | 636 | 664 |  |  | 63 | 63 | 67 | 68 | 66 | 64 | 68 | 65 | 1.238 | 1.257 | 991.9
CM9813 | power setting: 11.40 | 850 V | [23 empty] | 15.5 inches | yes | Nathan on 4/17/14
|  | 900 V |  |  | 750 | 765 | 731 | 745 | 693 | 702 | 520 | 524 |  |  | 25 | 25 | 25 | 25 | 24 | 24 | 23 | 23 | 1.525 | 1.521 | 805.9
|  | 1000 V |  |  | 1540 | 1570 | 1510 | 1530 | 1230 | 1240 | 661 | 673 |  |  | 43 | 43 | 44 | 44 | 40 | 40 | 39 | 40 | 1.342 | 1.301 | 899.4
|  | 1100 V |  |  | 2600 | 2640 | 2440 | 2470 | 1360 | 1390 | 634 | 658 |  |  | 65 | 66 | 64 | 63 | 61 | 61 | 64 | 60 | 1.27 | 1.248 | 992.55
CM9810 | power setting: 11.40 | 850 V | [23 empty] | 15.5 inches | yes | Nathan on 4/23/14
|  | 900 V |  |  | 780 | 786 | 770 | 777 | 724 | 728 | 541 | 543 |  |  | 22 | 22 | 25 | 25 | 26 | 25 | 24 | 24 | 1.477 | 1.471 | 807.9
|  | 1000 V |  |  | 1590 | 1590 | 1550 | 1550 | 1260 | 1250 | 685 | 682 |  |  | 41 | 41 | 43 | 43 | 44 | 43 | 45 | 45 | 1.279 | 1.328 | 900.9
|  | 1100 V |  |  | 2680 | 2680 | 2520 | 2520 | 1380 | 1200 | 663 | 601 |  |  | 70 | 71 | 65 | 67 | 68 | 175 | 68 | 105 | 1.326 | 1.494 | 994.05
|  |  | 10 kHz | 10 kHz | 1 MHz | 1 MHz | 2 MHz | 2 MHz | 5 MHz | 5 MHz | 10 MHz | 10 MHz | 10 kHz | 10 kHz | 1 MHz | 1 MHz | 2 MHz | 2 MHz | 5 MHz | 5 MHz | 10 MHz | 10 MHz | 5 MHz | 5 MHz | 1 MHz
CM9809 | power setting: 11.40 | 850 V | [23 empty] | 15.5 inches | yes | Nathan on 4/23/14
|  | 900 V |  |  | 740 | 743 | 728 | 730 | 687 | 688 | 531 | 530 |  |  | 24 | 24 | 25 | 24 | 24 | 24 | 22 | 22 | 1.46 | 1.448 | 807.2
|  | 1000 V |  |  | 1500 | 1500 | 1470 | 1470 | 1240 | 1230 | 686 | 676 |  |  | 41 | 40 | 44 | 44 | 43 | 44 | 41 | 40 | 1.266 | 1.289 | 900.05
|  | 1100 V |  |  | 2570 | 2570 | 2440 | 2440 | 1390 | 1230 | 661 | 592 |  |  | 69 | 70 | 65 | 65 | 66 | 135 | 64 | 91 | 1.287 | 1.408 | 992.85
CM9807 | power setting: 11.20 | 850 V | [23 empty] | 15.5 inches | yes | Nathan on 4/23/14
|  | 900 V |  |  | 757 | 749 | 748 | 741 | 702 | 698 | 531 | 530 |  |  | 24 | 24 | 25 | 25 | 24 | 23 | 24 | 24 | 1.524 | 1.542 | 804.95
|  | 1000 V |  |  | 1590 | 1570 | 1550 | 1530 | 1250 | 1240 | 684 | 672 |  |  | 43 | 43 | 41 | 40 | 43 | 44 | 40 | 40 | 1.315 | 1.363 | 898.25
|  | 1100 V |  |  | 2700 | 2680 | 2550 | 2530 | 1400 | 1360 | 654 | 595 |  |  | 69 | 67 | 68 | 67 | 65 | 65 | 64 | 90 | 1.27 | 1.331 | 991.2
CM9806 | BAD | 850 V | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | 15.5 inches | no | Nathan on 4/23/14 | No signals
|  | 900 V | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | #
|  | 1000 V | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | #
|  | 1100 V | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | #
CM9797 | power setting: 11.40 | 850 V | [23 empty] | 15.5 inches | yes | Nathan on 4/23/14 | signals differ more than usual at 1, 2 MHz
|  | 900 V |  |  | 718 | 736 | 714 | 727 | 671 | 677 | 516 | 517 |  |  | 25 | 25 | 26 | 26 | 21 | 21 | 23 | 24 | 1.503 | 1.515 | 805.35
|  | 1000 V |  |  | 1480 | 1530 | 1470 | 1500 | 1210 | 1220 | 646 | 675 |  |  | 44 | 46 | 43 | 45 | 41 | 42 | 40 | 39 | 1.32 | 1.274 | 898.65
|  | 1100 V |  |  | 2530 | 2620 | 2390 | 2450 | 1280 | 1390 | 564 | 673 |  |  | 68 | 70 | 69 | 70 | 64 | 65 | 92 | 68 | 1.223 | 1.194 | 991.65
|  |  | 10 kHz | 10 kHz | 1 MHz | 1 MHz | 2 MHz | 2 MHz | 5 MHz | 5 MHz | 10 MHz | 10 MHz | 10 kHz | 10 kHz | 1 MHz | 1 MHz | 2 MHz | 2 MHz | 5 MHz | 5 MHz | 10 MHz | 10 MHz | 5 MHz | 5 MHz | 1 MHz
CM9792 | power setting: 11.40 | 850 V | [23 empty] | 15.5 inches | yes | Nathan on 4/23/14
|  | 900 V |  |  | 770 | 777 | 765 | 772 | 724 | 731 | 536 | 543 |  |  | 25 | 25 | 22 | 23 | 24 | 24 | 23 | 23 | 1.41 | 1.471 | 807.5
|  | 1000 V |  |  | 1570 | 1570 | 1550 | 1550 | 1250 | 1250 | 638 | 671 |  |  | 40 | 40 | 43 | 43 | 41 | 42 | 52 | 47 | 1.261 | 1.308 | 900.35
|  | 1100 V |  |  | 2660 | 2660 | 2520 | 2520 | 1150 | 1320 | 580 | 587 |  |  | 68 | 65 | 65 | 66 | 193 | 68 | 60 | 61 | 1.388 | 1.408 | 993.1
CM9781 | power setting: 11.40 | 850 V | [23 empty] | 15.5 inches | yes | Nathan on 4/24/14
|  | 900 V |  |  | 720 | 735 | 703 | 717 | 669 | 676 | 515 | 521 |  |  | 25 | 25 | 24 | 24 | 22 | 23 | 24 | 25 | 1.573 | 1.541 | 805.55
|  | 1000 V |  |  | 1490 | 1520 | 1470 | 1490 | 1210 | 1220 | 668 | 680 |  |  | 44 | 44 | 39 | 40 | 42 | 43 | 38 | 38 | 1.361 | 1.316 | 898.8
|  | 1100 V |  |  | 2550 | 2590 | 2400 | 2430 | 1360 | 1400 | 623 | 679 |  |  | 64 | 64 | 66 | 65 | 62 | 62 | 74 | 65 | 1.315 | 1.223 | 991.9
CM9770 | power setting: 11.40 | 850 V | [23 empty] | 15.5 inches | yes | Nathan on 4/24/14
|  | 900 V |  |  | 738 | 736 | 728 | 725 | 679 | 677 | 524 | 523 |  |  | 21 | 21 | 23 | 23 | 25 | 25 | 23 | 23 | 1.536 | 1.511 | 807.15
|  | 1000 V |  |  | 1480 | 1470 | 1450 | 1440 | 1220 | 1220 | 665 | 677 |  |  | 42 | 41 | 42 | 42 | 40 | 40 | 37 | 37 | 1.353 | 1.314 | 900.15
|  | 1100 V |  |  | 2490 | 2480 | 2360 | 2350 | 1370 | 1410 | 626 | 690 |  |  | 62 | 61 | 65 | 66 | 60 | 61 | 86 | 58 | 1.284 | 1.222 | 993
CM9831 | power setting: 11.40 | 850 V | [23 empty] | 15.5 inches | yes | Nathan on 4/24/14
|  | 900 V |  |  | 775 | 766 | 760 | 751 | 703 | 697 | 532 | 528 |  |  | 23 | 23 | 22 | 22 | 24 | 24 | 23 | 23 | 1.482 | 1.499 | 806.35
|  | 1000 V |  |  | 1560 | 1540 | 1530 | 1510 | 1240 | 1230 | 694 | 675 |  |  | 42 | 42 | 42 | 41 | 38 | 39 | 40 | 39 | 1.293 | 1.336 | 899.45
|  | 1100 V |  |  | 2650 | 2610 | 2510 | 2480 | 1400 | 1370 | 667 | 601 |  |  | 65 | 63 | 61 | 60 | 63 | 66 | 64 | 97 | 1.243 | 1.315 | 992.45
|  |  | 10 kHz | 10 kHz | 1 MHz | 1 MHz | 2 MHz | 2 MHz | 5 MHz | 5 MHz | 10 MHz | 10 MHz | 10 kHz | 10 kHz | 1 MHz | 1 MHz | 2 MHz | 2 MHz | 5 MHz | 5 MHz | 10 MHz | 10 MHz | 5 MHz | 5 MHz | 1 MHz
CM9767 | power setting: 11.40 | 850 V | [23 empty] | 15.5 inches | yes | Nathan on 4/24/14
|  | 900 V |  |  | 773 | 767 | 758 | 752 | 702 | 700 | 527 | 528 |  |  | 21 | 21 | 25 | 25 | 25 | 25 | 23 | 23 | 1.515 | 1.53 | 805.4
|  | 1000 V |  |  | 1560 | 1540 | 1540 | 1520 | 1240 | 1240 | 675 | 682 |  |  | 42 | 43 | 42 | 42 | 44 | 44 | 40 | 41 | 1.341 | 1.318 | 898.35
|  | 1100 V |  |  | 2650 | 2630 | 2490 | 2480 | 1400 | 1410 | 612 | 655 |  |  | 65 | 64 | 65 | 67 | 61 | 60 | 89 | 59 | 1.297 | 1.312 | 991.1
CM9766 | power setting: 11.40 | 850 V | [23 empty] | 15.5 inches | yes | Nathan on 4/24/14
|  | 900 V |  |  | 765 | 765 | 745 | 745 | 700 | 700 | 527 | 530 |  |  | 26 | 25 | 24 | 24 | 25 | 25 | 22 | 22 | 1.53 | 1.519 | 807.4
|  | 1000 V |  |  | 1530 | 1530 | 1500 | 1500 | 1230 | 1230 | 660 | 676 |  |  | 44 | 45 | 43 | 43 | 42 | 42 | 35 | 36 | 1.345 | 1.321 | 900.3
|  | 1100 V |  |  | 2590 | 2590 | 2460 | 2460 | 1350 | 1390 | 598 | 655 |  |  | 65 | 64 | 67 | 64 | 65 | 60 | 97 | 61 | 1.311 | 1.293 | 993.25
CM9742 | power setting: 11.60 | 850 V | [23 empty] | 15.5 inches | yes | Nathan on 4/24/14
|  | 900 V |  |  | 710 | 703 | 702 | 699 | 649 | 649 | 504 | 507 |  |  | 25 | 25 | 23 | 23 | 24 | 24 | 23 | 23 | 1.593 | 1.599 | 805.45
|  | 1000 V |  |  | 1480 | 1460 | 1450 | 1440 | 1200 | 1200 | 673 | 677 |  |  | 42 | 41 | 45 | 45 | 42 | 42 | 40 | 41 | 1.397 | 1.366 | 898.7
|  | 1100 V |  |  | 2500 | 2500 | 2330 | 2340 | 1380 | 1380 | 682 | 685 |  |  | 63 | 62 | 66 | 66 | 65 | 65 | 55 | 56 | 1.371 | 1.337 | 991.75
CM9745 | power setting: 11.60 | 850 V | [23 empty] | 15.5 inches | yes | Nathan on 4/24/14
|  | 900 V |  |  | 786 | 794 | 770 | 778 | 712 | 719 | 526 | 530 |  |  | 25 | 25 | 22 | 22 | 23 | 23 | 23 | 22 | 1.565 | 1.558 | 805.85
|  | 1000 V |  |  | 1560 | 1580 | 1520 | 1540 | 1230 | 1240 | 675 | 668 |  |  | 40 | 40 | 41 | 41 | 41 | 41 | 44 | 44 | 1.384 | 1.371 | 898.9
|  | 1100 V |  |  | 2600 | 2640 | 2460 | 2500 | 1340 | 1300 | 645 | 587 |  |  | 60 | 60 | 64 | 66 | 64 | 94 | 66 | 101 | 1.396 | 1.444 | 991.7
|  |  | 10 kHz | 10 kHz | 1 MHz | 1 MHz | 2 MHz | 2 MHz | 5 MHz | 5 MHz | 10 MHz | 10 MHz | 10 kHz | 10 kHz | 1 MHz | 1 MHz | 2 MHz | 2 MHz | 5 MHz | 5 MHz | 10 MHz | 10 MHz | 5 MHz | 5 MHz | 1 MHz
CM9747 | power setting: 11.40 | 850 V | [23 empty] | 15.5 inches | yes | Nathan on 4/24/14
|  | 900 V |  |  | 782 | 775 | 767 | 762 | 716 | 713 | 535 | 535 |  |  | 25 | 25 | 26 | 26 | 22 | 22 | 22 | 21 | 1.523 | 1.521 | 807.65
|  | 1000 V |  |  | 1550 | 1540 | 1500 | 1490 | 1240 | 1230 | 662 | 675 |  |  | 43 | 43 | 45 | 45 | 42 | 41 | 38 | 39 | 1.363 | 1.311 | 900.55
|  | 1100 V |  |  | 2580 | 2560 | 2440 | 2430 | 1340 | 1420 | 612 | 671 |  |  | 70 | 68 | 65 | 67 | 72 | 67 | 82 | 60 | 1.387 | 1.312 | 993.45
CM9746 | power setting: 11.40 | 850 V | [23 empty] | 15.5 inches | yes | Nathan on 4/25/14
|  | 900 V |  |  | 762 | 758 | 754 | 751 | 706 | 703 | 527 | 531 |  |  | 26 | 25 | 24 | 23 | 24 | 24 | 20 | 20 | 1.446 | 1.477 | 807.3
|  | 1000 V |  |  | 1560 | 1540 | 1540 | 1520 | 1230 | 1230 | 671 | 686 |  |  | 44 | 43 | 45 | 45 | 42 | 41 | 39 | 38 | 1.279 | 1.31 | 900.2
|  | 1100 V |  |  | 2660 | 2620 | 2500 | 2460 | 1250 | 1360 | 603 | 664 |  |  | 67 | 65 | 68 | 65 | 120 | 61 | 92 | 59 | 1.425 | 1.375 | 992.95
CM9761 | power setting: 11.00 | 850 V | [23 empty] | 15.5 inches | yes | Nathan on 4/25/14
|  | 900 V |  |  | 757 | 763 | 747 | 753 | 701 | 707 | 526 | 532 |  |  | 25 | 24 | 23 | 23 | 26 | 26 | 25 | 24 | 1.482 | 1.493 | 806.95
|  | 1000 V |  |  | 1550 | 1560 | 1530 | 1540 | 1240 | 1250 | 675 | 687 |  |  | 43 | 43 | 42 | 42 | 45 | 45 | 40 | 40 | 1.271 | 1.294 | 900.05
|  | 1100 V |  |  | 2610 | 2630 | 2480 | 2500 | 1330 | 1410 | 583 | 654 |  |  | 68 | 70 | 67 | 65 | 71 | 70 | 97 | 60 | 1.189 | 1.242 | 992.9
CM9710 | power setting: 11.40 | 850 V | [23 empty] | 15.5 inches | yes | Nathan on 4/25/14
|  | 900 V |  |  | 770 | 765 | 753 | 750 | 709 | 708 | 525 | 528 |  |  | 25 | 24 | 21 | 21 | 23 | 23 | 23 | 22 | 1.528 | 1.534 | 806.75
|  | 1000 V |  |  | 1580 | 1570 | 1540 | 1530 | 1250 | 1250 | 678 | 683 |  |  | 45 | 46 | 42 | 42 | 44 | 44 | 46 | 46 | 1.363 | 1.319 | 899.9
|  | 1100 V |  |  | 2670 | 2660 | 2520 | 2520 | 1350 | 1410 | 571 | 648 |  |  | 69 | 67 | 71 | 71 | 62 | 62 | 101 | 58 | 1.31 | 1.296 | 992.95
|  |  | 10 kHz | 10 kHz | 1 MHz | 1 MHz | 2 MHz | 2 MHz | 5 MHz | 5 MHz | 10 MHz | 10 MHz | 10 kHz | 10 kHz | 1 MHz | 1 MHz | 2 MHz | 2 MHz | 5 MHz | 5 MHz | 10 MHz | 10 MHz | 5 MHz | 5 MHz | 1 MHz
CM9705 | power setting: 11.40 | 850 V | [23 empty] | 15.5 inches | yes | Nathan on 4/25/14
|  | 900 V |  |  | 730 | 731 | 720 | 721 | 670 | 675 | 520 | 524 |  |  | 24 | 24 | 27 | 26 | 23 | 23 | 21 | 21 | 1.472 | 1.499 | 806.1
|  | 1000 V |  |  | 1490 | 1490 | 1470 | 1470 | 1220 | 1220 | 672 | 687 |  |  | 45 | 45 | 41 | 41 | 43 | 43 | 38 | 39 | 1.275 | 1.283 | 899.3
|  | 1100 V |  |  | 2500 | 2510 | 2390 | 2400 | 1390 | 1430 | 643 | 684 |  |  | 65 | 64 | 67 | 66 | 63 | 62 | 63 | 64 | 1.164 | 1.248 | 992.15
CM9664 | power setting: 11.00 | 850 V | [23 empty] | 15.5 inches | yes | Nathan on 4/25/14
|  | 900 V |  |  | 765 | 761 | 739 | 738 | 695 | 695 | 522 | 525 |  |  | 26 | 26 | 26 | 25 | 25 | 25 | 24 | 25 | 1.53 | 1.589 | 805.7
|  | 1000 V |  |  | 1510 | 1500 | 1490 | 1480 | 1220 | 1220 | 664 | 681 |  |  | 44 | 43 | 42 | 42 | 44 | 44 | 42 | 42 | 1.345 | 1.387 | 898.75
|  | 1100 V |  |  | 2530 | 2520 | 2410 | 2400 | 1350 | 1430 | 601 | 679 |  |  | 62 | 61 | 63 | 60 | 66 | 64 | 89 | 67 | 1.297 | 1.347 | 991.5
CM9653 | power setting: 11.60 | 850 V | [23 empty] | 15.5 inches | yes | Nathan on 4/25/14
|  | 900 V |  |  | 726 | 730 | 714 | 719 | 672 | 675 | 518 | 517 |  |  | 25 | 24 | 22 | 22 | 25 | 24 | 25 | 25 | 1.51 | 1.485 | 806.4
|  | 1000 V |  |  | 1460 | 1470 | 1420 | 1430 | 1200 | 1190 | 688 | 676 |  |  | 43 | 43 | 41 | 41 | 41 | 42 | 37 | 35 | 1.333 | 1.315 | 899.4
|  | 1100 V |  |  | 2400 | 2410 | 2310 | 2320 | 1410 | 1390 | 671 | 615 |  |  | 67 | 69 | 66 | 66 | 58 | 59 | 60 | 97 | 1.341 | 1.334 | 992.1
CM9643 | power setting: 11.60 | 850 V | [23 empty] | 15.5 inches | yes | Nathan on 4/25/14
|  | 900 V |  |  | 762 | 763 | 755 | 757 | 694 | 699 | 528 | 533 |  |  | 25 | 25 | 23 | 24 | 22 | 22 | 23 | 23 | 1.521 | 1.532 | 805.8
|  | 1000 V |  |  | 1500 | 1500 | 1480 | 1480 | 1220 | 1220 | 672 | 668 |  |  | 42 | 42 | 40 | 40 | 41 | 41 | 40 | 42 | 1.356 | 1.345 | 898.7
|  | 1100 V |  |  | 2480 | 2470 | 2360 | 2360 | 1390 | 1380 | 623 | 664 |  |  | 62 | 61 | 63 | 61 | 63 | 62 | 87 | 59 | 1.34 | 1.329 | 991.3
|  |  | 10 kHz | 10 kHz | 1 MHz | 1 MHz | 2 MHz | 2 MHz | 5 MHz | 5 MHz | 10 MHz | 10 MHz | 10 kHz | 10 kHz | 1 MHz | 1 MHz | 2 MHz | 2 MHz | 5 MHz | 5 MHz | 10 MHz | 10 MHz | 5 MHz | 5 MHz | 1 MHz
CM9637 | power setting: 11.60 | 850 V | [23 empty] | 15.5 inches | yes | Nathan on 4/25/14
|  | 900 V |  |  | 776 | 773 | 765 | 764 | 694 | 691 | 525 | 519 |  |  | 24 | 25 | 26 | 26 | 25 | 25 | 24 | 24 | 1.512 | 1.467 | 807.8
|  | 1000 V |  |  | 1520 | 1510 | 1490 | 1480 | 1230 | 1210 | 692 | 668 |  |  | 44 | 43 | 42 | 42 | 41 | 40 | 40 | 40 | 1.354 | 1.318 | 900.8
|  | 1100 V |  |  | 2560 | 2560 | 2410 | 2410 | 1400 | 1300 | 654 | 587 |  |  | 62 | 61 | 61 | 60 | 60 | 87 | 63 | 109 | 1.41 | 1.438 | 993.7
CM9636 | power setting: 11.20 | 850 V | [23 empty] | 15.5 inches | yes | Nathan on 4/28/14 | fall times differ more than usual
|  | 900 V |  |  | 740 | 744 | 724 | 727 | 680 | 686 | 516 | 520 |  |  | 25 | 25 | 24 | 24 | 20 | 20 | 23 | 23 | 1.471 | 1.553 | 805.9
|  | 1000 V |  |  | 1520 | 1520 | 1490 | 1490 | 1220 | 1220 | 695 | 672 |  |  | 45 | 44 | 43 | 42 | 39 | 40 | 35 | 36 | 1.254 | 1.356 | 899.1
|  | 1100 V |  |  | 2610 | 2610 | 2450 | 2460 | 1420 | 1370 | 675 | 616 |  |  | 64 | 64 | 61 | 62 | 65 | 67 | 60 | 85 | 1.211 | 1.319 | 992
CM9630 | power setting: 11.20 | 850 V | [23 empty] | 15.5 inches | yes | Nathan on 4/28/14
|  | 900 V |  |  | 757 | 763 | 745 | 752 | 697 | 703 | 532 | 533 |  |  | 25 | 25 | 25 | 24 | 22 | 22 | 21 | 21 | 1.451 | 1.479 | 807.05
|  | 1000 V |  |  | 1520 | 1520 | 1500 | 1500 | 1240 | 1240 | 685 | 673 |  |  | 44 | 44 | 41 | 41 | 43 | 44 | 43 | 43 | 1.239 | 1.263 | 900.15
|  | 1100 V |  |  | 2580 | 2590 | 2460 | 2470 | 1390 | 1340 | 645 | 587 |  |  | 66 | 67 | 63 | 63 | 61 | 67 | 65 | 96 | 1.161 | 1.221 | 993
CM9629 | power setting: 11.20 | 850 V | [23 empty] | 15.5 inches | yes | Nathan on 4/29/14
|  | 900 V |  |  | 742 | 753 | 735 | 748 | 691 | 704 | 530 | 538 |  |  | 24 | 25 | 26 | 26 | 25 | 25 | 22 | 22 | 1.504 | 1.554 | 806.25
|  | 1000 V |  |  | 1530 | 1550 | 1510 | 1530 | 1240 | 1250 | 680 | 675 |  |  | 41 | 40 | 44 | 43 | 43 | 42 | 40 | 40 | 1.262 | 1.353 | 899.4
|  | 1100 V |  |  | 2630 | 2650 | 2490 | 2510 | 1410 | 1330 | 668 | 604 |  |  | 66 | 65 | 62 | 64 | 61 | 82 | 61 | 89 | 1.207 | 1.388 | 992.35
|  |  | 10 kHz | 10 kHz | 1 MHz | 1 MHz | 2 MHz | 2 MHz | 5 MHz | 5 MHz | 10 MHz | 10 MHz | 10 kHz | 10 kHz | 1 MHz | 1 MHz | 2 MHz | 2 MHz | 5 MHz | 5 MHz | 10 MHz | 10 MHz | 5 MHz | 5 MHz | 1 MHz
CM9651 | power setting: 11.20 | 850 V | [23 empty] | 15.5 inches | yes | Nathan on 4/29/14
|  | 900 V |  |  | 755 | 750 | 735 | 730 | 700 | 695 | 523 | 522 |  |  | 25 | 25 | 24 | 24 | 25 | 24 | 25 | 25 | 1.533 | 1.51 | 806.3
|  | 1000 V |  |  | 1520 | 1510 | 1500 | 1490 | 1240 | 1230 | 674 | 682 |  |  | 45 | 45 | 46 | 45 | 42 | 41 | 40 | 39 | 1.345 | 1.29 | 899.1
|  | 1100 V |  |  | 2540 | 2540 | 2420 | 2420 | 1350 | 1370 | 583 | 647 |  |  | 70 | 71 | 64 | 65 | 65 | 64 | 61 | 61 | 1.386 | 1.315 | 991.8
CM9655 | power setting: 11.60 | 850 V | [23 empty] | 15.5 inches | yes | Nathan on 4/29/14
|  | 900 V |  |  | 724 | 728 | 712 | 717 | 676 | 679 | 525 | 523 |  |  | 23 | 23 | 24 | 24 | 23 | 23 | 25 | 26 | 1.481 | 1.478 | 807.25
|  | 1000 V |  |  | 1480 | 1490 | 1450 | 1450 | 1220 | 1210 | 680 | 656 |  |  | 43 | 43 | 41 | 41 | 41 | 40 | 40 | 39 | 1.305 | 1.312 | 900.3
|  | 1100 V |  |  | 2490 | 2500 | 2380 | 2390 | 1430 | 1330 | 673 | 594 |  |  | 67 | 65 | 64 | 65 | 58 | 68 | 63 | 98 | 1.294 | 1.376 | 993.4
CM9661 | power setting: 11.20 | 850 V | [23 empty] | 15.5 inches | yes | Nathan on 4/29/14
|  | 900 V |  |  | 762 | 763 | 750 | 749 | 702 | 701 | 530 | 530 |  |  | 24 | 24 | 26 | 25 | 24 | 23 | 20 | 20 | 1.55 | 1.532 | 806.55
|  | 1000 V |  |  | 1510 | 1510 | 1490 | 1490 | 1240 | 1230 | 682 | 674 |  |  | 44 | 45 | 43 | 43 | 41 | 42 | 39 | 39 | 1.337 | 1.349 | 899.65
|  | 1100 V |  |  | 2520 | 2530 | 2390 | 2390 | 1400 | 1360 | 681 | 618 |  |  | 67 | 66 | 62 | 63 | 61 | 60 | 61 | 97 | 1.328 | 1.377 | 992.5
CM9674 | power setting: 11.60 | 850 V | [23 empty] | 15.5 inches | yes | Nathan on 4/29/14
|  | 900 V |  |  | 722 | 721 | 700 | 700 | 659 | 661 | 505 | 512 |  |  | 21 | 21 | 25 | 24 | 22 | 22 | 23 | 22 | 1.522 | 1.567 | 807.3
|  | 1000 V |  |  | 1430 | 1420 | 1390 | 1390 | 1200 | 1200 | 662 | 681 |  |  | 45 | 45 | 40 | 40 | 41 | 40 | 37 | 37 | 1.337 | 1.372 | 900.2
|  | 1100 V |  |  | 2400 | 2390 | 2300 | 2290 | 1370 | 1420 | 607 | 684 |  |  | 65 | 63 | 61 | 60 | 61 | 60 | 110 | 68 | 1.299 | 1.315 | 993.1
|  |  | 10 kHz | 10 kHz | 1 MHz | 1 MHz | 2 MHz | 2 MHz | 5 MHz | 5 MHz | 10 MHz | 10 MHz | 10 kHz | 10 kHz | 1 MHz | 1 MHz | 2 MHz | 2 MHz | 5 MHz | 5 MHz | 10 MHz | 10 MHz | 5 MHz | 5 MHz | 1 MHz
CM9631 | power setting: 11.20 | 850 V | [23 empty] | 15.5 inches | yes | Nathan on 4/29/14
|  | 900 V |  |  | 766 | 770 | 753 | 758 | 699 | 702 | 524 | 525 |  |  | 24 | 24 | 22 | 22 | 23 | 23 | 22 | 22 | 1.516 | 1.491 | 807.7
|  | 1000 V |  |  | 1520 | 1520 | 1480 | 1480 | 1220 | 1220 | 676 | 672 |  |  | 43 | 43 | 43 | 44 | 39 | 38 | 40 | 39 | 1.347 | 1.335 | 900.65
|  | 1100 V |  |  | 2510 | 2510 | 2370 | 2370 | 1400 | 1360 | 684 | 633 |  |  | 64 | 65 | 67 | 66 | 62 | 62 | 63 | 101 | 1.312 | 1.363 | 993.55
CM9744 | power setting: 11.20 | 850 V | [23 empty] | 15.5 inches | yes | Nathan on 4/29/14
|  | 900 V |  |  | 720 | 720 | 705 | 707 | 651 | 655 | 505 | 511 |  |  | 23 | 23 | 23 | 23 | 21 | 21 | 24 | 24 | 1.551 | 1.573 | 805.4
|  | 1000 V |  |  | 1470 | 1460 | 1440 | 1440 | 1190 | 1200 | 654 | 675 |  |  | 45 | 44 | 45 | 44 | 40 | 40 | 41 | 42 | 1.356 | 1.36 | 898.55
|  | 1100 V |  |  | 2470 | 2460 | 2340 | 2340 | 1320 | 1380 | 661 | 696 |  |  | 63 | 61 | 60 | 61 | 60 | 61 | 60 | 62 | 1.308 | 1.337 | 991.5
CM9748 | power setting: 11.60 | 850 V | [23 empty] | 15.5 inches | yes | Nathan on 4/29/14
|  | 900 V |  |  | 756 | 767 | 727 | 738 | 680 | 690 | 515 | 527 |  |  | 23 | 23 | 22 | 22 | 20 | 20 | 22 | 22 | 1.489 | 1.549 | 806.35
|  | 1000 V |  |  | 1480 | 1500 | 1460 | 1480 | 1200 | 1220 | 661 | 681 |  |  | 40 | 40 | 40 | 41 | 42 | 41 | 35 | 37 | 1.322 | 1.363 | 899.55
|  | 1100 V |  |  | 2510 | 2540 | 2360 | 2390 | 1360 | 1420 | 678 | 707 |  |  | 61 | 61 | 60 | 59 | 59 | 60 | 60 | 60 | 1.312 | 1.365 | 992.55
CM9749 | power setting: 11.60 | 850 V | [23 empty] | 15.5 inches | yes | Nathan on 4/30/14
|  | 900 V |  |  | 722 | 732 | 717 | 727 | 674 | 682 | 517 | 525 |  |  | 22 | 22 | 21 | 21 | 20 | 20 | 20 | 20 | 1.502 | 1.529 | 805.4
|  | 1000 V |  |  | 1500 | 1520 | 1480 | 1500 | 1210 | 1230 | 652 | 684 |  |  | 42 | 42 | 41 | 42 | 40 | 40 | 40 | 41 | 1.352 | 1.313 | 898.55
|  | 1100 V |  |  | 2510 | 2560 | 2380 | 2430 | 1320 | 1440 | 591 | 689 |  |  | 63 | 63 | 64 | 65 | 62 | 61 | 89 | 62 | 1.308 | 1.299 | 991.35
|  |  | 10 kHz | 10 kHz | 1 MHz | 1 MHz | 2 MHz | 2 MHz | 5 MHz | 5 MHz | 10 MHz | 10 MHz | 10 kHz | 10 kHz | 1 MHz | 1 MHz | 2 MHz | 2 MHz | 5 MHz | 5 MHz | 10 MHz | 10 MHz | 5 MHz | 5 MHz | 1 MHz
No label-10 | power setting: 11.20 | 850 V | [23 empty] | 15.5 inches | yes | Nathan on 4/30/14
|  | 900 V |  |  | 734 | 747 | 723 | 737 | 683 | 693 | 522 | 527 |  |  | 25 | 25 | 22 | 22 | 23 | 22 | 21 | 21 | 1.492 | 1.528 | 807.1
|  | 1000 V |  |  | 1480 | 1500 | 1450 | 1470 | 1220 | 1230 | 683 | 678 |  |  | 42 | 42 | 43 | 44 | 43 | 43 | 39 | 40 | 1.281 | 1.339 | 900.05
|  | 1100 V |  |  | 2500 | 2520 | 2410 | 2420 | 1400 | 1370 | 665 | 613 |  |  | 66 | 64 | 69 | 67 | 60 | 60 | 66 | 98 | 1.191 | 1.295 | 993
CM9824 | power setting: 11.40 | 850 V | [23 empty] | 15.5 inches | yes | Nathan on 4/30/14
|  | 900 V |  |  | 752 | 740 | 738 | 729 | 685 | 681 | 516 | 518 |  |  | 22 | 22 | 26 | 26 | 25 | 25 | 22 | 22 | 1.562 | 1.555 | 804.1
|  | 1000 V |  |  | 1570 | 1550 | 1540 | 1520 | 1230 | 1230 | 664 | 662 |  |  | 44 | 44 | 41 | 41 | 42 | 42 | 43 | 44 | 1.297 | 1.344 | 897.4
|  | 1100 V |  |  | 2650 | 2630 | 2500 | 2480 | 1400 | 1370 | 657 | 641 |  |  | 64 | 63 | 60 | 61 | 65 | 64 | 60 | 63 | 1.149 | 1.187 | 990.45
CM9826 | power setting: 11.60 | 850 V | [23 empty] | 15.5 inches | yes | Nathan on 4/30/14
|  | 900 V |  |  | 794 | 792 | 780 | 779 | 732 | 729 | 543 | 538 |  |  | 24 | 24 | 25 | 24 | 24 | 24 | 23 | 22 | 1.503 | 1.479 | 805.8
|  | 1000 V |  |  | 1640 | 1630 | 1610 | 1600 | 1270 | 1240 | 678 | 655 |  |  | 44 | 44 | 40 | 40 | 43 | 42 | 40 | 40 | 1.252 | 1.315 | 899
|  | 1100 V |  |  | 2760 | 2710 | 2590 | 2530 | 1400 | 1300 | 649 | 563 |  |  | 66 | 65 | 69 | 66 | 61 | 59 | 63 | 101 | 1.209 | 1.19 | 991.85
CM9828 | power setting: 11.40 | 850 V | [23 empty] | 15.5 inches | yes | Nathan on 4/30/14
|  | 900 V |  |  | 755 | 759 | 743 | 745 | 702 | 707 | 535 | 543 |  |  | 25 | 24 | 24 | 24 | 25 | 25 | 24 | 24 | 1.522 | 1.539 | 807
|  | 1000 V |  |  | 1550 | 1550 | 1520 | 1520 | 1240 | 1240 | 662 | 691 |  |  | 43 | 43 | 44 | 45 | 42 | 41 | 39 | 40 | 1.363 | 1.329 | 899.95
|  | 1100 V |  |  | 2620 | 2650 | 2480 | 2500 | 1190 | 1390 | 596 | 673 |  |  | 68 | 67 | 62 | 62 | 190 | 71 | 85 | 61 | 1.522 | 1.411 | 992.7
|  |  | 10 kHz | 10 kHz | 1 MHz | 1 MHz | 2 MHz | 2 MHz | 5 MHz | 5 MHz | 10 MHz | 10 MHz | 10 kHz | 10 kHz | 1 MHz | 1 MHz | 2 MHz | 2 MHz | 5 MHz | 5 MHz | 10 MHz | 10 MHz | 5 MHz | 5 MHz | 1 MHz
CM9837 | power setting: 11.60 | 850 V | [23 empty] | 15.5 inches | yes | Nathan on 5/1/14
|  | 900 V |  |  | 726 | 740 | 716 | 731 | 672 | 682 | 516 | 518 |  |  | 22 | 23 | 23 | 23 | 24 | 24 | 22 | 22 | 1.483 | 1.491 | 806
|  | 1000 V |  |  | 1500 | 1530 | 1470 | 1500 | 1220 | 1220 | 681 | 670 |  |  | 42 | 43 | 44 | 44 | 40 | 40 | 39 | 39 | 1.235 | 1.302 | 899.15
|  | 1100 V |  |  | 2560 | 2610 | 2400 | 2430 | 1420 | 1350 | 682 | 639 |  |  | 66 | 66 | 65 | 64 | 62 | 62 | 60 | 80 | 1.166 | 1.147 | 992.15
CM9840 | power setting: 11.60 | 850 V | [23 empty] | 15.5 inches | yes | Nathan on 5/1/14
|  | 900 V |  |  | 755 | 760 | 746 | 751 | 702 | 708 | 530 | 537 |  |  | 25 | 25 | 23 | 24 | 25 | 25 | 23 | 23 | 1.521 | 1.503 | 805.7
|  | 1000 V |  |  | 1560 | 1560 | 1530 | 1540 | 1250 | 1260 | 670 | 686 |  |  | 45 | 44 | 46 | 46 | 42 | 41 | 39 | 41 | 1.355 | 1.312 | 898.7
|  | 1100 V |  |  | 2650 | 2670 | 2510 | 2530 | 1360 | 1390 | 584 | 643 |  |  | 68 | 67 | 67 | 65 | 61 | 60 | 92 | 58 | 1.303 | 1.28 | 991.65
CM9841 | power setting: 11.40 | 850 V | [23 empty] | 15.5 inches | yes | Nathan on 5/1/14
|  | 900 V |  |  | 752 | 767 | 738 | 750 | 695 | 704 | 531 | 531 |  |  | 24 | 24 | 22 | 23 | 24 | 24 | 21 | 20 | 1.472 | 1.455 | 805.45
|  | 1000 V |  |  | 1540 | 1570 | 1520 | 1540 | 1240 | 1230 | 680 | 671 |  |  | 45 | 45 | 40 | 40 | 42 | 42 | 45 | 45 | 1.256 | 1.262 | 898.55
|  | 1100 V |  |  | 2600 | 2650 | 2470 | 2490 | 1390 | 1350 | 671 | 623 |  |  | 64 | 64 | 62 | 63 | 62 | 62 | 61 | 87 | 1.194 | 1.161 | 991.6
CM9843 | power setting: 11.00 | 850 V | [23 empty] | 15.5 inches | yes | Nathan on 5/1/14
|  | 900 V |  |  | 743 | 742 | 736 | 734 | 688 | 689 | 520 | 522 |  |  | 24 | 24 | 25 | 25 | 25 | 25 | 22 | 22 | 1.475 | 1.496 | 805.75
|  | 1000 V |  |  | 1530 | 1520 | 1510 | 1500 | 1220 | 1210 | 677 | 669 |  |  | 41 | 41 | 44 | 44 | 43 | 43 | 40 | 40 | 1.232 | 1.262 | 898.95
|  | 1100 V |  |  | 2580 | 2560 | 2450 | 2440 | 1410 | 1350 | 682 | 643 |  |  | 66 | 64 | 66 | 63 | 66 | 65 | 63 | 68 | 1.199 | 1.123 | 992
|  |  | 10 kHz | 10 kHz | 1 MHz | 1 MHz | 2 MHz | 2 MHz | 5 MHz | 5 MHz | 10 MHz | 10 MHz | 10 kHz | 10 kHz | 1 MHz | 1 MHz | 2 MHz | 2 MHz | 5 MHz | 5 MHz | 10 MHz | 10 MHz | 5 MHz | 5 MHz | 1 MHz
CM9845 | power setting: 11.60 | 850 V | [23 empty] | 15.5 inches | yes | Nathan on 5/1/14
|  | 900 V |  |  | 746 | 725 | 733 | 716 | 678 | 670 | 515 | 514 |  |  | 22 | 22 | 24 | 24 | 25 | 25 | 20 | 20 | 1.493 | 1.505 | 804.75
|  | 1000 V |  |  | 1540 | 1500 | 1500 | 1470 | 1220 | 1210 | 684 | 667 |  |  | 45 | 44 | 40 | 40 | 44 | 43 | 40 | 39 | 1.254 | 1.275 | 897.8
|  | 1100 V |  |  | 2600 | 2550 | 2450 | 2420 | 1430 | 1390 | 687 | 651 |  |  | 64 | 63 | 63 | 62 | 59 | 60 | 61 | 74 | 1.23 | 1.209 | 990.85