## Sheet1
| Consistency | High Voltage | Laser Frequency (Min1 Mean/Min2 Mean) (negative) (mV) |  |  |  |  | Standard Deviation (Output 1/Output 2) (mV) |  |  |  |  | Fall Time at 5MHz | Current at 10kHz | Distance between | Snapshot taken | Testing done by and date | Notes | Date taken to CUA
PMT Serial # | 800mV @900V | V | 10kHz | 1 MHz | 2 MHz | 5 MHz | 10 MHz | 10kHz | 1 MHz | 2 MHz | 5 MHz | 10 MHz | ns | μA | laser/edge PMT
CM9639 |  | 850 V | 339/338 | 350/348 | 369/368 | 391/390 | 367/361 | [5 empty] | 1.479/1.519 | 759.95 | n/a | yes | Natalie on 10/14/13 | no fire resistant | 2013-10-14
|  | 900 V | 450/450 | 501/503 | 528/536 | 585/585 | 524/523 | [5 empty] | 1.453/1.459 | 807.4
|  | 1000 V | 1.11/1.10 | 1.15/1.14 | 1.19/1.18 | 1.20/1.21 | 778/781 | [5 empty] | 1.366/1.340 | 899.55
|  | 1100 V | 2.16/2.11 | 2.17/2.13 | 2.24/2.20 | 1.63/1.65 | 882/895 | [5 empty] | 1.146/1.642 | 991.7
CM9594 |  | 850 V | 365/368 | 379/381 | 399/399 | 426/425 | 383/383 | [5 empty] | 1.510/1.512 | 759.5 | n/a | yes | Natalie on 10/14/13 |  | 2013-10-14
|  | 900 V | 550/558 | 570/575 | 598/605 | 640/643 | 552/550 | [5 empty] | 1.433/1.380 | 806
|  | 1000 V | 1.19/1.20 | 1.24/1.25 | 1.29/1.30 | 1.28/1.29 | 796/787 | [5 empty] | 1.319/1.340 | 899.1
|  | 1100 V | 2.29/2.31 | 2.31/2.32 | 2.36/2.37 | 1.75/1.68 | 880/878 | [5 empty] | 1.260/1.200 | 991.15
CM9592 |  | 850 V | 336/335 | 356/354 | 372/369 | 395/392 | 364/360 | [5 empty] | 1.441/1.505 | 760 | n/a | yes | Natalie on 10/14/13 |  | 2013-10-14
|  | 900 V | 518/515 | 551/447 | 578/574 | 616/612 | 537/532 | [5 empty] | 1.405/1.431 | 806.8
|  | 1000 V | 1.10/1.09 | 1.14/1.13 | 1.18/1.16 | 1.19/1.18 | 789/789 | [5 empty] | 1.310/1.361 | 899.5
|  | 1100 V | 1.99/1.96 | 2.03/2.00 | 2.10/2.08 | 1.67/1.68 | 820/825 | [5 empty] | 1.25/1.25 | 991.55
CM9593 |  | 850 V | 315/314 | 325/323 | 338/337 | 365/364 | 333/330 | [5 empty] | 1.496/1.455 | 758 | n/a | yes | Natalie on 10/14/13 |  | 2013-10-14
|  | 900 V | 525/528 | 464/468 | 494/496 | 532/534 | 487/484 | [5 empty] | 1.440/1.390 | 805
|  | 1000 V | 1.04/1.04 | 1.09/1.09 | 1.14/1.15 | 1.15/1.15 | 787/782 | [5 empty] | 1.340/1.350 | 898.25
|  | 1100 V | 1.98/1.97 | 2.00/2.00 | 2.07/2.07 | 1.65/1.64 | 898/902 | [5 empty] | 1.186/1.234 | 990.4
CM9571 |  | 850 V | 356/351 | 365/363 | 302/299 | 360/357 | 358/352 | [5 empty] | 1.67/1.68 | 760 | n/a | no | Franz on 10/24/13 |  | 2013-10-26
|  | 900 V | 403/399 | 401/398 | 442/435 | 538/532 | 498/494 | [5 empty] | 1.64/1.66 | 806
|  | 1000 V | 821/812 | 865/857 | 955/947 | 1.11/1.09 | 739/719 | [5 empty] | 1.73/1.71 | 907
|  | 1100 V | 1.57/1.55 | 1.60/1.59 | 1.68/1.67 | 1.59/1.52 | 776/748 | [5 empty] | 1.84/1.85 | 993
CM9791 | yes | 850 V | 708/709 | 721/722 | 715/714 | 707/708 | 565/568 | 67/68 | 59/59 | 58/57 | 67/68 | 58/59 | 1.436/1.411 | 758.8 | 15.5 inches | yes | Natalie on 11/7/13 | output 1 problem | taken by Natalie on 11/14/13
|  | 900 V | 1.09/1.09 | 1.09/1.09 | 1.08/1.08 | 1.02/1.02 | 785/786 | 80/80 | 76/76 | 72/75 | 92/93 | 118/117 | 1.3/1.3 | 806 |  |  |  | (see red)
|  | 1000 V | 2.16/2.15 | 2.17/2.17 | 2.15/2.16 | 1.52/1.53 | 831/828 | 162/164 | 157/155 | 169/161 | 144/146 | 179/180 | 1.274/1.079 | 898.25
|  | 1100 V | 3.67/3.67 | 3.55/3.54 | 3.34/3.33 | 1.17/1.54 | 491/846 | 197/198 | 272/264 | 213/215 | 125/150 | 110/141 | 1.290/1.216 | 990
|  |  | 10kHz | 1 MHz | 2 MHz | 5 MHz | 10 MHz | 10kHz | 1 MHz | 2 MHz | 5 MHz | 10 MHz
CM9393 | yes | 850 V | 582/573 | 586/575 | 568/567 | 553/540 | 478/472 | 40.7/45.2 | 50.9/50.8 | 45.2/44.8 | 51.3/51.8 | 55.2/51.6 | 1.332/1.447 | 758.9 | 15.25 inches | yes | Natalie on 11/7/13 |  | taken by Natalie on 11/14/13
|  | 900 V | 915/901 | 925/913 | 900/881 | 836/846 | 654/642 | 75/79.2 | 80.5/77.7 | 77.1/77.2 | 74.8/73.1 | 70.6/61.6 | 1.382/1.281 | 805.65
|  | 1000 V | 1.83/1.82 | 1.82/1.80 | 1.81/1.79 | 1.48/1.46 | 812/816 | 134/135 | 116/114 | 122/124 | 136/127 | 147/142 | 1.260/1.287 | 899.45
|  | 1100 V | 3.06/3.04 | 2.98/2.96 | 2.93/2.90 | 1.67/1.71 | 748/842 | 185/188 | 162/164 | 186/189 | 216/210 | 172/146 | 1.011/1.204 | 990.3
CM9635 | yes | 850 V | 684/680 | 685/680 | 678/674 | 636/638 | 532/533 | 54.9/53.1 | 55.0/55.4 | 67.2/65.0 | 62.5/60.4 | 50.8/52.7 | 1.375/1.333 | 760.1 | 15.5 inches | yes | Natalie on 11/7/13 |  | taken by Natalie on 11/14/13
|  | 900 V | 1.08/1.07 | 1.05/1.04 | 1.06/1.05 | 964/963 | 665/663 | 83.3/85.8 | 93.2/95.4 | 85.4/86.2 | 83.8/77.3 | 83.0/84.0 | 1.409/1.337 | 806.95
|  | 1000 V | 2.11/2.09 | 2.10/2.07 | 2.02/2.00 | 1.56/1.57 | 786/783 | 126/124 | 158/151 | 121/119 | 161/156 | 113/112 | 1.257/1.125 | 899.55
|  | 1100 V | 3.54/3.5 | 3.52/3.47 | 3.23/3.22 | 1.52/1.60 | 730/706 | 196/197 | 244/231 | 185/182 | 238/186 | 196/201 | 1.153/1.192 | 991.45
No Label-1 | yes | 850 V | 658/657 | 665/661 | 671/667 | 626/624 | 513/510 | 58.2/59.4 | 51.1/55.3 | 65.2/67.4 | 62.4/60.7 | 55.0/53.4 | 1.400/1.465 | 758.7 | 15.5 inches | yes | Natalie on 11/8/13 |  | taken by Natalie on 11/14/13
|  | 900 V | 1.03/1.03 | 1.02/1.02 | 1.03/1.03 | 934/931 | 657/658 | 86.3/88.2 | 95.3/95.8 | 82.2/84.5 | 80.0/78.9 | 74.0/744.7 | 1.385/1.403 | 805.45
|  | 1000 V | 2.06/2.06 | 2.03/2.03 | 2.00/2.00 | 1.55/1.55 | 787/783 | 141/137 | 128/128 | 121/117 | 136/137 | 140/148 | 1.181/1.215 | 898.45
|  | 1100 V | 3.35/3.34 | 3.37/3.38 | 3.15/3.15 | 1.64.1.72 | 785/825 | 311/310 | 193/192 | 200/201 | 206/204 | 201/197 | 1.117/1.120 | 990.2
No Label-2 | yes | 850 V | 661/599 | 635/572 | 672/602 | 631/567 | 512/460 | 481./43.5 | 50.8/45.2 | 82.6/72.2 | 57.0/48.1 | 58.6/53.0 | 1.340/1.418 | 760.1 | 15.5 inches | yes | Natalie on 11/8/13 | large difference | taken by Natalie on 11/14/13
|  | 900 V | 1.03/930 | 990/886 | 983/880 | 923/821 | 678/612 | 90.6/80.8 | 90.5/70.8 | 77.5/66.3 | 80.5/70.9 | 77.9/63.5 | 1.367/1.333 | 807.5 |  |  |  | between O1 and O2
|  | 1000 V | 2.01/1.81 | 1.99/1.78 | 1.97/1.76 | 1.56/1.39 | 826/741 | 131/127 | 148/126 | 141/131 | 124/115 | 116/108 | 1.291/1.244 | 899.95 |  |  |  | no fire resistant
|  | 1100 V | 3.44/3.10 | 3.35/3.00 | 3.19/2.82 | 1.50/1.51 | 735/720 | 703/604 | 247/238 | 186/172 | 296/150 | 157/146 | 1.255/1.168 | 991.5
CM9694 | yes | 850 V | 738/731 | 726/717 | 737/731 | 680/672 | 557/549 | 55.0/51.9 | 55.8/53.1 | 66.6/65.0 | 68.3/70.0 | 48.6/48.0 | 1.406/1/415 | 760.15 | 15.5 inches | yes | Natalie on 11/8/13 |  | taken by Natalie on 11/14/13
|  | 900 V | 1.13/1.11 | 1.15/1.14 | 1.14/1.13 | 1.02/1.00 | 672/673 | 85.1/85.9 | 90.4/88.0 | 92.7/89.3 | 93.9/92.7 | 86.2/83.0 | 1.345/1.275 | 807
|  | 1000 V | 2.17/2.15 | 2.17/2.14 | 2.10/2.08 | 1.54/1.55 | 812/821 | 133/134 | 152/148 | 137/138 | 128/122 | 182/179 | 1.121/1.202 | 899.9
|  | 1100 V | 3.47/3.44 | 3.54.3.50 | 3.29/3.27 | 1.56/1.66 | 678/758 | 236/233 | 196/195 | 190/190 | 211/195 | 136/137 | 1.213/1.106 | 991.55
|  |  | 10kHz | 1 MHz | 2 MHz | 5 MHz | 10 MHz | 10kHz | 1 MHz | 2 MHz | 5 MHz | 10 MHz
CM9756 | yes | 850 V | 717/723 | 703/705 | 711/715 | 658/657 | 544/544 | 51.1/53.7 | 58.9/62.3 | 52.3/53.0 | 68.0/69.1 | 55.7/55.8 | 1.405/1.440 | 759.95 | 15.25 inches | yes | Natalie on 11/8/13 |  | taken by Natalie on 11/14/13
|  | 900 V | 1.15/1.15 | 1.10/1.11 | 1.08/1.09 | 992/996 | 670/669 | 86.4/83.8 | 103/105 | 88.0/88.5 | 77.5/74.7 | 73.9/74.2 | 1.380/1.364 | 806.85
|  | 1000 V | 2.15/2.15 | 2.12/2.12 | 2.11/2.10 | 1.57/1.57 | 842/837 | 163/159 | 174/169 | 155/151 | 154/149 | 138/134 | 1.218/1.117 | 899.75
|  | 1100 V | 3.58/2.57 | 3.56/3.54 | 3.28/3.27 | 1.44/1.64 | 554/792 | 195/188 | 208/202 | 331/300 | 166/180 | 991/121 | 1.142/1.224 | 991.25
CM9634 | yes | 850 V | 736/723 | 737/726 | 744/732 | 697/692 | 549/543 | 56.6/56.4 | 64.0/63.0 | 61.4/60.7 | 62.3/58.5 | 64.1/63.9 | 1.417/1.314 | 759.55 | 15.5 inches | yes | Natalie on 11/11/13 |  | taken by Natalie on 11/14/13
|  | 900 V | 1.11/1.09 | 1.12/1.10 | 1.11/1.09 | 1.02/1.01 | 683/687 | 89.1/87.1 | 104/101 | 91.1/95.6 | 95.5/91.6 | 77.1/76.8 | 1.356/1.320 | 806.85
|  | 1000 V | 2.24/2.20 | 2.22/2.18 | 2.20/2.16 | 1.57/1.59 | 826/833 | 119/118 | 141/139 | 135/128 | 147/141 | 156/151 | 1.249/1.242 | 899.1
|  | 1100 V | 3.65/3.59 | 3.68/3.62 | 3.29/3.26 | 1.49/1.63 | 675/704 | 228/226 | 217/213 | 164/156 | 231/180 | 174/167 | 1.450/1.148 | 990.95
CM9832 | yes | 850 V | 477/482 | 467/469 | 473/471 | 440/440 | 395/398 | 39.0/38.3 | 39.2/38.7 | 44.6/43.5 | 41.8/44.1 | 53.0/52.3 | 1.470/1.408 | 726.75 | 15.5 inches | yes | Natalie on 11/11/13 | lower overall???? | taken by Natalie on 11/14/13
| checked twice | 900 V | 759/764 | 761/765 | 769/774 | 730/733 | 588/589 | 64.1/65.3 | 75.2/74.0 | 63.1/62.4 | 73.1/72.5 | 64.8/63.1 | 1.299/1.038 | 773.45 |  |  |  | currents are lower
|  | 1000 V | 1.62/1.64 | 1.59/1.61 | 1.62/1.64 | 1.38/1.39 | 833/828 | 117/116 | 111/113 | 110/113 | 121/128 | 107/112 | 1.281/1.319 | 865.9 |  |  |  | too
|  | 1100 V | 2.86/2.92 | 2.82/2.87 | 2.69/2.75 | 1.70/1.76 | 767/780 | 199/205 | 204/214 | 164/177 | 184/192 | 205/167 | 1.188/1.143 | 958.9
No label-3 | yes | 850 V | 602/600 | 595/590 | 600/604 | 563/560 | 496/492 | 50.9/50.1 | 60.4/58.2 | 55.9/54.9 | 50.2/49.9 | 46.0/45.8 | 1.404/1.384 | 761 | 15.5 inches | yes | Natalie on 11/11/13 |  | taken by Natalie on 11/14/13
|  | 900 V | 965/957 | 950/941 | 910/906 | 865/859 | 660/658 | 83.3/87.2 | 70.8/70.2 | 67.5/65.6 | 83.4/80.0 | 68.5/67.4 | 1.381/1.325 | 807.8
|  | 1000 V | 1.85/1.82 | 1.84/1.82 | 1.81/1.79 | 1.53/1.52 | 781/791 | 124/127 | 135/132 | 133/312 | 129/130 | 153/149 | 1.294/1.175 | 900.75
|  | 1100 V | 3.16/3.11 | 3.15/3.10 | 3.00/2.96 | 1.66/1.74 | 771/865 | 167/163 | 224/227 | 200/201 | 242/224 | 226/198 | 1.037/1.195 | 989.05
|  |  | 10kHz | 1 MHz | 2 MHz | 5 MHz | 10 MHz | 10kHz | 1 MHz | 2 MHz | 5 MHz | 10 MHz
No label-4 | BAD | 850 V | # | # | # | # | # | # | # | # | # | # | # | # |  |  | Natalie on 11/12/13 | BAD, put in box | BAD
|  | 900 V | # | # | # | # | # | # | # | # | # | # | # | # |  |  |  | for Vladimir to look
|  | 1000 V | # | # | # | # | # | # | # | # | # | # | # | # |  |  |  | at later
|  | 1100 V | # | # | # | # | # | # | # | # | # | # | # | #
CM9759 | BAD | 850 V | # | # | # | # | # | # | # | # | # | # | # | # |  |  | Natalie on 11/12/13 | BAD, put in box | BAD
|  | 900 V | # | # | # | # | # | # | # | # | # | # | # | # |  |  |  | for Vladimir to look
|  | 1000 V | # | # | # | # | # | # | # | # | # | # | # | # |  |  |  | at later
|  | 1100 V | # | # | # | # | # | # | # | # | # | # | # | #
No label-5 | yes | 850 V | 667/671 | 642/640 | 668/669 | 628/629 | 505/508 | 50.2/50.0 | 51.5/52.5 | 77.7/75.7 | 55.0/54.8 | 58.0/58.1 | 1.303/1.427 | 760.15 | 15.5 inches | yes | Natalie on 11/12/13 |  | taken by Natalie on 11/14/13
|  | 900 V | 1.02/1.02 | 981/974 | 962/959 | 903/899 | 657/661 | 95.9/92.5 | 96.0/92.2 | 86.1/85.0 | 82.6/79.4 | 99.7/97.1 | 1.349/1.339 | 807.05
|  | 1000 V | 1.98/1.97 | 2.02/2.00 | 1.97/1.96 | 1.55/1.55 | 834/845 | 131/129 | 136/132 | 144/138 | 149/149 | 165/167 | 1.160/1.164 | 899.75
|  | 1100 V | 3.34/3.31 | 3.37/3.35 | 3.14/3.13 | 1.49/1.69 | 746/807 | 207/200 | 223/225 | 190/193 | 318/195 | 147/152 | 1.138/1.193 | 991.45
No label-6 | yes | 850 V | 664/662 | 654/652 | 678/673 | 646/643 | 525/522 | 58.8/57.6 | 56.7/55.2 | 72.9/71.8 | 52.9/54.2 | 55.9/58.4 | 1.308/1.385 | 760.6 | 15.5 inches | yes | Natalie on 11/12/13 |  | taken by Natalie on 11/14/13
|  | 900 V | 1.04/1.04 | 1.03/1.03 | 1.02/1.02 | 948/946 | 720/718 | 75.8/84.4 | 80.0/77.3 | 84.6/81.8 | 81.7/82.4 | 87.1/84.6 | 1.406/1.409 | 807.6
|  | 1000 V | 2.00/1.98 | 2.00/1.99 | 1.95/1.92 | 1.58/1.57 | 820/827 | 126/123 | 134/124 | 110/113 | 149/145 | 120/128 | 1.283/1.253 | 900.4
|  | 1100 V | 3.34/3.32 | 3.31/3.29 | 3.05/3.02 | 1.67/1.74 | 778/751 | 206/205 | 192/194 | 178/176 | 298/226 | 170/178 | 1.123/1.167 | 991.85
CM9626 | yes | 850 V | 716/718 | 680/683 | 735/737 | 674/681 | 555/559 | 58.5/58.1 | 72.2/70.2 | 77.3/79.5 | 60.1/60.2 | 64.6/63.8 | 1.307/1.407 | 759.95 | 15.5 inches | yes | Natalie on 11/12/13 | pic named 9686 | taken by Natalie on 11/14/13
|  | 900 V | 1.11/1.12 | 11.0/1.10 | 1.08/1.08 | 1.02/1.03 | 717/720 | 84.0/87.6 | 101/101 | 89.0/88.1 | 92.3/90.5 | 93.6/92.0 | 1.360/1.355 | 806.8 |  |  |  | by accident
|  | 1000 V | 2.17/2.18 | 2.15/2.15 | 2.10/2.10 | 1.61/1.62 | 832/830 | 134/132 | 156/159 | 148/146 | 168/174 | 142/145 | 1.242/1.240 | 899.3
|  | 1100 V | 3.59/3.61 | 3.47/3.48 | 3.32/3.33 | 1.63/1.68 | 692/699 | 212/210 | 209/205 | 187/182 | 236/197 | 147/148 | 1.214/1.169 | 991.1
CM9565 | yes | 850 V | 781/771 | 773/767 | 778/771 | 710/706 | 582/575 | 67.6/68.2 | 66.8/64.0 | 71.5/70.4 | 73.9/70.1 | 74.7/70.1 | 1.422/1.398 | 758.9 | 15.5 inches | yes | Natalie on 11/12/13 | 1Mhz,1100V | taken by Natalie on 11/14/13
|  | 900 V | 1.21/1.20 | 1.19/1.18 | 1.18/1.17 | 1.07/1.06 | 762/760 | 88.8/90.6 | 81.6/77.0 | 107/105 | 83.5/81.9 | 154/150 | 1.323/1.387 | 806.35 |  |  |  | standard jumped entire time
|  | 1000 V | 2.33/2.28 | 2.27/2.23 | 2.22/2.18 | 1.61/1.61 | 794/829 | 140/131 | 170/162 | 160/151 | 206/198 | 166/158 | 1.200/1.266 | 898.55
|  | 1100 V | 3.86/3.78 | 3.80/3.72 | 3.49/3.39 | 895/1.41 | 474/863 | 219/217 | 395/380 | 201/198 | 170/155 | 110/155 | 1.574/1.187 | 990.2
|  |  | 10kHz | 1 MHz | 2 MHz | 5 MHz | 10 MHz | 10kHz | 1 MHz | 2 MHz | 5 MHz | 10 MHz
CM9639 | yes; power setting: 10.90 | 850 V | 642/643 | 643/643 | 645/646 | 642/640 | 637/633 | 21.2/20.8 | 22.3/21.4 | 23.8/22.5 | 19.2/20.1 | 15.5/14.6 | 1.469/1.285 | 759.7 | 15.5 inches | yes | Nathan on 12/5/13 | retest; no fire-resistant wrap | taken by Franz on 12/11/13
|  | 900 V | 804/802 | 813/815 | 803/799 | 761/756 | 680/670 | 71.2/70.3 | 66.3/63.5 | 68.2/69.4 | 67.7/65.6 | 45.3/44.2 | 1.329/1.292 | 808.85
|  | 1000 V | 1670/1670 | 1700/1690 | 1660/1650 | 1470/1450 | 788/770 | 129/122 | 135/137 | 160/161 | 139/146 | 91.8/90.2 | 1.321/1.296 | 899.3
|  | 1100 V | 3000/3000 | 2930/2930 | 2820/2830 | 1540/1530 | 870/811 | 215/214 | 202/201 | 173/175 | 220/216 | 123/151 | 1.210/1.153 | 991.3
CM9594 | power setting (10.9) fixed by last consistency test | 850 V | 648/637 | 645/636 | 643/637 | 650/640 | 640/634 | 25.1/26.3 | 22.1/21.9 | 29.3/28.4 | 24.3/23.2 | 17.1/16.7 | 1.360/1.471 | 759.2 | 15.5 inches | yes | Nathan on 12/6/13 | retest | taken by Franz on 12/11/13
|  | 900 V | 850/844 | 855/845 | 847/833 | 800/793 | 680/670 | 71.8/68.1 | 72.2/69.5 | 76.3/74.4 | 70.1/67.2 | 51.1/48.2 | 1.360/1.288 | 805.65
|  | 1000 V | 1750/1720 | 1780/1730 | 1750/1710 | 1420/1420 | 814/817 | 154/148 | 161/155 | 144/141 | 143/142 | 86.2/88.1 | 1.207/1.323 | 898.45
|  | 1100 V | 3120/3070 | 3030/2990 | 2900/2860 | 1590/1630 | 803/842 | 233/231 | 259/255 | 203/206 | 208/216 | 134/113 | 1.153/1.180 | 990.45
CM9592 | power setting (10.9) fixed by last consistency test | 850 V | 471/469 | 482/486 | 472/473 | 463/465 | 406/405 | 43.6/44.9 | 53.2/53.7 | 56.2/55.5 | 41.3/42.5 | 44.6/45.3 | 1.362/1.379 | 759.65 | 15.5 inches | yes | Nathan on 12/6/13 | retest | taken by Franz on 12/11/13
|  | 900 V | 756/757 | 763/770 | 757/762 | 714/725 | 574/576 | 70.1/70.9 | 73.3/75.1 | 73.3/74.5 | 75.5/74.3 | 71.7/71.2 | 1.382/1.273 | 805.9
|  | 1000 V | 1560/1570 | 1550/1560 | 1520/1550 | 1350/1360 | 757/755 | 137/133 | 127/129 | 118/120 | 131/134 | 138/141 | 1.309/1.288 | 898.9
|  | 1100 V | 2700/2720 | 2740/2760 | 2700/2710 | 1690/1660 | 835/783 | 182/187 | 183/189 | 215/220 | 194/197 | 186/192 | 1.135/1.031 | 990.9
CM9593 | power setting (10.9) fixed by last consistency test | 850 V | 503/501 | 509/507 | 504/500 | 499/497 | 426/424 | 51.7/50.4 | 44.1/43.8 | 52.1/52.3 | 53.2/52.9 | 56.1/55.3 | 1.425/1.392 | 758.95 | 15.5 inches | yes | Nathan on 12/6/13 | retest | taken by Franz on 12/11/13
|  | 900 V | 785/781 | 804/803 | 798/797 | 780/781 | 583/585 | 74.6/74.0 | 70.5/71.8 | 65.8/64.5 | 74.2/73.1 | 67.7/66.0 | 1.382/1.261 | 805.15
|  | 1000 V | 1630/1640 | 1620/1610 | 1590/1580 | 1350/1340 | 785/773 | 120/121 | 133/131 | 144/139 | 158/155 | 144/141 | 1.314/1.200 | 898.05
|  | 1100 V | 2800/2780 | 2860/2850 | 2600/2610 | 1640/1610 | 857/833 | 212/214 | 200/204 | 197/191 | 185/183 | 196/203 | 1.156/1.143 | 990.2
CM9571 | power setting (10.9) fixed by last consistency test | 850 V | 542/538 | 530/525 | 521/512 | 495/496 | 436/428 | 52.4/51.5 | 42.3/44.1 | 45.0/44.7 | 51.8/50.5 | 45.4/45.5 | 1.440/1.487 | 759.6 | 15.5 inches | yes | Nathan on 12/6/13 | retest | taken by Franz on 12/11/13
|  | 900 V | 817/811 | 815/808 | 804/797 | 766/755 | 586/579 | 69.9/70.4 | 71.8/71.3 | 76.1/77.2 | 83.1/82.1 | 72.2/71.3 | 1.352/1.267 | 805.55
|  | 1000 V | 1680/1650 | 1630/1610 | 1640/1610 | 1430/1410 | 788/768 | 139/141 | 128/130 | 132/127 | 147/144 | 102/99.8 | 1.293/1.269 | 898.4
|  | 1100 V | 2920/2860 | 2850/2790 | 2710/2660 | 1670/1570 | 833/759 | 175/178 | 203/207 | 195/189 | 176/171 | 173/182 | 1.169/1.188 | 990.65
CM9386 | power setting (10.9) fixed by last consistency test | 850 V | 490/487 | 486/485 | 500/502 | 493/487 | 424/421 | 49.6/50.1 | 57.4/56.4 | 44.1/45.2 | 50.1/49.5 | 48.3/49.8 | 1.428/1.405 | 759.7 | 15.5 inches | yes | Nathan on 12/6/13 |  | taken by Franz on 12/11/13
|  | 900 V | 742/739 | 760/755 | 756/757 | 723/721 | 563/561 | 60.0/58.7 | 67.9/68.0 | 68.9/69.7 | 67.8/68.4 | 79.9/77.8 | 1.305/1.316 | 806.1
|  | 1000 V | 1520/1530 | 1520/1520 | 1530/1540 | 1360/1350 | 777/771 | 132/135 | 124/123 | 118/123 | 156/158 | 120/116 | 1.262/1.279 | 899.1
|  | 1100 V | 2710/2700 | 2720/2730 | 2660/2650 | 1630/1600 | 789/743 | 186/193 | 180/185 | 209/206 | 198/196 | 196/203 | 1.144/1.156 | 991.3
|  |  | 10kHz | 1 MHz | 2 MHz | 5 MHz | 10 MHz | 10kHz | 1 MHz | 2 MHz | 5 MHz | 10 MHz
CM9374 | power setting (10.9) fixed by last consistency test | 850 V | 508/510 | 494/499 | 483/486 | 471/475 | 415/417 | 51.9/52.0 | 46.1/45.2 | 55.0/55.4 | 50.2/50.9 | 46.5/47.1 | 1.404/1.392 | 758.7 | 15.5 inches | yes | Nathan on 12/6/13 |  | taken by Franz on 12/11/13
|  | 900 V | 763/773 | 748/752 | 743/747 | 706/715 | 566/572 | 65.2/66.1 | 75.5/77.2 | 69.2/71.2 | 73.6/72.7 | 63.4/63.6 | 1.342/1.331 | 804.95
|  | 1000 V | 1560/1580 | 1520/1530 | 1520/1530 | 1340/1340 | 777/762 | 136/141 | 142/148 | 122/129 | 130/129 | 148/142 | 1.252/1.192 | 898.05
|  | 1100 V | 2670/2700 | 2620/2640 | 2550/2580 | 1600/1640 | 787/820 | 186/182 | 203/206 | 187/184 | 183/191 | 187/177 | 1.103/1.167 | 989.85
CN7558 | power setting (10.9) fixed by last consistency test | 850 V | 410/417 | 399/400 | 385/394 | 378/387 | 339/346 | 36.8/36.1 | 44.2/43.0 | 41.0/39.1 | 38.2/40.1 | 26.0/29.4 | 1.403/1.376 | 758.05 | 15.5 inches | yes | Nathan on 12/6/13 |  | taken by Franz on 12/11/13
|  | 900 V | 587/598 | 590/589 | 586/595 | 578/589 | 487/487 | 62.1/60.4 | 51.3/52.2 | 54.1/53.6 | 54.7/53.1 | 66.5/64.6 | 1.409/1.347 | 804.5
|  | 1000 V | 1230/1250 | 1220/1230 | 1220/1230 | 1110/1120 | 733/761 | 110/114 | 100/99.1 | 97.0/96.7 | 120/124 | 114/116 | 1.372/1.342 | 897.45
|  | 1100 V | 2170/2210 | 2180/2200 | 2090/2120 | 1590/1640 | 778/812 | 171/178 | 170/167 | 166/170 | 156/152 | 151/159 | 1.194/1.171 | 989.55
CM9361 | power setting (10.9) fixed by last consistency test | 850 V | 496/493 | 492/488 | 493/490 | 478/475 | 418/414 | 50.1/50.0 | 49.2/49.0 | 49.0/49.0 | 54.5/53.9 | 49.0/48.7 | 1.416/1.425 | 760.4 | 15.5 inches | yes | Nathan on 12/9/13 |  | taken by Franz on 12/11/13
|  | 900 V | 757/752 | 745/740 | 744/739 | 720/715 | 586/581 | 72.2/72.4 | 73.3/72.4 | 74.4/73.6 | 75.0/74.5 | 83.1/82.7 | 1.377/1.366 | 806.95
|  | 1000 V | 1540/1530 | 1560/1550 | 1550/1540 | 1330/1320 | 720/715 | 133/135 | 131/133 | 134/135 | 165/163 | 145/144 | 1.333/1.266 | 899.5
|  | 1100 V | 2610/2600 | 2600/2590 | 2590/2580 | 1660/1650 | 812/746 | 210/210 | 203/205 | 197/197 | 205/204 | 182/204 | 1.174/1.137 | 992.05
CM9373 | power setting (10.9) fixed by last consistency test | 850 V | 475/473 | 483/481 | 477/475 | 455/452 | 402/399 | 47.3/46.8 | 48.3/49.1 | 49.2/49.3 | 49.2/48.5 | 49.4/49.9 | 1.420/1.416 | 759.95 | 15.5 inches | yes | Nathan on 12/9/13 |  | taken by Franz on 12/11/13
|  | 900 V | 718/714 | 723/720 | 725/722 | 683/680 | 559/558 | 67.1/67.5 | 64.1/64.2 | 67.1/67.3 | 69.5/70.2 | 70.1/69.3 | 1.335/1.366 | 805.95
|  | 1000 V | 1450/1440 | 1460/1460 | 1440/1440 | 1290/1290 | 750/755 | 115/114 | 126/126 | 125/124 | 119/120 | 126/123 | 1.312/1.300 | 898.15
|  | 1100 V | 2490/2470 | 2510/2490 | 2440/2420 | 1580/1600 | 800/824 | 171/171 | 183/181 | 174/171 | 159/162 | 161/164 | 1.085/1.169 | 990.85
CM9753 | power setting (10.9) fixed by last consistency test | 850 V | 563/566 | 567/570 | 555/559 | 525/530 | 445/449 | 52.9/53.7 | 52.1/52.4 | 52.7/52.4 | 51.5/52.3 | 54.1/53.7 | 1.392/1.433 | 759.1 | 15.5 inches | yes | Nathan on 12/9/13 |  | taken by Franz on 12/11/13
|  | 900 V | 833/838 | 828/831 | 828/832 | 781/786 | 597/599 | 77.8/76.3 | 75.0/74.6 | 75.1/75.7 | 73.5/73.1 | 71.4/71.7 | 1.368/1.342 | 805.2
|  | 1000 V | 1710/1720 | 1680/1690 | 1680/1690 | 1390/1400 | 762/750 | 131/130 | 136/137 | 133/132 | 125/127 | 136/137 | 1.275/1.257 | 898.2
|  | 1100 V | 2930/2940 | 2940/2960 | 2800/2820 | 1620/1590 | 766/704 | 204/208 | 207/210 | 203/205 | 209/207 | 176/195 | 1.1697/1.134 | 990.2
CM9752 | power setting (10.9) fixed by last consistency test | 850 V | 619/614 | 592/587 | 580/577 | 560/556 | 465/463 | 53.2/52.9 | 54.5/53.2 | 55.1/54.3 | 56.5/56.3 | 53.3/53.5 | 1.428/1.375 | 760.15 | 15.5 inches | yes | Nathan on 12/9/13 | Output 1 signal is less stable at 1100V, 5 MHz | taken by Franz on 12/11/13
|  | 900 V | 894/883 | 890/880 | 897/886 | 836/826 | 611/609 | 75.3/76.1 | 73.5/74.3 | 76.2/77.1 | 76.2/74.9 | 71.2/69.4 | 1.389/1.355 | 806.25
|  | 1000 V | 1850/1820 | 1810/1780 | 1790/1750 | 1410/1400 | 758/767 | 144/142 | 142/140 | 146/142 | 147/151 | 137/135 | 1.243/1.232 | 899.2
|  | 1100 V | 3120/3050 | 3050/2990 | 2910/2870 | 1400/1560 | 711/724 | 210/213 | 206/205 | 197/202 | 278/188 | 170/169 | 1.158/1.105 | 991
|  |  | 10kHz | 1 MHz | 2 MHz | 5 MHz | 10 MHz | 10kHz | 1 MHz | 2 MHz | 5 MHz | 10 MHz
CM9384 | power setting (10.9) fixed by last consistency test | 850 V | 466/462 | 465/461 | 470/465 | 444/441 | 397/392 | 51.3/49.3 | 48.4/47.7 | 50.3/50.4 | 51.8/50.3 | 46.1/47.0 | 1.375/1.433 | 759.5 | 15.5 inches | yes | Nathan on 12/9/13 |  | taken by Franz on 12/11/13
|  | 900 V | 725/715 | 719/707 | 720/711 | 690/684 | 555/549 | 69.4/68.8 | 70.8/69.4 | 63.5/64.4 | 64.6/65.2 | 66.1/65.3 | 1.356/1.347 | 805.7
|  | 1000 V | 1470/1440 | 1460/1430 | 1500/1480 | 1320/1310 | 760/754 | 123/121 | 126/128 | 123/126 | 129/125 | 128/125 | 1.261/1.242 | 898.75
|  | 1100 V | 2590/2530 | 2600/2550 | 2570/2520 | 1640/1580 | 809/741 | 202/199 | 201/198 | 199/196 | 182/180 | 149/163 | 1.057/1.123 | 990.7
CM9382 | power setting (10.9) fixed by last consistency test | 850 V | 473/469 | 500/496 | 496/491 | 490/486 | 413/410 | 49.5/48.3 | 48.6/49.1 | 48.4/48.8 | 48.1/47.5 | 47.3/46.5 | 1.432/1.400 | 760.85 | 15.5 inches | yes | Nathan on 12/10/13 |  | taken by Franz on 12/11/13
|  | 900 V | 766/758 | 755/750 | 765/759 | 740/737 | 571/570 | 70.6/69.5 | 70.3/69.8 | 70.1/70.7 | 71.5/70.4 | 70.4/70.3 | 1.349/1.372 | 806.75
|  | 1000 V | 1560/1550 | 1560/1550 | 1540/1530 | 1340/1340 | 743/755 | 126/124 | 132/135 | 125/124 | 115/114 | 127/126 | 1.281/1.276 | 899.6
|  | 1100 V | 2720/2710 | 2730/2700 | 2660/2640 | 1580/1660 | 705/790 | 185/182 | 182/179 | 177/180 | 183/184 | 193/172 | 1.146/1.118 | 991.65
CM9378 | power setting (10.9) fixed by last consistency test | 850 V | 460/461 | 465/463 | 459/457 | 438/437 | 392/393 | 45.1/44.4 | 47.6/46.8 | 47.7/47.3 | 44.1/44.0 | 46.1/45.5 | 1.477/1.442 | 760.15 | 15.5 inches | yes | Nathan on 12/10/13 |  | taken by Franz on 12/11/13
|  | 900 V | 677/675 | 695/693 | 702/701 | 673/672 | 553/554 | 64.9/65.1 | 63.8/63.7 | 65.2/65.0 | 60.2/59.8 | 60.6/60.3 | 1.395/1.388 | 806.3
|  | 1000 V | 1410/1400 | 1430/1420 | 1420/1410 | 1280/1270 | 762/753 | 120/118 | 120/121 | 118/116 | 117/116 | 115/114 | 1.283/1.311 | 899.1
|  | 1100 V | 2450/2440 | 2530/2510 | 2460/2440 | 1660/1610 | 813/726 | 177/176 | 180/182 | 183/186 | 185/187 | 172/195 | 1.173/1.146 | 991.15
CM9879 | power setting (10.9) fixed by last consistency test | 850 V | 386/383 | 386/385 | 372/373 | 355/356 | 320/321 | 39.5/38.9 | 40.3/39.7 | 35.0/34.3 | 34.7/34.2 | 35.8/35.1 | 1.4467/1.508 | 760.2 | 15.5 inches | yes | Nathan on 12/10/13 |  | taken by Franz on 12/11/13
|  | 900 V | 550/552 | 541/542 | 551/552 | 536/537 | 458/457 | 50.7/49.9 | 45.6/46.7 | 49.5/50.1 | 56.1/57.3 | 50.5/50.7 | 1.383/1.414 | 806.25
|  | 1000 V | 1140/1140 | 1130/1130 | 1130/1130 | 1020/1020 | 720/721 | 110/108 | 108/109 | 105/104 | 103/104 | 101/101 | 1.303/1.351 | 899.1
|  | 1100 V | 1980/1970 | 1960/1960 | 1950/1940 | 1590/1610 | 808/820 | 150/151 | 152/149 | 156/155 | 150/149 | 133/130 | 1.165/1.188 | 991.05
No label-7 | power setting (10.9) fixed by last consistency test | 850 V | 495/490 | 486/482 | 473/470 | 455/450 | 388/386 | 50.1/49.7 | 53.3/52.9 | 52.3/52.5 | 49.5/48.2 | 46.2/46.2 | 1.400/1.425 | 759.6 | 15.5 inches | yes | Nathan on 12/10/13 |  | taken by Franz on 12/11/13
|  | 900 V | 722/716 | 712/705 | 720/714 | 685/680 | 550/549 | 64.3/65.5 | 69.0/69.8 | 67.2/67.9 | 70.7/69.9 | 70.0/69.3 | 1.370/1.406 | 805.9
|  | 1000 V | 1510/1490 | 1470/1450 | 1480/1460 | 1280/1270 | 765/764 | 121/120 | 127/126 | 123/125 | 120/118 | 122/121 | 1.332/1.250 | 898.75
|  | 1100 V | 2610/2580 | 2590/2560 | 2510/2480 | 1640/1650 | 844/842 | 172/173 | 171/172 | 171/168 | 173/171 | 172/174 | 1.123/1.150 | 991
No label-8 | power setting (10.9) fixed by last consistency test | 850 V | 502/501 | 494/494 | 491/491 | 463/462 | 410/411 | 45.0/44.3 | 47.2/47.1 | 45.2/44.3 | 46.6/45.9 | 47.0/47.3 | 1.443/1.380 | 760.6 | 15.5 inches | yes | Nathan on 12/10/13 |  | taken by Franz on 12/11/13
|  | 900 V | 737/737 | 733/731 | 740/740 | 726/723 | 558/556 | 67.1/66.5 | 63.4/64.5 | 69.1/69.0 | 70.7/71.2 | 68.9/69.2 | 1.269/1.305 | 806.9
|  | 1000 V | 1520/1520 | 1510/1510 | 1520/1520 | 1340/1330 | 771/755 | 131/127 | 131/130 | 128/130 | 128/130 | 123/122 | 1.212/1.225 | 899.9
|  | 1100 V | 2600/2620 | 2660/2660 | 2590/2580 | 1680/1640 | 822/743 | 187/184 | 180/181 | 188/191 | 183/185 | 178/190 | 1.186/1.166 | 991.85
|  |  | 10kHz | 1 MHz | 2 MHz | 5 MHz | 10 MHz | 10kHz | 1 MHz | 2 MHz | 5 MHz | 10 MHz | 5MHz | 10kHz
No label-9 | power setting (10.9) fixed by last consistency test | 850 V | 477/475 | 490/486 | 480/477 | 454/452 | 400/399 | 49.1/49.3 | 48.9/49.3 | 48.1/48.5 | 50.4/50.1 | 47.3/48.1 | 1.445/1.432 | 760.1 | 15.5 inches | yes | Nathan on 12/10/13 |  | taken by Franz on 12/11/13
|  | 900 V | 724/720 | 732/728 | 741/738 | 698/694 | 553/552 | 69.1/68.5 | 69.5/68.7 | 69.7/70.0 | 71.2/72.1 | 71.2/71.7 | 1.350/1.361 | 806.3
|  | 1000 V | 1540/1520 | 1510/1500 | 1510/1490 | 1320/1320 | 756/751 | 129/130 | 128/127 | 125/126 | 126/125 | 115/116 | 1.308/1.272 | 899.15
|  | 1100 V | 2680/2660 | 2660/2640 | 2600/2580 | 1620/1640 | 707/764 | 185/183 | 182/181 | 183/185 | 173/176 | 184/173 | 1.121/1.175 | 991.1
CM9922 | 13.0 power | 850 V | 392/412 | 400/436 | 424/412 | 400/388 | 360/380 | 36.0/36.4 | 35.2/35.0 | 38.6/37.6 | 34.3/34.5 | 40.2/40.5 | 1.520/1.635 | 760.8 | 15.5 inches | yes | Natalie on 12/17/13 |  | taken by Franz on 12/19/13
|  | 900 V | 608/612 | 672/692 | 624/684 | 640/580 | 616/516 | 50.4/50.8 | 67.0/66.7 | 58.1/58.7 | 53.8/54.1 | 57.9/54.4 | 1.507/1.439 | 807.3
|  | 1000 V | 1.24/1.23 | 1.22/12.5 | 1.22/1.21 | 1.18/1.03 | 860/810 | 101/105 | 97.3/94.0 | 104/102 | 96.6/90/3 | 99.7/97.1 | 1.305/1.320 | 900.25
|  | 1100 V | 2.20/2.45 | 2.04/2.13 | 2.32/2.15 | 1.56/1.69 | 960/790 | 145/141 | 170/166 | 155/152 | 155/152 | 160/143 | 1.216/1.246 | 992.65
CM9910 | yes | 850 V | 512/476 | 440/460 | 352/452 | 288/412 | 256/380 | 35.6/40.4 | 47.0/42.8 | 32.2/40.5 | 34.2/44.3 | 36.0/36.2 | 1.197/1.045 | 759.8 | 15.5 inches | yes | Natalie on 12/17/13 | stopped testing because output 1 went CRAZY | BAD
|  | 900 V | 744/764 | 648/718 | 648/660 | 704/668 | 584/684 | 62.3/63.3 | 65.4/66.0 | 52.3/58.2 | 53.8/56.7 | 58.0/62.5 | 1.441/1.446 | 806.75 |  |  |  | see photo taken
|  | 1000 V | # | # | # | # | # | # | # | # | # | # | # | # |  |  |  | changing frequency did not help
|  | 1100 V | # | # | # | # | # | # | # | # | # | # | # | # |  |  |  | box for Vladimir
CM9920 | 13.4 power | 850 V | 496/476 | 488/452 | 416/478 | 424/452 | 376/372 | 40.1/37.5 | 44.6/40.8 | 38.7/36.5 | 38.8/37.7 | 43.0/42.8 | 1.470/1.366 | 759.35 | 15.5 inches | yes | Natalie on 12/18/13 |  | taken by Franz on 12/19/13
|  | 900 V | 648/668 | 600/604 | 736/732 | 600/644 | 576/516 | 60.2/61.8 | 58.2/58.9 | 57.0/61.1 | 61.4/60.9 | 57.8/55.9 | 1.381/1.391 | 806.05
|  | 1000 V | 1.26/1.43 | 1.41/1.44 | 1.56/1.37 | 1.24/1.29 | 790/720 | 111/113 | 105/109 | 113/113 | 126/127 | 109/111 | 1.303/1.241 | 899.05
|  | 1100 V | 2.58/2.27 | 2.38.2.42 | 2.38/2.57 | 1.52/1.55 | 680/870 | 171/170 | 159/158 | 163/161 | 160/159 | 143/144 | 1.173/1.205 | 991.25
CM9948 | yes | 850 V | 480/532 | 560/540 | 496/516 | 520/484 | 488/440 | 55.9/58.8 | 49.5/50.4 | 44.5/44.1 | 40.7/39.3 | 45.0/46.0 | 1.510/1.478 | 760.45 | 15.5 inches | yes | Natalie on 12/18/13 |  | taken by Franz on 12/19/13
|  | 900 V | 864/852 | 760/820 | 800/812 | 674/724 | 592/564 | 75.4/74.8 | 66.0/67.8 | 70.5/72.0 | 65.8/67.6 | 66.7/68.2 | 1.450/1.441 | 807.2
|  | 1000 V | 1.44/1.55 | 1.70/1.57 | 1.56/1.70 | 1.42/1.27 | 780/810 | 108/115 | 114/117 | 107/113 | 115/118 | 120/121 | 1.204/1.313 | 900.25
|  | 1100 V | 2.56/2.61 | 2.50/2.55 | 2.66/2.33 | 1.72/1.53 | 600/670 | 170/175 | 177/181 | 186/185 | 167/174 | 148/152 | 1.144/1.172 | 992.6
|  |  | 10kHz | 1 MHz | 2 MHz | 5 MHz | 10 MHz | 10kHz | 1 MHz | 2 MHz | 5 MHz | 10 MHz | 5MHz | 10kHz
CM9950 | yes | 850 V | 536/524 | 536/484 | 448/436 | 424/412 | 384/428 | 64.8/62.4 | 47.6/47.2 | 41.4/39.1 | 43.5/42.5 | 41.5/46.4 | 1.481/1.510 | 760.65 | 15.5 inches | yes | Natalie on 12/18/13 |  | taken by Franz on 12/19/13
|  | 900 V | 720/764 | 729/692 | 648/692 | 656/676 | 536/508 | 67.9/68.2 | 57.7/60.3 | 55.1/54.2 | 60.3/61.6 | 71.6/70.0 | 1.440/1.491 | 807.2
|  | 1000 V | 1.48/1.43 | 1.46/1.33 | 1.36/1.51 | 1.30/1.27 | 712/764 | 103/103 | 104/105 | 102/103 | 115/111 | 90.1/88.8 | 1.34/1.27 | 900.45
|  | 1100 V | 2.42/2.49 | 2.40/2.61 | 2.56/2.27 | 1.80/1.69 | 440/790 | 175/170 | 174/167 | 167/162 | 147/148 | 186/167 | 1.221/1.123 | 992.75
CM9952 | yes | 850 V | 440/436 | 464/428 | 456/468 | 464/452 | 408/412 | 39.0/38.0 | 37.3/35.5 | 40.0/39.7 | 38.4/37.0 | 42.3/41.4 | 1.500/1.407 | 758.9 | 15.5 inches | yes | Natalie on 12/18/13 |  | taken by Franz on 12/19/13
|  | 900 V | 736/716 | 664/620 | 720/700 | 688/660 | 528/564 | 56.0/55.9 | 60.9/60.0 | 56.7/56.8 | 54.8/54.1 | 61.7/61.9 | 1.410/1.362 | 811.35
|  | 1000 V | 1.42/1.29 | 1.52/1.39 | 1.32/1.27 | 1.38/1.11 | 760/790 | 90.3/85.9 | 109/103 | 110/107 | 103/97.1 | 140/138 | 1.280/1.219 | 898.6
|  | 1100 V | 2.29/2.38 | 2.28/2.41 | 2.66/2.41 | 1.64/1.69 | 880/770 | 163/156 | 171/167 | 176/178 | 171/166 | 213/212 | 1.198/1.231 | 990.95
CM9914 | 14.0 power | 850 V | 464/502 | 444/494 | 496/458 | 476/478 | 368/382 | 39.5/41.7 | 38.7/38.4 | 36.1/36.2 | 42.7/43.7 | 44.8/44.3 | 1.446/1.523 | 758.3 | 15.5 inches | yes | Natalie on 12/19/13 |  | taken by Franz on 12/19/13
|  | 900 V | 624/660 | 704/716 | 712/692 | 608/620 | 552/564 | 68.1/69.7 | 68.7/68.3 | 56.3/57.0 | 59.8/60.6 | 60.1/60.2 | 1.317/1.374 | 804.7
|  | 1000 V | 1.46/1.43 | 1.46/1.51 | 1.56/1.65 | 1.28/1.43 | 740/690 | 105/109 | 112/117 | 109/112 | 108/112 | 118/116 | 1.298/1.240 | 898.1
|  | 1100 V | 2.458/2.537 | 2.502/2.449 | 2.378/2.492 | 1.506/1.668 | 833/881 | 65.55/63.95 | 64.09/60.01 | 62.45/66.90 | 112.3/102.8 | 154/159 | 1.213/1.214 | 990.45
CM9949 | 14.0 power | 850 V | 553/516 | 531/534 | 530/547 | 518/479 | 409/433 | 19/19 | 17/17 | 19/18 | 16/16 | 24/23 | 1.436/1.475 | 758.8 | 15.5 inches | yes | Natalie on 12/20/13
|  | 900 V | 775/721 | 794/782 | 824/819 | 748/752 | 617/619 | 35/35 | 24/24 | 25/25 | 26/26 | 31/31 | 1.435/1.478 | 805.5
|  | 1000 V | 1.68/1.70 | 1.63/1.62 | 1.55/1.58 | 1.34/1.27 | 712/725 | 39/39 | 49/49 | 41/42 | 57/58 | 53/52 | 1.247/1.309 | 898.35
|  | 1100 V | 2.80/2.78 | 2.71/2.76 | 2.63/2.70 | 1.61/1.68 | 827/760 | 71/75 | 66/69 | 68/69 | 89/87 | 106/88 | 1.214/1.162 | 990.5
|  |  | 10kHz | 1 MHz | 2 MHz | 5 MHz | 10 MHz | 10kHz | 1 MHz | 2 MHz | 5 MHz | 10 MHz | 5MHz | 10kHz
CM9930 | yes | 850 V | 558/524 | 557/531 | 564/542 | 500/506 | 451/436 | 21/21 | 17/17 | 18/18 | 22/21 | 21/21 | 1.288/1.472 | 759.65 | 15.5 inches | yes | Natalie on 12/20/13
|  | 900 V | 831/867 | 831/815 | 837/895 | 793/804 | 575/581 | 24/24 | 28/28 | 29/29 | 31/31 | 28/28 | 1.423/1.403 | 807.45
|  | 1000 V | 1.69/1.67 | 1.68/1.72 | 1.65/1.62 | 1.35/1.43 | 761/685 | 37/37 | 49/49 | 43/44 | 48/48 | 57/57 | 1.247/1.273 | 899.3
|  | 1100 V | 2.71/2.78 | 2.74/2.72 | 2.68/2.60 | 1.65/1.55 | 800/845 | 66/67 | 67/69 | 64/63 | 83/81 | 217/79 | 1.269/1.171 | 991.4
CM9918 | yes | 850 V | 490/479 | 488/485 | 499/486 | 462/488 | 404/417 | 18/18 | 15/15 | 16/15 | 16/16 | 22/21 | 1.477/1.439 | 759.7 | 15.5 inches | yes | Natalie on 12/20/13
|  | 900 V | 799/780 | 770/765 | 767/749 | 736/718 | 585/567 | 23/23 | 25/25 | 24/24 | 26/27 | 22/23 | 1.415/1.400 | 806.55
|  | 1000 V | 1.54/1.51 | 1.57/1.52 | 1.57/1.51 | 1.27/1.39 | 755/771 | 40/44 | 37/39 | 45/47 | 56/57 | 54/53 | 1.295/1.319 | 899.4
|  | 1100 V | 2.57/2.67 | 2.55/2.45 | 2.46/2.62 | 1.51/1.65 | 783/912 | 58/59 | 62/62 | 65/63 | 77/77 | 117/83 | 1.230/1.139 | 991.65
CM9925 | yes | 850 V | 574/554 | 542/567 | 549/530 | 512/534 | 435/464 | 16/17 | 17/18 | 16/16 | 19/19 | 24/24 | 1.462/1.556 | 759.75 | 15.5 inches | yes | Natalie on 12/20/13
|  | 900 V | 862/834 | 852/883 | 893/881 | 796/776 | 565/618 | 27/27 | 27/28 | 28/27 | 27/27 | 19/19 | 1.444/1.408 | 806.7
|  | 1000 V | 1.70/1.67 | 1.74/1.71 | 1.71/1.66 | 1.40/1.41 | 788/750 | 44/47 | 36/37 | 45/44 | 59/58 | 60/58 | 1.239/1.250 | 899.45
|  | 1100 V | 2.85/2.92 | 2.84/2.79 | 2.68/2.89 | 1.72/1.62 | 737/705 | 76/77 | 74/77 | 63/62 | 70/67 | 115/73 | 1.248/1.146 | 991.65
|  |  | 10kHz | 1 MHz | 2 MHz | 5 MHz | 10 MHz | 10kHz | 1 MHz | 2 MHz | 5 MHz | 10 MHz | 5MHz | 10kHz
CM9921 | yes | 850 V | 516/507 | 497/487 | 502/522 | 459/461 | 429/402 | 19/19 | 16/16 | 19/18 | 19/18 | 24/22 | 1.345/1.431 | 759.5 | 15.5 inches | yes | Natalie on 12/20/13
|  | 900 V | 792/751 | 770/782 | 783/777 | 739/746 | 619/635 | 26/27 | 23/22 | 26/25 | 33/33 | 31/30 | 1.436/1.451 | 806.8
|  | 1000 V | 1.59/1.53 | 1.56/1.45 | 1.55/1.49 | 1.37/1.36 | 785/701 | 38/37 | 46/46 | 41/42 | 48/47 | 51/50 | 1.349/1.360 | 899.2
47 total? |  | 1100 V | 2.65/2.63 | 2.57/2.52 | 2.46/2.53 | 1.56/1.68 | 778/825 | 77/75 | 69/68 | 64/62 | 64/63 | 74/69 | 1.030/1.161 | 991.55
CM9878 | yes; power setting: 13.30 | 850 V | 495/492 | 502/498 | 492/489 | 480/474 | 420/413 | 15/15 | 18/18 | 17/18 | 16/16 | 17/17 | 1.503/1.496 | 760.05 | 15.5 inches | yes | Nathan on 1/2/14
|  | 900 V | 802/795 | 799/792 | 777/772 | 737/732 | 582/573 | 27/28 | 29/28 | 25/24 | 28/27 | 26/25 | 1.466/1.417 | 806.45
|  | 1000 V | 1550/1550 | 1540/1530 | 1530/1520 | 1330/1310 | 764/745 | 43/44 | 44/43 | 42/42 | 42/41 | 40/39 | 1.340/1.348 | 899.4
|  | 1100 V | 2690/2670 | 2680/2650 | 2600/2550 | 1630/1600 | 820/810 | 63/62 | 68/69 | 67/66 | 64/64 | 64/65 | 1.217/1.176 | 991.4
CM9875 | power setting (13.3) fixed by last consistency test | 850 V | 448/446 | 440/440 | 427/427 | 412/410 | 357/355 | 15/15 | 15/15 | 15/14 | 15/15 | 14/15 | 1.492/1.467 | 759.65 | 15.5 inches | yes | Nathan on 1/2/14
|  | 900 V | 646/644 | 637/637 | 631/632 | 615/614 | 521/518 | 21/20 | 22/21 | 20/20 | 21/21 | 21/21 | 1.381/1.350 | 805.8
|  | 1000 V | 1330/1320 | 1320/1310 | 1310/1300 | 1170/1160 | 753/750 | 41/43 | 42/41 | 40/40 | 44/42 | 38/39 | 1.328/1.313 | 898.75
|  | 1100 V | 2320/2290 | 2310/2280 | 2260/2220 | 1610/1620 | 840/844 | 63/62 | 55/55 | 65/63 | 58/56 | 60/60 | 1.189/1.200 | 990.7
CM9871 | power setting (13.3) fixed by last consistency test | 850 V | 414/413 | 412/412 | 409/408 | 389/385 | 337/335 | 14/14 | 14/14 | 14/15 | 15/16 | 15/14 | 1.530/1.493 | 758.85 | 15.5 inches | yes | Nathan on 1/2/14
|  | 900 V | 609/610 | 614/615 | 605/605 | 583/581 | 492/487 | 21/21 | 21/21 | 22/22 | 24/24 | 20/20 | 1.411/1.361 | 805.1
|  | 1000 V | 1290/1290 | 1290/1290 | 1270/1260 | 1150/1150 | 743/731 | 41/42 | 42/42 | 41/41 | 40/40 | 40/39 | 1.311/1.344 | 897.95
|  | 1100 V | 2280/2280 | 2280/2270 | 2250/2230 | 1630/1580 | 828/810 | 64/66 | 61/62 | 68/68 | 60/60 | 66/66 | 1.162/1.197 | 990.1
|  |  | 10kHz | 1 MHz | 2 MHz | 5 MHz | 10 MHz | 10kHz | 1 MHz | 2 MHz | 5 MHz | 10 MHz | 5MHz | 10kHz
CM9876 | power setting (13.3) fixed by last consistency test | 850 V | 538/544 | 530/537 | 527/532 | 500/503 | 427/428 | 18/19 | 17/17 | 15/15 | 19/19 | 19/19 | 1.470/1.447 | 760.25 | 15.5 inches | yes | Nathan on 1/2/14
|  | 900 V | 779/784 | 776/781 | 772/779 | 735/744 | 586/587 | 26/25 | 23/22 | 26/25 | 27/27 | 25/26 | 1.415/1.427 | 806.35
|  | 1000 V | 1560/1580 | 1540/1560 | 1530/1540 | 1330/1340 | 774/760 | 45/45 | 44/45 | 44/45 | 45/45 | 44/44 | 1.312/1.286 | 899.15
|  | 1100 V | 2610/2630 | 2580/2600 | 2530/2540 | 1650/1610 | 808/744 | 63/64 | 73/72 | 68/69 | 59/60 | 60/93 | 1.156/1.145 | 991.25
CM9870 | power setting (13.3) fixed by last consistency test | 850 V | 440/439 | 443/442 | 435/434 | 418/416 | 370/367 | 16/16 | 15/16 | 16/16 | 16/16 | 16/16 | 1.530/1.489 | 760.35 | 15.5 inches | yes | Nathan on 1/3/14
|  | 900 V | 666/664 | 665/664 | 664/663 | 637/635 | 536/532 | 24/24 | 23/23 | 24/24 | 23/23 | 21/21 | 1.458/1.442 | 806.7
|  | 1000 V | 1380/1370 | 1370/1360 | 1350/1340 | 1230/1220 | 745/742 | 43/44 | 41/40 | 40/40 | 40/41 | 40/40 | 1.306/1.331 | 899.6
|  | 1100 V | 2380/2360 | 2360/2340 | 2330/2300 | 1580/1590 | 758/808 | 64/63 | 65/64 | 62/63 | 60/61 | 91/65 | 1.235/1.266 | 991.65
CM9884 | power setting (13.3) fixed by last consistency test | 850 V | 432/440 | 431/437 | 425/431 | 400/405 | 351/353 | 13/14 | 15/15 | 14/14 | 15/15 | 14/14 | 1.499/1.502 | 758.6 | 15.5 inches | yes | Nathan on 1/3/14
|  | 900 V | 632/640 | 630/639 | 634/643 | 608/616 | 511/513 | 20/20 | 21/21 | 19/19 | 22/22 | 22/22 | 1.404/1.375 | 805
|  | 1000 V | 1320/1340 | 1300/1320 | 1290/1310 | 1170/1180 | 736/741 | 41/42 | 39/40 | 38/40 | 42/43 | 40/40 | 1.319/1.332 | 897.75
|  | 1100 V | 2300/2350 | 2290/2330 | 2260/2290 | 1570/1620 | 814/834 | 66/68 | 65/67 | 55/56 | 50/50 | 61/61 | 1.196/1.212 | 990
CM9877 | power setting (13.3) fixed by last consistency test | 850 V | 467/466 | 470/469 | 460/461 | 443/440 | 388/385 | 17/17 | 14/14 | 16/16 | 17/17 | 16/16 | 1.504/1.487 | 760 | 15.5 inches | yes | Nathan on 1/3/14
|  | 900 V | 694/693 | 692/692 | 687/688 | 658/658 | 542/540 | 21/21 | 24/24 | 23/23 | 24/25 | 22/23 | 1.438/1.363 | 806
|  | 1000 V | 1400/1400 | 1390/1390 | 1380/1370 | 1240/1240 | 754/764 | 45/45 | 40/41 | 44/43 | 40/41 | 41/40 | 1.358/1.349 | 898.75
|  | 1100 V | 2400/2390 | 2410/2400 | 2360/2340 | 1600/1640 | 755/827 | 58/58 | 61/60 | 60/61 | 58/59 | 84/56 | 1.211/1.237 | 990.75
|  |  | 10kHz | 1 MHz | 2 MHz | 5 MHz | 10 MHz | 10kHz | 1 MHz | 2 MHz | 5 MHz | 10 MHz | 5MHz | 10kHz
CM9881 | power setting (13.3) fixed by last consistency test | 850 V | 440/437 | 445/442 | 442/439 | 423/419 | 372/368 | 17/16 | 16/15 | 17/17 | 16/16 | 14/14 | 1.476/1.492 | 758.25 | 15.5 inches | yes | Nathan on 1/3/14
|  | 900 V | 677/672 | 682/678 | 675/671 | 655/650 | 538/531 | 21/22 | 23/23 | 22/22 | 23/23 | 23/22 | 1.427/1.395 | 804.5
|  | 1000 V | 1400/1390 | 1400/1390 | 1390/1380 | 1250/1240 | 744/739 | 40/41 | 40/40 | 40/41 | 44/43 | 44/43 | 1.334/1.351 | 897.2
|  | 1100 V | 2440/2430 | 2470/2460 | 2410/2390 | 1590/1590 | 775/802 | 63/65 | 65/65 | 63/61 | 63/62 | 72/62 | 1.186/1.199 | 989.2
CM9867 | power setting (13.3) fixed by last consistency test | 850 V | 443/443 | 434/433 | 423/423 | 409/408 | 356/355 | 16/16 | 14/14 | 15/15 | 16/16 | 14/14 | 1.524/1.509 | 758.75 | 15.5 inches | yes | Nathan on 1/3/14
|  | 900 V | 649/650 | 643/644 | 640/641 | 617/618 | 520/517 | 22/22 | 23/23 | 21/21 | 24/24 | 22/23 | 1.383/1.409 | 805.05
|  | 1000 V | 1370/1370 | 1340/1330 | 1330/1330 | 1190/1190 | 747/748 | 40/40 | 40/40 | 41/41 | 43/42 | 40/40 | 1.299/1.318 | 897.9
|  | 1100 V | 2400/2390 | 2390/2370 | 2340/2320 | 1630/1650 | 839/848 | 67/68 | 65/64 | 64/65 | 58/59 | 61/60 | 1.159/1.216 | 990.05
CM9372 | power setting (13.3) fixed by last consistency test | 850 V | 579/577 | 566/565 | 562/561 | 543/541 | 457/454 | 19/19 | 20/20 | 21/21 | 20/20 | 20/20 | 1.466/1.470 | 759.75 | 15.5 inches | yes | Nathan on 1/3/14
|  | 900 V | 868/862 | 860/852 | 862/856 | 812/809 | 613/611 | 27/26 | 29/28 | 29/29 | 25/25 | 28/28 | 1.402/1.427 | 806.05
|  | 1000 V | 1780/1760 | 1760/1740 | 1740/1730 | 1430/1430 | 750/765 | 50/50 | 52/52 | 53/54 | 47/48 | 48/49 | 1.319/1.286 | 898.95
|  | 1100 V | 3010/3000 | 2980/2970 | 2870/2870 | 1560/1620 | 690/780 | 80/81 | 87/88 | 71/72 | 72/67 | 125/68 | 1.200/1.189 | 991.05
CM9890 | power setting (13.3) fixed by last consistency test | 850 V | 493/482 | 481/473 | 475/468 | 455/451 | 388/386 | 17/17 | 16/16 | 16/17 | 17/17 | 17/17 | 1.488/1.507 | 758.05 | 15.5 inches | yes | Nathan on 1/3/14
|  | 900 V | 735/721 | 732/720 | 726/716 | 689/686 | 548/550 | 25/24 | 23/22 | 24/25 | 23/23 | 24/24 | 1.427/1.415 | 804.6
|  | 1000 V | 1550/1520 | 1530/1500 | 1510/1490 | 1300/1300 | 743/744 | 47/48 | 47/46 | 48/49 | 48/47 | 45/45 | 1.332/1.318 | 897.45
|  | 1100 V | 2710/2660 | 2700/2650 | 2600/2570 | 1600/1630 | 803/808 | 69/70 | 61/62 | 60/61 | 64/64 | 61/62 | 1.154/1.166 | 989.75
|  |  | 10kHz | 1 MHz | 2 MHz | 5 MHz | 10 MHz | 10kHz | 1 MHz | 2 MHz | 5 MHz | 10 MHz | 5MHz | 10kHz
CM9888 | power setting (13.3) fixed by last consistency test | 850 V | 482/433 | 480/430 | 473/425 | 452/406 | 398/357 | 17/16 | 15/14 | 15/14 | 19/17 | 17/15 | 1.507/1.510 | 759.95 | 15.5 inches | yes | Nathan on 1/3/14 | difference between amplitude of signal 1 and 2 is larger than usual
|  | 900 V | 723/646 | 727/651 | 723/644 | 693/618 | 557/502 | 26/24 | 27/25 | 27/23 | 22/19 | 21/19 | 1.423/1.442 | 806.05
|  | 1000 V | 1500/1330 | 1500/1340 | 1480/1340 | 1290/1190 | 745/681 | 47/41 | 37/37 | 41/40 | 44/42 | 41/38 | 1.310/1.247 | 898.9
|  | 1100 V | 2600/2430 | 2580/2390 | 2520/2300 | 1570/1470 | 752/726 | 61/58 | 70/69 | 60/59 | 70/65 | 75/55 | 1.174/1.207 | 990.95
CM9889 | 14.0 power | 850 V | 470/475 | 470/467 | 454/448 | 443/441 | 403/411 | 17/17 | 16/16 | 16/16 | 15/16 | 21/21 | 1.580/1.423 | 760.1 | 15.5 inches | yes | Natalie on 1/4/14
|  | 900 V | 693/699 | 740/738 | 708/711 | 702/715 | 581/563 | 28/28 | 22/23 | 23/23 | 25/25 | 23/24 | 1.501/1.426 | 806.6
|  | 1000 V | 1.496/1.483 | 1.479/1.499 | 1.46/1.473 | 1.237/1.291 | 751/730 | 37/39 | 40/43 | 40/42 | 45/44 | 42/41 | 1.289/1.314 | 899.9
|  | 1100 V | 2.606/2.457 | 2.577/2.580 | 2.410/2.455 | 1.683/1.769 | 765/767 | 67/65 | 65/65 | 62/63 | 72/73 | 74/58 | 1.159/1.220 | 992.2
CM9387 | 12.8 power | 850 V | 566/531 | 535/550 | 511/507 | 497/490 | 427/441 | 17/25 | 17/17 | 18/17 | 16/16 | 23/23 | 1.412/1.466 | 759.1 | 15.5 inches | yes | Natalie on 1/4/14
|  | 900 V | 849/866 | 804/796 | 828/822 | 762/772 | 621/630 | 30/28 | 28/27 | 29/28 | 30/30 | 35/35 | 1.382/1.388 | 806.15
|  | 1000 V | 1.685/1.710 | 1.770/1.635 | 1.655/1.630 | 1.392/1.315 | 819/803 | 53/52 | 55/53 | 53/52 | 57/55 | 46/45 | 1.231/1.322 | 899.9
|  | 1100 V | 2.874/2.953 | 2.879/2.764 | 2.750/2.755 | 1.502/1.645 | 785/811 | 74/71 | 78/76 | 62/61 | 75/77 | 114/80 | 1.150/1.210 | 991.1
CM9743 | yes previous | 850 V | 538/546 | 553/563 | 544/536 | 538/508 | 449/432 | 31/31 | 20/21 | 17/17 | 23/24 | 29/29 | 1.461/1.451 | 759.75 | 15.5 inches | yes | Natalie on 1/4/14
|  | 900 V | 877/867 | 843/823 | 834/815 | 751/760 | 642/636 | 27/28 | 31/31 | 30/30 | 31/31 | 39/40 | 1.421/1.403 | 806.45
|  | 1000 V | 1.802/1.769 | 1.666/1.755 | 1.757/1.655 | 1.369/1.335 | 757/775 | 54/54 | 55/54 | 63/63 | 54/53 | 51/50 | 1.247/1.241 | 899.6
|  | 1100 V | 2.989/2.926 | 2.719/2.826 | 2.815/2.880 | 1.642/1.553 | 867/878 | 83/84 | 86/84 | 79/78 | 74/74 | 108/93 | 73/74 | 991.8
|  |  | 10kHz | 1 MHz | 2 MHz | 5 MHz | 10 MHz | 10kHz | 1 MHz | 2 MHz | 5 MHz | 10 MHz | 5MHz | 10kHz
CM9392 | yes previous | 850 V | 512/496 | 506/498 | 470/488 | 487/479 | 392/393 | 19/20 | 18/18 | 18/19 | 20/21 | 21/22 | 1.484/1.390 | 759.6 | 15.5 inches | yes | Natalie on 1/4/14
|  | 900 V | 747/746 | 784/779 | 748/751 | 732/723 | 547/561 | 26/26 | 30/30 | 23/23 | 22/22 | 27/28 | 1.467/1.429 | 806.4
|  | 1000 V | 1.548/1.522 | 1.509/1.529 | 1.494/1.546 | 1.351/1.359 | 768/771 | 42/45 | 39/42 | 43/47 | 50/48 | 40/40 | 1.319/1.250 | 899.45
|  | 1100 V | 2.493/2.677 | 2.712/2.503 | 2.531/2.471 | 1.460/1.659 | 631/657 | 66/69 | 83/83 | 68/70 | 63/66 | 90/80 | 1.228/1.187 | 991.8
test 300 |  | 850 V | # | # | # | # | # | # | # | # | # | # | # | #
|  | 900 V | # | # | # | # | # | # | # | # | # | # | # | #
|  | 1000 V | # | # | # | # | # | # | # | # | # | # | # | #
|  | 1100 V | # | # | # | # | # | # | # | # | # | # | # | #
test 300 |  | 850 V | # | # | # | # | # | # | # | # | # | # | # | #
|  | 900 V | # | # | # | # | # | # | # | # | # | # | # | #
|  | 1000 V | # | # | # | # | # | # | # | # | # | # | # | #
|  | 1100 V | # | # | # | # | # | # | # | # | # | # | # | #
test 300 |  | 850 V | # | # | # | # | # | # | # | # | # | # | # | #
|  | 900 V | # | # | # | # | # | # | # | # | # | # | # | #
|  | 1000 V | # | # | # | # | # | # | # | # | # | # | # | #
|  | 1100 V | # | # | # | # | # | # | # | # | # | # | # | #
|  |  | 10kHz | 1 MHz | 2 MHz | 5 MHz | 10 MHz | 10kHz | 1 MHz | 2 MHz | 5 MHz | 10 MHz | 5MHz | 10kHz
test 300 |  | 850 V | # | # | # | # | # | # | # | # | # | # | # | #
|  | 900 V | # | # | # | # | # | # | # | # | # | # | # | #
|  | 1000 V | # | # | # | # | # | # | # | # | # | # | # | #
|  | 1100 V | # | # | # | # | # | # | # | # | # | # | # | #
test 300 |  | 850 V | # | # | # | # | # | # | # | # | # | # | # | #
|  | 900 V | # | # | # | # | # | # | # | # | # | # | # | #
|  | 1000 V | # | # | # | # | # | # | # | # | # | # | # | #
|  | 1100 V | # | # | # | # | # | # | # | # | # | # | # | #
test 300 |  | 850 V | # | # | # | # | # | # | # | # | # | # | # | #
|  | 900 V | # | # | # | # | # | # | # | # | # | # | # | #
|  | 1000 V | # | # | # | # | # | # | # | # | # | # | # | #
|  | 1100 V | # | # | # | # | # | # | # | # | # | # | # | #
test 300 |  | 850 V | # | # | # | # | # | # | # | # | # | # | # | #
|  | 900 V | # | # | # | # | # | # | # | # | # | # | # | #
|  | 1000 V | # | # | # | # | # | # | # | # | # | # | # | #
|  | 1100 V | # | # | # | # | # | # | # | # | # | # | # | #
|  |  | 10kHz | 1 MHz | 2 MHz | 5 MHz | 10 MHz | 10kHz | 1 MHz | 2 MHz | 5 MHz | 10 MHz | 5MHz | 10kHz
test 300 |  | 850 V | # | # | # | # | # | # | # | # | # | # | # | #
|  | 900 V | # | # | # | # | # | # | # | # | # | # | # | #
|  | 1000 V | # | # | # | # | # | # | # | # | # | # | # | #
|  | 1100 V | # | # | # | # | # | # | # | # | # | # | # | #
test 300 |  | 850 V | # | # | # | # | # | # | # | # | # | # | # | #
|  | 900 V | # | # | # | # | # | # | # | # | # | # | # | #
|  | 1000 V | # | # | # | # | # | # | # | # | # | # | # | #
|  | 1100 V | # | # | # | # | # | # | # | # | # | # | # | #
test 300 |  | 850 V | # | # | # | # | # | # | # | # | # | # | # | #
|  | 900 V | # | # | # | # | # | # | # | # | # | # | # | #
|  | 1000 V | # | # | # | # | # | # | # | # | # | # | # | #
|  | 1100 V | # | # | # | # | # | # | # | # | # | # | # | #
test 300 |  | 850 V | # | # | # | # | # | # | # | # | # | # | # | #
|  | 900 V | # | # | # | # | # | # | # | # | # | # | # | #
|  | 1000 V | # | # | # | # | # | # | # | # | # | # | # | #
|  | 1100 V | # | # | # | # | # | # | # | # | # | # | # | #
|  |  | 10kHz | 1 MHz | 2 MHz | 5 MHz | 10 MHz | 10kHz | 1 MHz | 2 MHz | 5 MHz | 10 MHz | 5MHz | 10kHz
test 300 |  | 850 V | # | # | # | # | # | # | # | # | # | # | # | #
|  | 900 V | # | # | # | # | # | # | # | # | # | # | # | #
|  | 1000 V | # | # | # | # | # | # | # | # | # | # | # | #
|  | 1100 V | # | # | # | # | # | # | # | # | # | # | # | #
test 300 |  | 850 V | # | # | # | # | # | # | # | # | # | # | # | #
|  | 900 V | # | # | # | # | # | # | # | # | # | # | # | #
|  | 1000 V | # | # | # | # | # | # | # | # | # | # | # | #
|  | 1100 V | # | # | # | # | # | # | # | # | # | # | # | #
test 300 |  | 850 V | # | # | # | # | # | # | # | # | # | # | # | #
|  | 900 V | # | # | # | # | # | # | # | # | # | # | # | #
|  | 1000 V | # | # | # | # | # | # | # | # | # | # | # | #
|  | 1100 V | # | # | # | # | # | # | # | # | # | # | # | #