## Sheet1
Hall-D Broadband Hodoscope COUNTER ASSEMBLY
Label: [size]-[item] (for inventory)
LG&SC#: received packing [bag#]-[item#]
PMT#: Hamamatsu serial#
Counter#: installed position number in support frame (1…274) | [6 empty] | (see https://halldweb1.jlab.org/wiki/index.php/Specification_for_broad-band_tagging_hodoscope )
Position & angle alignment: 17 units surveyed by JLAB alignment group; all others checked using 1mm-grid paper
Visual inspection at CUA: BN, JP
Visual inspection results: to test (=check with LED in light-tight box using R9800 PMT); okay (=minor imperfection, try to polish)
Gluing: refer to procedure D00000-19-01-P002
Note: resin and hardener have to kit match - discard both components when one runs out!
RTV615A kit (1): resin bottle #038096, hardener jar #038100 | [5 empty] | 13KWFA006A kit (2): resin bottle #038096, hardener jar #038100
Glue joint inspection after 6hr curing process!
Label | Size | LG&SC# | PMT# | link to PMT test results | Installed Counter# | Position align | Angle align | Surface notes | Glue Notes | Inspection results (FK) | Imperfections fixed? (date, initial) | gluing (date, initial) | glue joint inspection (initial) | mylar, tedlar wrapping (date, initial) | collar attachment (date, initial) | mu-metal wrapping (date, initial) | HV test (date, initial) | HV test: light tight? | HV test: spectrum saved? | store for shipping (date, initial) | install in frame (date, initial) | counter position # | final HV test (dete, initial) | final HV test: spectrum saved?
3-1-1 | 3mm | 1_1 |  |  | 0 | 0.264 mm | 0.219° | Scintilator nicked | Good | okay
3-1-2 | 3mm | 1_2 |  |  |  | 0.108 mm | 0.219° | bubbles in corner | Good | to test
3-1-3 | 3mm | 1_3 |  |  |  | 0.053 mm | 0.104° | Good | Very little bubbles | good
3-1-4 | 3mm | 1_4 | CM 9451 |  |  | < 0.5 mm | <1° | Good | Good | good |  | 2/20 fk | jp | 2/24 jp | 3/3 jp | 3/10 jp
3-1-5 | 3mm | 1_5 | CM 9235 |  |  | < 0.5 mm | <1° | Good | Good | good |  | 2/4 fk | bn | 2/5 bn | 2/19 jp | 3/4 jp
3-1-6 | 3mm | 1_6 | CM 9229 |  |  | < 0.5 mm | <1° | Good | Good | good |  | 2/4 fk | bn | 2/5 bn | 2/19 jp | 2/26 jp
3-1-7 | 3mm | 1_7 | CM 9232 |  |  | < 0.5 mm | <1° | Good | Good | good |  | 2/4 fk | bn | 2/5 bn | 2/19 jp | 2/26 jp
3-1-8 | 3mm | 1_8 | CM 9233 |  |  | < 0.5 mm | <1° | Good | Good | good |  | 2/4 fk | bn | 2/5 bn | 2/19 jp | 2/26 jp
3-1-9 | 3mm | 1_9 |  |  |  | < 0.5 mm | ~1-1.5° | Good | Large bubble seen at the corner | to test
3-1-10 | 3mm | 1_10 | CM 9833 |  |  | < 0.5 mm | <1° | Good | Good | good |  | 2/18 fk | jp | 2/20 jp | 3/3 jp | 3/10 jp
3-1-11 | 3mm | 1_11 | CM 9244 |  |  | < 0.5 mm | <1° | Good | Good | good |  | 2/20 fk | jp | 2/24 jp | 3/3 jp | 3/10 jp
3-1-12 | 3mm | 1_12 | CM 9383 |  |  | < 0.5 mm | <1° | Tiny chip in scintilator corner | Good | to test |  | 2/20 fk | jp | 2/24 jp | 3/3 jp | 3/10 jp
3-1-13 | 3mm | 1_13 | CM 9241 |  |  | < 0.5 mm | <1° | Good | Good | good |  | 2/4 fk | bn | 2/5 bn | 2/19 jp | 2/26 jp
3-1-14 | 3mm | 1_14 | CM 9248 |  |  | < 0.5 mm | <1° | Good | Good | good |  | 2/20 fk | jp | 2/24 jp | 3/3 jp | 3/10 jp
3-1-15 | 3mm | 1_15 | CM 9413 |  |  | < 0.5 mm | <1° | Good | Good | good |  | 2/20 fk | jp | 2/24 jp | 3/3 jp | 3/10 jp
4-1-16 | 3mm | 1_16 | CM 9471 |  |  | < 0.5 mm | <1° | Good | Little cloudy | to test |  | 2/20 fk | jp | 2/24 jp | 3/3 jp | 3/10 jp
3-1-17 | 3mm | 1_17 | CM 9393 |  |  | < 0.5 mm | <1° | Good | Good | good |  | 11/26 fk | bn | 11/28 bn | 1/11 jp | 1/16 jp
3-1-18 | 3mm | 1_18 | CM 9942 |  |  | < 0.5 mm | <1° | Good | Good | good |  | 2/4 fk | bn | 2/5 bn | 2/19 jp | 3/4 jp
3-2-1 | 3mm | 2_1 | CM 9592 |  |  | < 0.5 mm | <1° | Good | Messy glue | to test |  | 12/10 fk | fk | 12/12 bn | 1/11 jp | 1/16 jp
3-2-2 | 3mm | 2_2 |  |  |  | < 0.5 mm | <1° | Good | Good | good
3-2-3 | 3mm | 2_3 |  |  |  | < 0.5 mm | <1° | Good | Good | good
3-2-4 | 3mm | 2_4 |  |  |  | < 0.5 mm | <1° | Good | Good | good
3-2-5 | 3mm | 2_5 |  |  |  | < 0.5 mm | <1° | Good | Good | good
3-2-6 | 3mm | 2_6 |  |  |  | < 0.5 mm | ~1.5° | Good | not perfectly glued at corner | to test
3-2-7 | 3mm | 2_7 |  |  |  | < 0.5 mm | <1° | Good | Good | good
3-2-8 | 3mm | 2_8 |  |  |  | < 0.5 mm | <1° | Good | Small bubble seen | to test
3-2-9 | 3mm | 2_9 |  |  |  | < 0.5 mm | <1° | Good | Good | good
3-2-10 | 3mm | 2_10 |  |  |  | < 0.5 mm | <1° | Good | Good | good
3-2-11 | 3mm | 2_11 |  |  |  | < 0.5 mm | <1° | Good | Good | good
3-2-12 | 3mm | 2_12 |  |  |  | < 0.5 mm | <1° | Good | Good | good
3-2-13 | 3mm | 2_13 |  |  |  | < 0.5 mm | <1° | Good | Good | good
3-2-14 | 3mm | 2_14 |  |  |  | < 0.5 mm | <1° | Good | Good | good
3-2-15 | 3mm | 2_15 |  |  |  | < 0.5 mm | <1° | Good | Good | good
3-2-16 | 3mm | 2_16 |  |  |  | < 0.5 mm | <1° | Bubble in the scintilator | Good | to test
3-2-17 | 3mm | 2_17 |  |  |  | < 0.5 mm | <1° | Good | Good | good
3-2-18 | 3mm | 2_18 | CM 9756 |  |  | < 0.5 mm | <1° | Good | Good | good |  | 11/26 fk | bn | 11/28 bn | 12/5 jp | 1/8 jp
3-2-19 | 3mm | 2_19 |  |  |  | < 0.5 mm | <1° | Good | Good | good
3-3-1 | 3mm | 3_1 |  |  |  | < 0.5 mm | <1° | Good | Completely white | BAD
3-3-2 | 3mm | 3_2 |  |  |  | < 0.5 mm | <1° | Good | Good | good
3-3-3 | 3mm | 3_3 |  |  |  | < 0.5 mm | <1° | Good | Good | good
3-3-4 | 3mm | 3_4 |  |  |  | < 0.5 mm | <1° | Good | Small bubble seen at the corner | good
3-3-5 | 3mm | 3_5 |  |  |  | < 0.5 mm | <1° | Good | Good | good
3-3-6 | 3mm | 3_6 |  |  |  | < 0.5 mm | <1° | Good | Good | good
3-3-7 | 3mm | 3_7 |  |  |  | < 0.5 mm | <1° | Good | Small bubble at one corner and edge | to test
3-3-8 | 3mm | 3_8 |  |  |  | < 0.5 mm | <1° | Good | Small bubbles all around the corner | to test
3-3-9 | 3mm | 3_9 |  |  |  | < 0.5 mm | ~1-1.5° | Small scratch on scintilator | Good | to test
3-3-10 | 3mm | 3_10 |  |  |  | < 0.5 mm | <1° | Good | Slightly cloudy | to test
3-3-11 | 3mm | 3_11 | no label 1 |  |  | < 0.5 mm | <1° | Good | Slightly cloudy | to test |  | 11/26 fk | bn | 11/28 bn | 12/5 fk | 12/16 bn
3-3-12 | 3mm | 3_12 |  |  |  | < 0.5 mm | <1° | Good | Good | good
3-3-13 | 3mm | 3_13 | no label 5 |  |  | < 0.5 mm | <1° | Small bubbles in Light guide | Good | to test |  | 11/21 fk | fk | 11/25 fk | 12/5 fk | 12/16 bn
3-3-14 | 3mm | 3_14 |  |  |  | < 0.5 mm | <1° | Good | Good | good
3-3-15 | 3mm | 3_15 |  |  |  | < 0.5 mm | <1° | Good | Good | good
3-3-16 | 3mm | 3_16 |  |  |  | < 0.5 mm | <1° | Good | Good | good
3-3-17 | 3mm | 3_17 |  |  |  | < 0.5 mm | <1° | Good | Good | good
3-3-18 | 3mm | 3_18 |  |  |  | < 0.5 mm | <1° | Good | Good | good
3-3-19 | 3mm | 3_19 |  |  |  | < 0.5 mm | <1° | Good | Good | good
3-4-1 | 3mm | 4_1 |  |  |  | < 0.5 mm | <1° | Good | Good | good
3-4-2 | 3mm | 4_2 |  |  |  | < 0.5 mm | <1° | Good | Good | good
3-4-3 | 3mm | 4_3 |  |  |  | < 0.5 mm | <1° | Good | Good | good
3-4-4 | 3mm | 4_4 |  |  |  | < 0.5 mm | <1° | Good | Good | good
3-4-5 | 3mm | 4_5 |  |  |  | < 0.5 mm | <1° | Good | Small bubble seen in corner | good
3-4-6 | 3mm | 4_6 | CM 9752 |  |  | < 0.5 mm | <1° | Good | Slightly cloudy | to test |  | 12/3 fk | bn | 12/5 bn | 12/10 jp | 12/16 bn
3-4-7 | 3mm | 4_7 |  |  |  | < 0.5 mm | <1° | Good | Good | good
3-4-8 | 3mm | 4_8 | CM 9565 |  |  | < 0.5 mm | <1° | Good | Slightly cloudy | to test |  | 11/21 fk | bn | 11/25 bn | 12/5 jp | 12/16 bn
3-4-9 | 3mm | 4_9 |  |  |  | < 0.5 mm | <1° | Good | Good | good
3-4-10 | 3mm | 4_10 |  |  |  | < 0.5 mm | <1° | Good | Good | good
3-4-11 | 3mm | 4_11 |  |  |  | < 0.5 mm | <1° | Good | Bubble seen in corner | good
3-4-12 | 3mm | 4_12 |  |  |  | < 0.5 mm | <1° | Good | Tiny bubbles in corner | good
3-4-13 | 3mm | 4_13 |  |  |  | < 0.5 mm | <1° | Good | Tiny bubbles in corner and at center | to test
3-4-14 | 3mm | 4_14 |  |  |  | < 0.5 mm | <1° | Good | Good | good
3-4-15 | 3mm | 4_15 | CM 9635 |  |  | < 0.5 mm | ~1.5° | Small scratch on light guide | Slightly cloudy | to test |  | 11/26 fk | bn | 11/28 bn | 12/5 jp | 12/17 jp
3-4-16 | 3mm | 4_16 |  |  |  | < 0.5 mm | <1° | Good | Good | good
3-4-17 | 3mm | 4_17 |  |  |  | < 0.5 mm | <1° | Good | Good | good
3-4-18 | 3mm | 4_18 |  |  |  | < 0.5 mm | <1° | Good | Good | good
4-5-1 | 4mm | 5_1 |  |  |  | < 0.5 mm | <1° | Good | Good | good
4-5-2 | 4mm | 5_2 |  |  |  | < 0.5 mm | <1° | Good | Good | good
4-5-3 | 4mm | 5_3 |  |  |  | < 0.5 mm | <1° | Good | Good | good
4-5-4 | 4mm | 5_4 |  |  |  | < 0.5 mm | <1° | Good | Good | good
4-5-5 | 4mm | 5_5 | no label 3 |  |  | < 0.5 mm | <1° | Tiny chip in scintilator corner | Good | polish | 11/26 fk - redo
4-5-6 | 4mm | 5_6 |  |  |  | < 0.5 mm | <1° | Good | Good | good
4-5-7 | 4mm | 5_7 |  |  |  | < 0.5 mm | <1° | Good | Slightly cloudy | to test
4-5-8 | 4mm | 5_8 |  |  |  | < 0.5 mm | <1° | Good | Good | good
4-5-9 | 4mm | 5_9 | CM 9832 |  |  | < 0.5 mm | <1° | Little scratch on the light guide | Good | to test
4-5-10 | 4mm | 5_10 |  |  |  | < 0.5 mm | <1° | Good | Good | good
4-5-11 | 4mm | 5_11 |  |  |  | < 0.5 mm | <1° | Good | Good | good
4-5-12 | 4mm | 5_12 | no label 6 |  |  | < 0.5 mm | <1° | Tiny chip in scintilator corner | Good | good
4-5-13 | 4mm | 5_13 |  |  |  | < 0.5 mm | <1° | Good | Good | good
4-5-14 | 4mm | 5_14 |  |  |  | < 0.5 mm | <1° | Good | Small spot seen | good
4-5-15 | 4mm | 5_15 | no label 8 |  |  | < 0.5 mm | <1° | Tiny chip in scintilator corner | Good | good |  | 12/3 fk | bn | 12/4 bn | 12/5 jp | 12/17 jp
4-6-1 | 4mm | 6_1 |  |  |  | 0.068 mm | 0.013° | Good | Good | good
4-6-2 | 4mm | 6_2 |  |  |  | 0.129 mm | 0.144° | Good | Messy glue | to test
4-6-3 | 4mm | 6_3 |  |  |  | 0.621 mm | 0.107° | Good | Good | good
4-6-4 | 4mm | 6_4 |  |  |  | < 0.5 mm | <1° | Good | Small clouds | to test
4-6-5 | 4mm | 6_5 |  |  |  | < 0.5 mm | <1° | Good | Good | good
4-6-6 | 4mm | 6_6 |  |  |  | < 0.5 mm | <1° | Good | Good | good
4-6-7 | 4mm | 6_7 |  |  |  | < 0.5 mm | <1° | Dirty | Good | polish
4-6-8 | 4mm | 6_8 |  |  |  | < 0.5 mm | <1° | Dirty | Good | polish
4-6-9 | 4mm | 6_9 |  |  |  | < 0.5 mm | <1° | Dirty | Good | polish
4-6-10 | 4mm | 6_10 |  |  |  | < 0.5 mm | <1° | Dirty | Good | polish
4-6-11 | 4mm | 6_11 |  |  |  | < 0.5 mm | <1° | Dirty | Good | polish
4-6-12 | 4mm | 6_12 |  |  |  | < 0.5 mm | <1° | Good | Good | good
4-6-13 | 4mm | 6_13 | CM 9694 |  |  | < 0.5 mm | <1° | Good | Cloudy | bad
4-6-14 | 4mm | 6_14 |  |  |  | < 0.5 mm | <1° | Good | Bubble | to test
4-6-15 | 4mm | 6_15 |  |  |  | < 0.5 mm | <1° | Dirty | Good | polish
4-6-16 | 4mm | 6_16 |  |  |  | < 0.5 mm | <1° | Dirty | Good | polish
4-7-1 | 4mm | 7_1 | CM 9634 |  |  | < 0.5 mm | <1° | Good | Small bubble seen | good
4-7-2 | 4mm | 7_2 | CM 9626 |  |  | < 0.5 mm | <1° | Good | Many small bubbles in corner | to test
4-7-3 | 4mm | 7_3 |  |  |  | < 0.5 mm | <1° | Good | Good | good
4-7-4 | 4mm | 7_4 |  |  |  | < 0.5 mm | <1° | Good | Tiny bubble near edge | good
4-7-5 | 4mm | 7_5 |  |  |  | < 0.5 mm | <1° | Good | Very tiny bubble in center | good
4-7-6 | 4mm | 7_6 | no label 9 |  |  | < 0.5 mm | <1° | Dirty | Cloudy | bad |  | 12/3 fk | bn | 12/4 bn | 12/10 jp | 12/17 bn
4-7-7 | 4mm | 7_7 |  |  |  | < 0.5 mm | <1° | Dirty | Tiny bubbles seen | good
4-7-8 | 4mm | 7_8 |  |  |  | < 0.5 mm | <1° | Discoloration at tip | Good | polish
4-7-9 | 4mm | 7_9 |  |  |  | < 0.5 mm | <1° | Small scratch on light guide | Good | to test
4-7-10 | 4mm | 7_10 |  |  |  | < 0.5 mm | <1° | Good | Good | good
4-7-11 | 4mm | 7_11 |  |  |  | < 0.5 mm | <1° | Good | Good | good
4-7-12 | 4mm | 7_12 |  |  |  | < 0.5 mm | <1° | Good | Good | good
4-7-13 | 4mm | 7_13 |  |  |  | < 0.5 mm | <1° | Good | Good | good
4-7-14 | 4mm | 7_14 |  |  |  | < 0.5 mm | <1° | Good | Good | good
4-7-15 | 4mm | 7_15 |  |  |  | < 0.5 mm | <1° | Dirty | Good | polish
4-7-16 | 4mm | 7_16 |  |  |  | < 0.5 mm | <1° | Dirty | Good | polish
5-8-1 | 5mm | 8_1 |  |  |  | 0.043 mm | 0.204° | Good | Good | good
5-8-2 | 5mm | 8_2 |  |  |  | 0.067 mm | 0.119° | Good | Small bubbles | to test
5-8-3 | 5mm | 8_3 |  |  |  | 0.128 mm | 0.027° | Good | Small bubbles | to test
5-8-4 | 5mm | 8_4 | CM 9879 |  |  | < 0.5 mm | <1° | Good | Cloudy | to test |  | 12/3 fk | bn | 12/4 bn | 12/12 jp | 12/16 jp
5-8-5 | 5mm | 8_5 |  |  |  | < 0.5 mm | <1° | Good | Good | good
5-8-6 | 5mm | 8_6 | CM 9862 |  |  | < 0.5 mm | <1° | Good | Tiny bubbles | good |  | 4/24 fk | bn | 4/25 bn | 4/28 bn
5-8-7 | 5mm | 8_7 | CM 9778 |  |  | < 0.5 mm | <1° | Dirty | Good | polish |  | 4/24 fk | bn | 4/25 bn | 4/28 bn
5-8-8 | 5mm | 8_8 | CM 9863 |  |  | < 0.5 mm | <1° | Good | Cloudy | to test |  | 4/24 fk | bn | 4/25 bn | 4/28 bn
5-8-9 | 5mm | 8_9 |  |  |  | < 0.5 mm | <1° | Good | Good | good
5-8-10 | 5mm | 8_10 | CM 9787 |  |  | < 0.5 mm | <1° | Dirty | Good | polish |  | 4/24 fk | bn | 4/25 bn | 4/28 bn
5-8-11 | 5mm | 8_11 |  |  |  | < 0.5 mm | <1° | Good | Tiny bubbles near center | to test
5-8-12 | 5mm | 8_12 |  |  |  | < 0.5 mm | <1° | Dirty | Good | polish
5-8-13 | 5mm | 8_13 |  |  |  | < 0.5 mm | <1° | Dirty | Good | polish
5-9-1 | 5mm | 9_1 |  |  |  | < 0.5 mm | <1° | Good | Good | good
5-9-2 | 5mm | 9_2 |  |  |  | < 0.5 mm | <1° | Good | Little bubble on side | to test
5-9-3 | 5mm | 9_3 |  |  |  | < 0.5 mm | <1° | Good | Good | good
5-9-4 | 5mm | 9_4 | CM 9352 |  |  | < 0.5 mm | <1° | Dirty | Cloudy | to test |  | 3/26 fk | jp | 3/27 jp | 3/31 jp
5-9-5 | 5mm | 9_5 | CM 9259 |  |  | < 0.5 mm | <1° | Good | Tiny bubble | good |  | 3/25 fk | jp | 3/26 jp | 3/31 jp
5-9-6 | 5mm | 9_6 | CM 9252 |  |  | < 0.5 mm | <1° | Good | Good | good |  | 3/25 fk | jp | 3/26 jp | 3/31 jp
5-9-7 | 5mm | 9_7 | CM 9502 |  |  | < 0.5 mm | <1° | Good | Good | good |  | 3/26 fk | jp | 3/27 jp | 3/31 jp
5-9-8 | 5mm | 9_8 | CM 9255 |  |  | < 0.5 mm | <1° | Good | Good | good |  | 3/25 fk | jp | 3/26 jp | 3/31 jp
5-9-9 | 5mm | 9_9 | CM 9253 |  |  | < 0.5 mm | <1° | Good | Good | good |  | 3/25 fk | jp | 3/26 jp | 3/31 jp
5-9-10 | 5mm | 9_10 | no label 7 |  |  | < 0.5 mm | <1° | Good | Cloudy | to test |  | 12/3 fk | bn | 12/5 bn | 12/12 jp | 1/8 jp
5-9-11 | 5mm | 9_11 | CM 9352 |  |  | < 0.5 mm | <1° | Good | Small bubbles | to test
5-9-12 | 5mm | 9_12 |  |  |  | < 0.5 mm | ~1.5° | Dirty | Cloudy | bad
5-9-13 | 5mm | 9_13 | CM 9456 |  |  | < 0.5 mm | <1° | Good | Good | good |  | 3/25 fk | jp | 3/26 jp | 4/1 jp | 4/16 jp
5-9-14 | 5mm | 9_14 | CM 9555 |  |  | < 0.5 mm | <1° | Good | Small bubble on side | to test
5-9-15 | 5mm | 9_15 |  |  |  | < 0.5 mm | <1° | Good | Cloudy | to test
5-9-16 | 5mm | 9_16 | CM 9493 |  |  | < 0.5 mm | <1° | Good | Good | good |  | 3/26 fk | jp | 3/27 jp | 4/1 jp | 4/16 jp
5-9-17 | 5mm | 9_17 | CM 9501 |  |  | < 0.5 mm | <1° | Good | Good | good |  | 3/26 fk | jp | 3/27 jp | 4/1 jp | 4/16 jp
5-9-18 | 5mm | 9_18 | CM 9353 |  |  | < 0.5 mm | <1° | Good | Tiny bubbles seen | good |  | 3/26 fk | jp | 3/27 jp | 4/1 jp | 4/16 jp
5-10-1 | 5mm | 10_1 |  |  |  | < 0.5 mm | <1° | Good | Cloudy | to test
5-10-2 | 5mm | 10_2 | CM 9897 |  |  | < 0.5 mm | <1° | Good | Small bubble in corner | to test |  | 2/6 fk | jp | 2/10 jp | 2/12 jp | 2/27 jp
5-10-3 | 5mm | 10_3 | CM 9231 |  |  | < 0.5 mm | <1° | Good | Bubbles | to test |  | 2/6 fk | jp | 2/10 jp | 2/12 jp | 2/27 jp
5-10-4 | 5mm | 10_4 | CM 9901 |  |  | < 0.5 mm | <1° | Good | Cloudy | to test |  | 2/6 fk | jp | 2/10 jp | 2/12 jp | 2/27 jp
5-10-5 | 5mm | 10_5 | CM 9928 |  |  | < 0.5 mm | <1° | Good | Good | good |  | 2/11 fk | jp | 2/12 jp | 2/13 jp | 3/5 jp
5-10-6 | 5mm | 10_6 | CM 9887 |  |  | < 0.5 mm | <1° | Good | Small bubble on side | to test |  | 2/18 fk | jp | 2/19 jp | 2/20 jp | 2/27 jp
5-10-7 | 5mm | 10_7 | CM 9946 |  |  | < 0.5 mm | <1° | Good | Tiny bubbles near corner | good |  | 2/6 fk | jp | 2/10 jp | 2/12 jp | 3/5 jp
5-10-8 | 5mm | 10_8 | CM 9935 |  |  | < 0.5 mm | <1° | Good | Good | good |  | 2/6 fk | jp | 2/10 jp | 2/13 jp | 3/5 jp
5-10-9 | 5mm | 10_9 | CM 9389 |  |  | < 0.5 mm | <1° | Good | Bubbles | to test |  | 2/11 fk | jp | 2/12 jp | 2/13 jp | 3/5 jp
5-10-10 | 5mm | 10_10 | CM 9908 |  |  | < 0.5 mm | <1° | Good | Good | good |  | 2/11 fk | jp | 2/12 jp | 2/13 jp | 3/19 jp
5-10-11 | 5mm | 10_11 |  |  |  | < 0.5 mm | <1° | Good | Cloudy | to test
5-10-12 | 5mm | 10_12 | CM 9915 |  |  | < 0.5 mm | <1° | Good | Good | good |  | 2/6 fk | jp | 2/10 jp | 2/12 jp | 3/19 jp
5-10-13 | 5mm | 10_13 | CM 9909 |  |  | < 0.5 mm | <1° | Good | Good | good |  | 2/18 fk | jp | 2/20 jp | 2/12 jp | 3/20 jp
5-10-14 | 5mm | 10_14 | CM 9391 |  |  | < 0.5 mm | <1° | Good | Good | good |  | 2/11 fk | jp | 2/13 jp | 2/13 jp | 3/20 jp
5-10-15 | 5mm | 10_15 | CM 9371 |  |  | < 0.5 mm | <1° | Good | Cloudy | to test |  | 2/18 fk | jp | 2/19 jp | 2/24 jp | 3/19 jp
5-10-16 | 5mm | 10_16 | CM 9886 |  |  | < 0.5 mm | <1° | Good | Good | good |  | 2/11 fk | jp | 2/13 jp | 2/24 jp | 3/20 jp
5-10-17 | 5mm | 10_17 | CM 9889 |  |  | < 0.5 mm | <1° | Good | Good | good |  | 2/18 fk | jp | 2/19 jp | 2/24 jp | 3/17 jp
5-10-18 | 5mm | 10_18 | CM 9892 |  |  | < 0.5 mm | <1° | Good | Good | good |  | 2/11 fk | jp | 2/13 jp | 2/24 jp | 3/17 jp
5-11-1 | 5mm | 11_1 |  |  |  | < 0.5 mm | <1° | Good | Good | good
5-11-2 | 5mm | 11_2 | CM 9952 |  |  | < 0.5 mm | <1° | Good | Good | good |  | 2/24 fk | jp | 2/25 jp | 3/3 jp | 3/17 jp
5-11-3 | 5mm | 11_3 | CM 9948 |  |  | < 0.5 mm | <1° | Good | Small bubbles | to test |  | 2/24 fk?
5-11-4 | 5mm | 11_4 | CM 9960 |  |  | < 0.5 mm | <1° | Good | Good | good |  | 1/9 fk | bn | 1/10 bn | 1/15 jp | 1/24 bn
5-11-5 | 5mm | 11_5 | CM 9950 |  |  | < 0.5 mm | <1° | Good | Good | good |  | 1/9 fk | bn | 1/10 bn | 1/15 jp | 1/24 bn
5-11-6 | 5mm | 11_6 | CM 9457 |  |  | < 0.5 mm | <1° | Good | Good | good |  | 2/24 fk | jp | 2/25 jp | 3/3 jp | 3/24 jp
5-11-7 | 5mm | 11_7 | CM 9914 |  |  | < 0.5 mm | <1° | Good | Good | good |  | 1/9 fk | bn | 1/10 bn | 1/15 jp | 1/24 bn
5-11-8 | 5mm | 11_8 | CM 9385 |  |  | < 0.5 mm | <1° | Good | Good | good |  | 2/24 fk | jp | 2/26 jp | 3/3 jp | 3/24 jp
5-11-9 | 5mm | 11_9 | CM 9463 |  |  | < 0.5 mm | <1° | Good | Good | good |  | 2/24 fk | jp | 2/25 jp | 3/3 jp | 3/24 jp
5-11-10 | 5mm | 11_10 | CM 9464 |  |  | < 0.5 mm | <1° | Good | Good | good |  | 2/24 fk | jp | 2/25 jp | 3/3 jp | 3/24 jp
5-11-11 | 5mm | 11_11 |  |  |  | < 0.5 mm | <1° | Good | Good | good
5-11-12 | 5mm | 11_12 |  |  |  | < 0.5 mm | <1° | Dirty | Good | polish
5-11-13 | 5mm | 11_13 | CM 9922 |  |  | < 0.5 mm | <1° | Good | Good | good |  | 1/9 fk | bn | 1/13 bn | 1/15 jp | 1/24 bn
8-12-1 | 8mm | 12_1 | CM 9878 |  |  | < 0.5 mm | <1° | Good | Good | good |  | 1/14 fk | fk | 1/16 bn | 1/20 jp | 1/24 bn
8-12-2 | 8mm | 12_2 | CM 9875 |  |  | < 0.5 mm | <1° | Dirty | Good | polish |  | 1/14 fk | fk | 1/16 bn | 1/20 jp | 1/24 bn
8-12-3 | 8mm | 12_3 | CM 9870 |  |  | < 0.5 mm | <1° | Good | Good | good |  | 1/14 fk | fk | 1/16 bn | 1/20 jp | 1/23 jp
8-12-4 | 8mm | 12_4 | CM 9876 |  |  | < 0.5 mm | <1° | Dirty | Good | polish |  | 1/14 fk | fk | 1/16 bn | 1/20 jp | 1/23 jp
8-12-5 | 8mm | 12_5 | CM 9921 |  |  | < 0.5 mm | <1° | Good | Good | good |  | 1/16 fk | bn | 1/17 bn | 1/20 jp | 1/23 jp
8-12-6 | 8mm | 12_6 | CM 9918 |  |  | < 0.5 mm | <1° | Good | Tiny bubbles | to test |  | 1/16 fk | bn | 1/17 bn | 1/20 jp | 1/23 jp
8-12-7 | 8mm | 12_7 | CM 9387 |  |  | < 0.5 mm | <1° | Dirty | Good | polish |  | 1/16 fk | bn | 1/17 bn | 1/20 jp | 1/24 bn
8-12-8 | 8mm | 12_8 | CM 9925 |  |  | < 0.5 mm | <1° | Good | Good | good |  | 1/16 fk | bn | 1/17 bn | 1/20 jp | 1/23 jp
8-12-9 | 8mm | 12_9 | CM 9930 |  |  | < 0.5 mm | <1° | Dirty | Good | polish |  | 1/16 fk | bn | 1/17 bn | 1/20 jp | 1/24 bn
8-12-10 | 8mm | 12_10 | CM 9881 |  |  | < 0.5 mm | <1° | Good | Good | good |  | 1/14 fk | fk | 1/16 bn | 1/20 jp | 1/24 bn
8-12-11 | 8mm | 12_11 | CM 9871 |  |  | < 0.5 mm | <1° | Good | Good | good |  | 1/16 fk | bn | 1/20 bn | 1/20 jp | 1/24 bn
8-12-12 | 8mm | 12_12 | CM 9877 |  |  | < 0.5 mm | <1° | Good | Good | good |  | 1/14 fk | fk | 1/17 bn | 1/20 jp | 1/23 bn
8-13-1 | 8mm | 13_1 | CM 9381 |  |  | 0.112 mm | 0.038° | Good | Good | good |  | 1/28 fk | bn | 1/30 jp | 2/5 jp | 2/19 jp
8-13-2 | 8mm | 13_2 | CM 9445 |  |  | 0.025 mm | 0.065° | Good | Good | good |  | 1/28 fk | bn | 1/30 bn | 2/5 jp | 2/19 jp
8-13-3 | 8mm | 13_3 | CM 9938 |  |  | < 0.5 mm | <1° | Good | Good | good |  | 1/30 fk | jp | 2/4 jp | 2/5 jp | 2/26 jp
8-13-4 | 8mm | 13_4 | CM 9250 |  |  | < 0.5 mm | <1° | Good | Good | good |  | 1/28 fk | bn | 1/30 bn | 2/5 jp | 2/26 jp
8-13-5 | 8mm | 13_5 | CM 9446 |  |  | < 0.5 mm | <1° | Good | Good | good |  | 1/28 fk | bn | 1/30 bn | 2/5 jp | 2/26 jp
8-13-6 | 8mm | 13_6 | CM 9791 |  |  | < 0.5 mm | <1° | Good | Cloudy | to test |  | 12/3 fk | bn | 12/12 bn | 1/6 jp | 1/23 bn
8-13-7 | 8mm | 13_7 | CM 9441 |  |  | < 0.5 mm | <1° | Good | Good | good |  | 1/28 fk | bn | 1/30 jp | 2/5 jp | 2/19 jp
8-13-8 | 8mm | 13_8 | no label 2 |  |  | < 0.5 mm | <1° | Dirty | Good | polish |  | 12/17 fk | bn | 1/6 bn | 1/9 jp | 1/23 bn
8-13-9 | 8mm | 13_9 |  |  |  | < 0.5 mm | <1° | Good | Cloudy | to test
8-13-10 | 8mm | 13_10 | CM 9444 |  |  | < 0.5 mm | <1° | Good | Good | good |  | 1/28 fk | bn | 2/4 jp | 2/5 jp | 3/24 jp
10-14-1 | 10mm | 14_1 |  |  |  | 0.240 mm | 0.533° | Good | Good | good
10-14-2 | 10mm | 14_2 |  |  |  | 0.203 mm | 0.316° | Good | Lots of small bubbles | to test
10-14-3 | 10mm | 14_3 | CM 9234 |  |  | < 0.5 mm | <1° | Good | Good | good |  | 1/30 fk | jp | 2/4 jp | 2/5 jp | 2/25 jp
10-14-4 | 10mm | 14_4 | CM 9890 |  |  | < 0.5 mm | <1° | Good | Good | good |  | 1/13 fk | bn | 1/16 bn | 1/20 jp | 1/24 bn
10-14-5 | 10mm | 14_5 | CM 9236 |  |  | < 0.5 mm | <1° | Dirty | Good | polish |  | 1/30 fk | jp | 2/4 jp | 2/5 jp | 2/25 jp
10-14-6 | 10mm | 14_6 | CM 9639 |  |  | < 0.5 mm | <1° | Good | Good | good |  | 12/17 fk | JP | 1/7 jp | 1/20 jp | 1/28 bn
10-14-7 | 10mm | 14_7 | CM 9594 |  |  | < 0.5 mm | <1° | Good | Good | good |  | 12/17 fk | jp | 1/7 jp | 1/20 jp | 1/28 bn
10-14-8 | 10mm | 14_8 | CM 9392 |  |  | < 0.5 mm | <1° | Good | Good | good |  | 1/13 fk | bn | 1/16 bn | 1/20 jp | 1/28 bn
10-14-9 | 10mm | 14_9 | CM 9941 |  |  | < 0.5 mm | <1° | Good | Good | good |  | 1/30 fk | jp | 2/6 jp | 2/11 jp | 2/24 jp
10-14-10 | 10mm | 14_10 | CM 9936 |  |  | < 0.5 mm | <1° | Good | Good | good |  | 1/30 fk | jp | 2/6 jp | 2/11 jp | 2/25 jp
10-14-11 | 10mm | 14_11 | CM 9372 |  |  | < 0.5 mm | <1° | Good | Tiny bubbles | good |  | 1/13 fk | bn | 1/16 bn | 1/20 jp | 1/28 bn
10-14-12 | 10mm | 14_12 | CM 9225 |  |  | < 0.5 mm | <1° | Good | Good | good |  | 1/30 fk | jp | 2/6 jp | 2/11 jp | 2/27 jp
10-14-13 | 10mm | 14_13 | CM 9743 |  |  | < 0.5 mm | <1° | Good | Cloudy | to test |  | 1/13 fk | bn | 1/17 bn | 1/21 jp | 1/28 bn
10-14-14 | 10mm | 14_14 | CM 9867 |  |  | < 0.5 mm | <1° | Good | Good | good |  | 1/13 fk | bn | 1/17 bn | 1/21 jp | 1/24 bn
10-14-15 | 10mm | 14_15 | CM 9884 |  |  | < 0.5 mm | <1° | Good | Small bubbles | to test |  | 1/13 fk | bn | 1/16 bn | 1/20 jp | 1/24 bn
10-15-1 | 10mm | 15_1 | CM 9780 |  |  | < 0.5 mm | <1° | Good | Good | good |  | 4/24 fk | bn | 4/28 bn | 4/30 bn
10-15-2 | 10mm | 15_2 | CM 9782 |  |  | < 0.5 mm | <1° | Good | Good | good |  | 4/15 fk | bn | 4/17 bn | 4/23 bn
10-15-3 | 10mm | 15_3 | CM 9796 |  |  | < 0.5 mm | <1° | Dirty | Good | polish |  | 4/22 fk | bn | 4/24 bn | 4/30 bn
10-15-4 | 10mm | 15_4 | CM 9945 |  |  | < 0.5 mm | <1° | Good | Good | good |  | 4/22 fk | bn | 4/24 bn | 4/30 bn
10-15-5 | 10mm | 15_5 | CM 9857 |  |  | < 0.5 mm | <1° | Good | Tiny bubbles throughout | to test |  | 4/22 fk | bn | 4/24 bn | 4/30 bn
10-15-6 | 10mm | 15_6 | CM 9820 |  |  | < 0.5 mm | <1° | Good | Tiny bubbles | good |  | 4/22 fk | bn | 4/24 bn | 4/30 bn
10-15-7 | 10mm | 15_7 | CM 9893 |  |  | < 0.5 mm | ~1.5° | Small scratch on scintilator | Good | polish |  | 4/24 fk | bn | 4/28 bn | 4/30 bn
10-15-8 | 10mm | 15_8 | CM 9904 |  |  | < 0.5 mm | <1° | Dirty | Good | polish |  | 4/15 fk | bn | 4/17 bn | 4/23 bn
10-15-9 | 10mm | 15_9 | CM 9854 |  |  | < 0.5 mm | <1° | Dirty | Good | polish |  | 4/15 fk | bn | 4/18 bn | 4/23 bn
10-15-10 | 10mm | 15_10 | CM 9798 |  |  | < 0.5 mm | <1° | Good | Good | good |  | 4/15 fk | bn | 4/18 bn | 4/30 bn
10-15-11 | 10mm | 15_11 | CM 9799 |  |  | < 0.5 mm | <1° | Good | Good | good |  | 4/15 fk | bn | 4/18 bn | 4/23 bn
10-15-12 | 10mm | 15_12 | CM 9795 |  |  | < 0.5 mm | <1° | Good | Good | good |  | 4/22 fk | bn | 4/24 bn | 4/30 bn
10-15-13 | 10mm | 15_13 | CM 9926 |  |  | < 0.5 mm | <1° | Good | Good | good |  | 4/22 fk | bn | 4/28 bn | 4/30 bn
10-15-14 | 10mm | 15_14 | CM 9819 |  |  | < 0.5 mm | <1° | Good | Good | good |  | 4/22 fk | bn | 4/28 bn | 4/30 bn
16-16-1 | 16mm | 16_1 |  |  |  | 0.112 mm | 1.086° | Good | Good | good
16-16-2 | 16mm | 16_2 | CM 9504 |  |  | 0.191 mm | 0.096° | Good | Tiny bubbles | good |  | 4/7 fk | bn | 4/9 bn | 4/25 bn
16-16-3 | 16mm | 16_3 | CM 9433 |  |  | < 0.5 mm | <1° | Scratch on the scintilator | Good | polish |  | 4/7 fk | bn | 4/9 bn | 4/25 bn
16-16-4 | 16mm | 16_4 | CM 9511 |  |  | < 0.5 mm | <1° | Good | Good | good |  | 4/7 fk | bn | 4/9 bn | 4/25 bn
16-16-5 | 16mm | 16_5 | CM 9593 |  |  | < 0.5 mm | <1° | Good | Cloudy | good |  | 12/10 fk | fk | 12/16 bn | 1/8 jp
16-16-6 | 16mm | 16_6 | CM 9850 |  |  | < 0.5 mm | <1° | Small chip on scintilator | Good | polish |  | 4/7 fk | bn | 4/9 bn | 4/25 bn | 4/30 jp
16-16-7 | 16mm | 16_7 |  |  |  | < 0.5 mm | <1° | Good | Cloudy | to test
16-16-8 | 16mm | 16_8 | CM 9858 |  |  | < 0.5 mm | <1° | Good | Tiny bubbles throughout | to test |  | 4/7 fk | bn | 4/16 bn | 4/30 bn
16-16-9 | 16mm | 16_9 | CM 9417 |  |  | < 0.5 mm | ~1.5° | Good | Large bubble seen at the corner | to test |  | 4/4 fk | fk | 4/9 bn | 4/25 bn | 5/1 bn
16-16-10 | 16mm | 16_10 | CM 9415 |  |  | < 0.5 mm | <1° | Good | Tiny bubbles | good |  | 4/8 fk | bn | 4/16 jp | 4/30 bn
16-16-11 | 16mm | 16_11 | CM 9873 |  |  | < 0.5 mm | <1° | Good | Good | good |  | 4/7 fk | bn | 4/9 bn | 4/28 jp | 5/1 bn
16-17-1 | 16mm | 17_1 | CM 9418 |  |  | < 0.5 mm | <1° | Good | Good | good |  | 4/8 fk | bn | 4/16 jp | 4/28 jp | 4/30 jp
16-17-2 | 16mm | 17_2 | CM 9847 |  |  | < 0.5 mm | <1° | Dirty | Tiny bubble | polish |  | 4/8 fk | bn | 4/16 bn | 4/28 jp | 4/30 jp
16-17-3 | 16mm | 17_3 | CM 9437 |  |  | < 0.5 mm | <1° | Good | Tiny bubbles on edge | good |  | 4/8 fk | bn | 4/16 bn | 4/30 bn
16-17-4 | 16mm | 17_4 | CM 9849 |  |  | < 0.5 mm | <1° | Good | Good | good |  | 4/8 fk | bn | 4/16 bn | 4/30 bn
16-17-5 | 16mm | 17_5 | CM 9856 |  |  | < 0.5 mm | <1° | Good | Small bubble on side | to test |  | 4/8 fk | bn | 4/16 jp | 4/25 bn | 4/30 jp
16-17-6 | 16mm | 17_6 | CM 9560 |  |  | < 0.5 mm | <1° | Dirty | Good | polish |  | 4/16 fk | bn | 4/21 bn | 4/29 jp | 5/1 bn
16-17-7 | 16mm | 17_7 | CM 9512 |  |  | < 0.5 mm | <1° | Good | Tiny bubbles throughout | to test |  | 4/16 fk | bn | 4/21 bn | 4/29 jp | 5/1 bn
16-17-8 | 16mm | 17_8 |  |  |  | < 0.5 mm | <1° | Good | Good | good
16-17-9 | 16mm | 17_9 | CM 9518 |  |  | < 0.5 mm | <1° | Dirty | Good | polish |  | 4/16 fk | bn | 4/21 bn | 4/30 jp
16-17-10 | 16mm | 17_10 | CM 9872 |  |  | < 0.5 mm | <1° | Good | Good | good |  | 4/16 fk | bn | 4/21 bn | 4/29 jp | 4/30 jp
16-17-11 | 16mm | 17_11 | CM 9509 |  |  | < 0.5 mm | <1° | Tiny chip in scintilator corner | Good | polish |  | 4/16 fk | bn | 4/21 bn | 4/30 jp
21-18-1 | 21mm | 18_1 | CM 9374 |  |  | < 0.5 mm | <1° | Good | Good | good |  | 12/10 fk | bn | 12/12 bn | 1/23 jp | 2/12 jp
21-18-2 | 21mm | 18_2 | CM 9373 |  |  | < 0.5 mm | <1° | Good | Good | good |  | 12/5 fk | fk | 12/9 bn | 1/23 jp | 2/12 jp
21-18-3 | 21mm | 18_3 | CM 9378 |  |  | < 0.5 mm | <1° | Good | Good | good |  | 12/5 fk | fk | 12/9 bn | 1/23 jp | 2/12 jp
21-18-4 | 21mm | 18_4 | CM 7558 |  |  | < 0.5 mm | <1° | Good | Tiny bubbles | good |  | 12/10 fk | bn | 12/12 bn | 1/15 jp | 2/24 jp
21-18-5 | 21mm | 18_5 | CM 9382 |  |  | < 0.5 mm | <1° | Good | Good | good |  | 12/5 fk | fk | 12/10 jp | 1/20 jp | 2/12 jp
21-18-6 | 21mm | 18_6 | CM 9384 |  |  | < 0.5 mm | <1° | Good | Good | good |  | 12/5 fk | fk | 12/10 jp | 1/20 jp | 2/13 jp
21-18-7 | 21mm | 18_7 | CM 9571 |  |  | < 0.5 mm | <1° | Good | Small bubbles near corner | good |  | 12/10 fk | bn | 12/16 bn | 1/15 jp | 2/13 jp
21-18-8 | 21mm | 18_8 | CM 9753 |  |  | < 0.5 mm | <1° | Good | Good | good |  | 12/10 fk | bn | 12/16 bn | 1/23 jp | 2/12 jp
21-18-9 | 21mm | 18_9 | CM 9386 |  |  | < 0.5 mm | <1° | Good | Good | good | 12/5 fk - redo
21-18-10 | 21mm | 18_10 | CM 9361 |  |  | < 0.5 mm | <1° | Good | Good | good |  | 12/5 fk | fk | 12/10 jp | 1/23 jp | 2/13 jp
21-19-1 | 21mm | 19_1 |  |  |  | 0.124 mm | 0.087° | Good | Good | good
21-19-2 | 21mm | 19_2 | CM 9317 |  |  | 0.132 mm | 0.253° | Good | Corner problem | to test |  | 4/1 fk | jp | 4/3 jp | 4/15 jp | 4/16 jp
21-19-3 | 21mm | 19_3 | CM 9551 |  |  | < 0.5 mm | <1° | Good | Good | good |  | 4/1 fk | jp | 4/3 jp | 4/15 jp | 4/16 jp
21-19-4 | 21mm | 19_4 | CM 9432 |  |  | < 0.5 mm | <1° | Good | Good | good |  | 4/4 fk | fk | 4/7 jp | 4/15 jp | 4/15 jp
21-19-5 | 21mm | 19_5 | CM 9801 |  |  | < 0.5 mm | <1° | Good | Good | good |  | 4/4 fk | fk | 4/7 jp | 4/15 jp | 4/15 jp
21-19-6 | 21mm | 19_6 | CM 9516 |  |  | < 0.5 mm | <1° | Good | Good | good |  | 4/4 fk | fk | 4/7 jp | 4/15 jp | 4/16 jp
21-19-7 | 21mm | 19_7 |  |  |  | < 0.5 mm | <1° | Good | Good | good
21-19-8 | 21mm | 19_8 | CM 9855 |  |  | < 0.5 mm | <1° | Good | Good | good |  | 4/4 fk | fk | 4/8 bn | 4/16 jp | 4/17 jp
21-19-9 | 21mm | 19_9 | CM 9478 |  |  | < 0.5 mm | <1° | Good | Good | good |  | 4/1 fk | jp | 4/3 jp | 4/22 bn | 4/22 bn
21-19-10 | 21mm | 19_10 | CM 9302 |  |  | < 0.5 mm | <1° | Good | Good | good |  | 4/1 fk | jp | 4/3 jp | 4/22 bn | 4/22 bn
21-19-11 | 21mm | 19_11 | CM 9888 |  |  | < 0.5 mm | <1° | Good | Cloudy | to test |  | 4/1 fk | jp | 4/2 jp | 4/16 jp | 4/17 jp
21-19-12 | 21mm | 19_12 | CM 9379 |  |  | < 0.5 mm | <1° | Good | Good | good |  | 4/4 fk | fk | 4/8 bn | 4/16 jp | 4/17 jp
21-19-13 | 21mm | 19_13 | CM 9485 |  |  | < 0.5 mm | <1° | Good | Good | good |  | 4/1 fk | jp | 4/2 jp | 4/16 jp | 4/16 jp
rotation in slot direction
translation offset to electron axis
electron path for 8, 10, 16, 21 mm counters
electron path for 3, 4, 5 mm counters